## Element categories_v00-01
| Category | Sub-category | Element
1 | MicroscopicBasic |  | Instrument, MicroscopyBody, MicroscopeTable
2 | SampleHolderAndFocusing |  | SampleHolder
| SampleHolderAndFocusing | SampleHolderAndFocusing.Focusing | MechanicalStage, PiezoElectricStage, ObjectiveTurret, PiezoElectricObjective, AutoFocus
3 | Magnification |  | Objective, Optovar
4 | LightSource |  | Laser, Arc, Filament, LightEmittingDiode, GenericExcitationSource
5 | LightSourceCoupling |  | LightGuide, Direct
6 | LightPath |  | FilterSet, ExcitationFilter, Dichroic, EmissionFilter
7 | Mirror |  | Reflector, Beamsplitter
8 | AdditionalOptics |  | Lens, Prism, PolarizationOptics, AdditionalFilter, Beamsplitter, Reflector
9 | Detector |  | PhotoMultiplier, GenricDetector
| Detector | Detector.Camera | CCD, IntensifiedCCD, AnalogVideo, CMOS
10 | Calibration |  | IntensityCalibrationTool
11 | ChildrenElement |  | Pump, LEDModule, WavelengthRange, TransmittanceRange, ReflectanceRance

## 4DN proposal_Specs_v01-03
4DN-OME model
| Tier | Element | Attribute or Subelement | Tier | Complex type | Description | Cardinality | Data type | Allowed values
MICROSCOPE Speecifications | 1 | Instrument |  | 1 |  | This element describes the instrument used to capture the Image. The Filters elements are described in the Light Path and can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. In addition the Light Path might contain additional optics such as Lens, Prism, Mirrors, Polarization Optics and other Filters. Within the Image itself, a reference is made to the Instrument elements and, if appropriate, their settings. There may be multiple light sources, detectors, objectives and filters on a microscope. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the light path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. | 1
| 1 | Instrument | ID | 1 |  |  | R | InstrumentID
| 1 | Instrument | Name | 1 |  | A user defined name for this instrument. | R | string
| 2 | Instrument | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 2 | AnnotationRef |  | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation.
| 2 | AnnotationRef | ID | 2 |  |  |  | AnnotationID
| 2 | CommentAnnotation |  | 2 |  | An annotation that consists of a simple multi-line comment.
| 2 | CommentAnnotation | ID | 2 |  |  |  | AnnotationID
| 5 | CommentAnnotation | Namespace | 5 |  | We recommend the inclusion of a namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type.
| 2 | CommentAnnotation | Annotator | 2 |  | The Annotator is the person who attached this annotation. e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator.
| 2 | CommentAnnotation/Description | Descrption | 2 |  | This element contains a description for the annotation.[plain-text multi-line string] |  | Element
| 2 | CommentAnnotation/Value | Value | 2 |  | This element contains the actual content of this annotation. [multi-line string] |  | Element
| 2 | CommentAnnotation/AnnotationRef | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. |  | Extension of Reference
| 1 | MicroscopeBody |  | 1 |  | A description of the body of this microscope. | 1 … ∞
| 1 | MicroscopeBody | Manufacturer | 1 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | MicroscopeBody | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | MicroscopeBody | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | MicroscopeBody | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | MicroscopeBody | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | MicroscopeBody | ID | 1 |  |  | R | MicroscopeID
| 1 | MicroscopeBody | Name | 1 |  | A user defined name for this element. | R | string
| 1 | MicroscopeBody | Type | 1 | MicroscopeType | This element describes the type of this microscope. | R | enum | Upright, Inverted, Dissection, Stereomicroscope, Electrophysiology, Custom, Other
| 3 | MicroscopeBody | EyepieceFieldNumber | 3 |  | The Field Number represents the diaphragm size of the microscope eyepiece expressed in mm. |  | float with unit:none
| 2 | MicroscopeBody | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 …1 | Extension of Reference
| 4 | MicroscopeTable |  | 4 |  | This element describes the table used to support this microscope | 1
| 4 | MicroscopeTable | Manufacturer | 4 | Manufacturer Spec | The manufacturer of the component. | R | string
| 4 | MicroscopeTable | Model | 4 | Manufacturer Spec | The model of the component. | R | string
| 5 | MicroscopeTable | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | MicroscopeTable | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | MicroscopeTable | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 4 | MicroscopeTable | ID | 4 |  |  | R | LSID
| 4 | MicroscopeTable | Name | 4 |  | A user defined name for this element | R | string
| 4 | MicroscopeTable | Type | 4 | MicroscopeTableTYpe | This element describes the type of this microscope table. | R | enum | Simple Bench, Non-isolating Frame, Active Air, Passive Air, Stone, Other
| 5 | MicroscopeTable | TableWidth | 5 |  | The width of this microscope table |  | float with unit:TableDepthUnit
| 5 | MicroscopeTable | TableDepth | 5 |  | The depth of this microscope table |  | float with unit:TableWidthUnit
| 4 | MicroscopeTable | AnnotationRef | 4 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 1 | LightSource-APE |  | 1 | LightSource | This element describes the light source(s) available in this microscope. | 1 … ∞ | Substitution : choose one of the children elements from the list below | Laser, Filament, Arc, LightEmittingDiode, GenericLightSource
| 1 | LightSource-APE | Manufacturer | 1 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | LightSource-APE | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | LightSource-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | LightSource-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | LightSource-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | LightSource-APE | ID | 1 |  | A LightSource ID must be specified for each light source, and the individual light sources can be referred to by their LightSource IDs (eg from Channel). | R | LightSourceID
| 1 | LightSource-APE | Name | 1 |  | A user defined name for this element | R | string
| 3 | LightSource-APE | NominalPower | 3 |  | The nominal power of the light-source | R | float with unit:PowerUnit
| 2 | LightSource-APE | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 1 | Laser |  | 1 |  | A laser ("light amplification by stimulated emission of radiation") is a device that emits light through a process of optical amplification based on the stimulated emission of electromagnetic radiation.
| 5 | Laser | Type | 5 | LaserType | The general category of this laser. | R | enum | Gas, MetalVapor, LaserDiode, DiodePumpedSolidState, LED, Dye, Semiconductor, FreeElectron, Other
| 5 | Laser | LaserMedium | 5 | LaserMedium | This is the material this laser consists of. | R | enum | Cu, Ag, ArFl, ArCl, KrFl, KrCl, XeFl, XeCl, XeBr, N, Ar, Kr, Xe, HeNe, HeCd, CO, CO2, H2O, HFl, NdGlass, NdYAG, ErGlass, ErYAG, HoYLF, HoYAG, Ruby, TiSapphire, Alexandrite, Rhodamine6G, CoumarinC30, GaAs, GaAlAs, EMinus, Other
| 2 | Laser | Wavelength | 2 |  | The center Wavelength of the light emitted by this laser. | R | float with unit:Wavelength
| 5 | Laser | FrequencyMultiplication | 5 |  | FrequencyMultiplication that may be specified. |  | positiveInt with unit:none
| 5 | Laser | Tuneable | 5 |  | Whether or not the laser is Tuneable. | R | boolean
| 5 | Laser | Pulse | 5 | LaserPulse | The Pulse mode of the laser. | R | enum | CW, Single, QSwitched, Repetitive, ModeLocked, Other
| 5 | Laser | PockelCell | 5 |  | If true the laser has a PockelCell to rotate the polarization of the beam. |  | boolean
| 5 | Laser | RepetitionRate | 5 |  | The rate in Hz at which the laser pulses if the Pulse type is 'Repetitive'. |  | foat with unit:RepetitionRateUnit
| 5 | Laser | Pump | 5 |  | The pump element is a reference to a LightSource. It is used within the Laser element to specify the light source for the laser's pump (if any). | 0 … 1 | Element
| 5 | Pump |  | 5 |  | The pump element is a reference to a LightSource. It is used within the Laser element to specify the light source for the laser's pump (if any).
| 5 | Pump | ID (LightSourceID) | 5
| 1 | Filament |  | 1 |  | The Filament element is used to describe various kinds of filament bulbs such as Incadescent or Halogen.
| 3 | Filament | Type | 3 | FilamentType | The type of Filament lamp. | R | enum | Incandescent, Halogen, Other
| 1 | Arc |  | 1 |  | This element describes various kinds of Arc lamps
| 3 | Arc | Type | 3 | ArcType | The type of Arc lamp. | R | enum | Hg, Xe, HgXe, Other
| 1 | LightEmittingDiode |  | 1 |  | The LightEmittingDiode element is used to describe various kinds of LED lamps.
| 3 | LightEmittingDiode | Type | 3 | LightEmittingDiodeType | The type of LightEmittingDiode light source | R | enum | Unmounted, Mounted, Collimate, Other
| 2 | LightEmittingDiode | LEDModule | 2 |  | A Light Emitting Diode (LED) light source is often composed of multiple modules each with its own wavelenght characteristics. This element describes an individual LED module. Multiple of this elements may be used. | 1 … ∞ | Element
| 2 | LEDModule | PeakWavelenght | 2 |  | The wavelength at which there is maximal response rate | R | float with unit: PeakWavelengthUnit
| 3 | LEDModule | LumenPower | 3 |  | The amount of light this light source generate at maximum expressed in lumen. The lumen is the SI derived unit of luminous flux, a measure of the total quantity of visible light emitted by a source | R | float with unit:lumen
| 5 | LEDModule | ProjectionAngle | 5 |  | The angle at which this light source projects light |  | integer with unit:ProjectionAngleUnit
| 3 | LEDModule | CutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% | R | integer with unit: CutInUnit
| 3 | LEDModule | CutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% | R | integer with unit: CutOutUnit
| 5 | LEDModule | WavelengthProfileFile | 5 |  | The location of a file containing the wavelenght profile of this component. | R | anyURI
| 1 | GenericExcitationSource |  | 1 |  | The GenericExcitationSource element is used to represent a source as a collection of key/value pairs, stored in a Map. The other lightsource objects should always be used in preference to this if possible.
| 4 | LightSourceCoupling-APE |  | 4 | LightSourceCoupling | This element describes the method(s) used to couple the light source with the light path moving through speciment to the detector. | 1 … ∞ | Substitution : choose one child from the list below | LighthGuide, Direct
| 4 | LightSourceCoupling-APE | Manufacturer | 4 | Manufacturer Spec | The manufacturer of the component. | R | string
| 4 | LightSourceCoupling-APE | Model | 4 | Manufacturer Spec | The model of the component. | R | string
| 5 | LightSourceCoupling-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | LightSourceCoupling-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | LightSourceCoupling-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 4 | LightSourceCoupling-APE | ID | 4 |  |  | R | LightSourceCouplingID
| 4 | LightSourceCoupling-APE | Name | 4 |  | A user defined name for this element | R | string
| 5 | LightSourceCoupling-APE | Transmittance | 5 |  | The amount of light this LightSourceCoupling apparatus transmits at a maximum. A fraction, as a value from 0.0 to 1.0 | R | float with unit:none
| 5 | LightSourceCoupling-APE | CouplingLensRef | 5 |  | This is a reference to a lens positioned at connector between this light source coupling element and the microscope body. | 0 … 1 | Extension of Reference
| 4 | LightSourceCoupling-APE | AnnotationRef | 4 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 4 | LightGuide |  | 4 |  | Light Guides interface with illuminators to transfer light to the speciment in a usable manner.
| 5 | LightGuide | Type | 5 | LightGuideType | The type of light guide used in this case. | R | enum | Fiber optics, Liquid, Other
| 5 | LightGuide | PeakWavelength | 5 |  | The wavelength associated with maximum light intensity transmitted by this light guide | R | float with unit:PeakWavelengthUnit
| 5 | LightGuide | RefractiveIndex | 5 |  | The refractive index of the material this light guide is made of. | R | float with unit:none
| 5 | LightGuide | NumericalAperture | 5 |  | The numerical aperture of the light guide expressed as a floating point (real) number. Expected range 0.02 - 1.5. | R | float with unit:none
| 5 | LightGuide | AcceptanceAngle | 5 |  | The angle within which light beam may be coupled into the bound modes of an optical fiber. | R | float with unit:AcceptanceAngleUnit
| 5 | LightGuide | Geometry | 5 | LightGuideGeometry | The shape of this light guide |  | enum | Spot Light, Line Light, Ring Light, Backlight, Diffuse Axial, Other
| 5 | LightGuide | MaterialName | 5 | LightGuideMaterialName | This records the chemical composition or the nonproprietary trade designation of the material used for this light guide. |  | enum | Quartz, Silica, Borosilicate, Other
| 5 | LightGuide | SingleMode | 5 |  | Records whether or not the light guide is single-mode. If not, then the guide is assumend to be multi-modal. [flag] |  | boolean
| 5 | LightGuide | WaveguideMode | 5 | LightGuideWaveguideMode | When boundary conditions are imposed by a physical structure, a wave of a particular frequency can be described in terms of a transverse mode (or superposition of such modes). These modes generally follow different propagation constants. |  | enum | Transverse electromagnetic, Transverse electric, Transverse magnetic, Hybrid, Other
| 5 | LightGuide | Diameter | 5 |  | The outer diameter of the light guide |  | float with unit:DiameterUnit
| 5 | LightGuide | WavelengthRange | 5 |  | This records a range of wavelengths associated with the associated microscope component. Multiple of this elements may be used. | 1 | Element
| 5 | WavelengthRange | CutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component. | R | anyURI
| 4 | Direct |  | 4 |  | Direct coupling consists of either direct mounting or a optical light path that connects the Light Source with the microscope.
| 5 | Direct | Type | 5 | DirectType | The type of direct coupling available for this Instrument. | R | enum | Mounted, Optical light path, Other
| 5 | Direct | Mount | 5 | DirectMount | The mount utilized to attach a LightSource to the microscope. | R | string
| 3 | SampleHolder |  | 3 |  | This element describes an optional sample holder mounted on the microscope stage | 1 … ∞
| 3 | SampleHolder | Manufacturer | 3 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | SampleHolder | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | SampleHolder | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | SampleHolder | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | SampleHolder | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | SampleHolder | ID | 3 | Stage |  | R | SampleHolderID
| 3 | SampleHolder | Name | 3 | Stage | A user defined name for this element | R | string
| 5 | SampleHolder | MovementAxes | 5 | Stage | The axis along which this device moves | R | enum | X, Y, Z, XY, XYZ, Other
| 5 | SampleHolder | XTravelRange | 5 | Stage | The X axis range this device is capable of moving in | R | Float with unit:XAxisTravelRangeUnit
| 5 | SampleHolder | YTravelRange | 5 | State | The Y axis range this device is capable of moving in | R | Float with unit:YAxisTravelRangeUnit
| 5 | SampleHolder | ZTravelRange | 5 | Stage | The Z axis range this device is capable of moving in | R | Float with unit:ZAxisTravelRangeUnit
| 5 | SampleHolder | MaxVelocity | 5 | Stage | The maximum velocity this device is capable of moving at |  | Float with unit:MaxVelocityUnit
| 5 | SampleHolder | Type | 5 |  | The type of sample holder. | R | enum | Rigid Stand, Translational, Other
| 5 | SampleHolder | RotationAngle | 5 |  | The angle by which this device is capable of rotating |  | Float with unit:RotationAngleUnit
| 3 | SampleHolder | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 3 | Focusing-APE |  | 3 | Substitution : choose one child from the list below | This element describes the objective lens and stage positioning device(s) available in this microscope. An instrument may have more than one of these devices. The type of focusing device is specified by one of the child-element which are 'MechanicalStage', 'PiezoElectricSage', 'ObjectiveTurret' or 'PiezoElectricObjective'. | 1 … ∞ |  | MechanicStage, PiezoelectricStage, PiezoelectricObjective
| 3 | Focusing-APE | Manufacturer | 3 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Focusing-APE | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Focusing-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Focusing-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Focusing-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | Focusing-APE | ID | 3 | Stage |  | R | FocusingID
| 3 | Focusing-APE | Name | 3 | Stage | A user defined name for this element | R | string
| 5 | Focusing-APE | MovementAxes | 5 | Stage | The axis along which this device moves | R | enum | X, Y, Z, XY, XYZ, Other
| 5 | Focusing-APE | XTravelRange | 5 | Stage | The X axis range this device is capable of moving in. | R | float with unit:XTravelRangeUnit
| 5 | Focusing-APE | YTravelRange | 5 | Stage | The Y axis range this device is capable of moving in | R | float with unit:YTravelRangeUnit
| 5 | Focusing-APE | ZTravelRange | 5 | State | The Z axis range this device is capable of moving in | R | float with unit:ZTravelRangeUnit
| 5 | Focusing-APE | MaxVelocity | 5 | Stage | The maximum velocity this device is capable of moving at |  | float with unit:MaximumVelocity
| 5 | Focusing-APE | Resolution | 5 |  | The expected motion resolution of this device. | R | float with unit:ResolutionUnit
| 5 | Focusing-APE | PositionLinearityError | 5 |  | The expected positioning error associated with this device. Typically expressed as % error. | R | float with unit:none
| 5 | Focusing-APE | Repetability | 5 |  | The expected precision with which this device is capable of returning to a previously visited position. Typically expressed in % error. | R | float with unit:none
| 5 | Focusing-APE | SettlingTime | 5 |  | The expected time it takes for this device to stabilize at the target position. | R | float with unit:SettlingTimeUnit
| 2 | Focusing-APE | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 3 | MechanicalStage |  | 3 |  | This is a mechanical stage positioning device. It is not very common but can be found in simple surveying microscopes.
| 3 | PiezoElectricStage |  | 3 |  | This is a piezoelectric stage positioning device.
| 3 | ObjectiveTurret |  | 3 |  | This is a objective turret positoning device
| 3 | PiezoElectricObjective |  | 3 |  | This is a piezoelectric objective-lens positioning device.
| 3 | AutoFocus |  | 3 |  | This element describes the focus stabilization device(s) available for this microscope | 0 … ∞
| 3 | AutoFocus | Manufacturer | 3 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | AutoFocus | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | AutoFocus | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | AutoFocus | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | AutoFocus | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | AutoFocus | ID | 3 |  |  | R | AutoFocusID
| 3 | AutoFocus | Name | 3 |  | A user defined name for this element | R | string
| 3 | AutoFocus | Type | 3 | AutoFocusType | The type of this autofocus device. | R | enum | Image based, Optical, Other
| 3 | AutoFocus | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 1 | Optovar |  | 1 |  | Optovars are magnification changing devices originally developed by Zeiss, which provide additional magnification that can be beneficial when a desired field of view is required for the imaging sensor. Magnification changer components are placed before the tube lens within the infinity space of the optical system. Typical magnification increments are: 1.25x, 1.6x, and 2x. | 0 … ∞
| 1 | Optovar | Manufacturer | 1 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Optovar | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Optovar | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Optovar | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Optovar | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | Optovar | ID | 1 |  |  | R | ID
| 1 | Optovar | Name | 1 |  | A user defined name for this element | R | Denomination
| 3 | Optovar | LensType | 3 |  | The type of lens used by the Magnification Changer available with this microscope. |  | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 2 | Optovar | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 1 | Objective |  | 1 |  | The microscope's objective lens is a compound lens consisting of several simple lenses (elements), usually arranged along a common axis. | 1 … ∞
| 1 | Objective | Manufacturer | 1 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Objective | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Objective | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Objective | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Objective | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | Objective | ID | 1 |  |  | R | ObjectiveID
| 1 | Objective | Name | 1 |  | A user defined name for this element | R | Denomination
| 2 | Objective | Correction | 2 | ObjectiveCorrection | The correction applied to the lens | R | enum | Achro, Achromat, Achroplan, Apo, Fl, Fluar, Fluor, Fluotar, Neofluar, Pl, Plan, Plano, PlanApo, PlanFluor, PlanNeoFluar, SuperFluor, UV, VioletCorrected, Other
| 1 | Objective | NominalMagnification | 1 |  | The magnification of the lens as specified by the manufacturer - i.e. '60' is a 60X lens | R | float with unit:none
| 5 | Objective | CalibratedMagnification | 5 |  | The magnification of the lens as measured by a calibration process- i.e. '59.987' for a 60X lens |  | float with unit:none
| 1 | Objective | LensNA | 1 |  | The numerical aperture of the lens expressed as a floating point (real) number. Expected range 0.02 - 1.5. | R | float with unit:none
| 2 | Objective | NominalImmersionMedium | 2 | ObjectiveNominalImmmersionMedium | The immersion medium the lens is designed for | R | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 3 | Objective | Infinity | 3 |  | Records whether or not the objective is infinity corrected | R | boolean
| 2 | Objective | CoverslipTickness | 2 |  | The coverlip thickness this objective is recommended for | R | float with unit:none
| 3 | Objective | WorkingDistance | 3 |  | The working distance of the lens expressed as a floating point (real) number | R | float with unit:WorkingDistanceUnit
| 4 | Objective | Iris | 4 |  | Records whether or not the objective was fitted with an Iris. [flag] | R | boolean
| 3 | Objective | ObjectiveFieldNumber | 3 |  | The maximal diameter of the viewfield allowed by this objective expressed in mm and measured at the intermediate image plane. NB: The field of view is influenced by the magnification and diameter of the ocular (eyepiece) field diaphragm, but also by the limits imposed by the design of the objective. | R | float with unit:none
| 5 | Objective | ParfocalizingDistance | 5 |  | The distance between the objective lens mounting plane and the specimen. | R | float with unit:ParafocalizingDistanceUnit[default mm]
| 5 | Objective | BackFocalPlanePosition | 5 |  | The position of the back focal plane with respect to the specimen position. | R | float with unit:BackFocalPlanePositionUnit[default mm]
| 2 | Objective | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 2 | Lens |  | 2 |  | A lens is a component of glass or transparent plastic material, usually circular in diameter, which has two primary surfaces that are ground and polished in a specific manner designed to produce either a convergence or divergence of light passing through the material. This element describes a lens utilized in this microscope | 0 … ∞
| 2 | Lens | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Lens | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Lens | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Lens | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Lens | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Lens | ID | 2 |  |  | R | LensID
| 2 | Lens | Name | 2 |  | A user defined name for this Lens element. | R | string
| 5 | Lens | Type | 5 | LensType | The general category this lens belongs to. | R | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 5 | Lens | Geometry | 5 | LensGeometry | The geometrical shape of the lens surfaces |  | enum | Spherical, Aspherical, Cylindrical, Ball, Hemispheric-ball, Cone, Drum, Rod, Toric
| 5 | Lens | Construction | 5 | LensConstruction | Whether or not the lens is simple or composite | R | enum | Simple, Composite, Other
| 2 | Lens | Magnification | 2 |  | The nominal magnification provided by this lens | R | integer with unit:none
| 5 | Lens | AttenuationCoefficient | 5 |  | The attenuation coefficient (aka blocking coefficient) is a material property that characterizes the degree of transparency of a lens. The attenuation coefficient is often defined in units of optical density (OD), wherease OD= -log10(T), and T=Transmittance (0≤T≤1). | R | float with unit:none
| 5 | Lens | RadiusOfCurvature | 5 |  | A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusUnit
| 5 | Lens | FocalDistanceToObject | 5 |  | The distance between the front principal plane of the lens and the specimen | R | float with unit:FocalDistanceToObjectUnit
| 5 | Lens | FocalDistanceToImage | 5 |  | The distance from the rear principal plane to the image | R | float with unit:FocalDistanceToImageUnit
| 5 | Lens | FocalLenght | 5 |  | The distance from the rear principal plane of the lens and its focal point | R | float with unit:FocalLengthUnit
| 5 | Lens | MaterialName | 5 | LensMeterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | alpha-Barium Borate, N-BK7 (Borosilicate Crown), Calcium Fluoride , UV-Grade Fused Silica (UVFS), Germanium, Magnesium Fluoride, N-SF11 (Dense Flint), Polytetrafluoroethylene (PFTE), Silicon (Si), Zinc Selenide (ZnSe), Other
| 5 | Lens | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | Lens | RefractiveIndex | 5 |  | The refractive index of the material the prism is made of | R | float with unit:none
| 5 | Lens | AbbeNumber | 5 |  | The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion | R | float with unit:none
| 5 | Lens | Density | 5 |  | The density of the material this lens is made of |  | float with unit:DensityUnit
| 5 | Lens | Diameter | 5 |  | The physical diamter of this lens |  | float with unit:none
| 2 | Lens | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 5 | CouplingLens |  | 5 |  | A lens positioned at connector between this light source coupling element and the microscope body.
| 3 | RelayLens |  | 3 |  | A lens positioned at the mount connecting this detector to the microscope body.
| 2 | AdditionalLens |  | 2 |  | A lens poistioned in the light path in the light path for a given channel.
| 2 | Prism |  | 2 |  | A prism is a block of glass or other transparent material with flat, polished surfaces. A prism can be employed to deflect or deviate a light beam, rotate or invert an image, separate polarization states, or disperse light into its component wavelengths. Prisms can be classified in three general categories: reflecting prisms, polarizing prisms, and refracting or dispersion prisms. Reflecting prisms are useful for redirecting light beams by total internal reflection. Refracting prisms can be employed to bend and separate light into its component colors. Polarizing prisms are birefringent crystals that divide incident non-polarized light into separate components polarized orthogonally to each other. These prisms are employed to produce polarized light for optical instruments such as microscopes and polarimeters. A prism can be viewed as an ensamble of different rigidly mounted mirrors, with each face having a permanent orientation with respect to all the others. | 0 … ∞
| 2 | Prism | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Prism | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Prism | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Prism | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Prism | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Prism | ID | 2 |  |  | R | PrismID
| 2 | Prism | Name | 2 |  | A user defined name for this Prism element. | R | string
| 3 | Prism | Type | 3 | PrismType | The category this prism belongs to. | R | enum | Reflecting, Polarizing, Refracting, Other
| 5 | Prism | Geometry | 5 | PrismGeometry | The geometrical shape of this prism |  | enum | Right-Angle, Porro, Dove, Amici, Penta, Romboid, Nicol, Glan-Thompson, Wollaston, Rochon, Senarmont, Triangular, Abbe prism, Pellin–Broca, Grim, Compound, Other
| 5 | Prism | Reflectance | 5 |  | The amount of light the prism reflects at a maximum. It is represented as a fraction, as a value from 0.0 to 1.0. | R | float with unit:none
| 5 | Prism | PrismAngle | 5 |  | The angle between two surfaces of a prism |  | float with unit:PrismAngleUnit
| 5 | Prism | AngleOfIncidence | 5 |  | The angle of incidence (AOI) is the angle between a beam of light incident on the prism surface and the line perpendicular to the surface at the point of incidence, called the normal. | R | float with unit:IncidenceAngleUnit
| 5 | Prism | DeviationAngle | 5 |  | A beam passing through an object like a prism is deflected twice: once entering, and again when exiting. The sum of these two deflections is called the deviation angle and is measured with respect to the incident ray | R | float with unit:DeviationUnitAngle
| 5 | Prism | MaterialName | 5 | PrismMaterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | Calcium Fluoride, F2 (Flint), N-BK7 (Borosilicate Crown), N-SF11 (Dense Flint), Rutile (TiO2), UV-Grade Fused Silica (UVFS), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | Prism | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | Prism | RefractiveIndex | 5 |  | The refractive index of the material the prism is made of | R | float with unit:none
| 5 | Prism | AbbeNumber | 5 |  | The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion | R | float with unit:none
| 5 | Prism | Density | 5 |  | The density of the material the prism is made of |  | float with unit:DensityUnit
| 5 | Prism | WavelengthRange | 5 |  | This records a range of wavelengths associated with the associated microscope component. Multiple of this elements may be used. | 1 | Element
| 2 | Prism | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | PolarizationOptics |  | 2 |  | This is an optical filter that lets light waves of a specific polarization pass through while blocking light waves of other polarizations | 0 … ∞
| 2 | PolarizationOptics | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | PolarizationOptics | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | PolarizationOptics | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | PolarizationOptics | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | PolarizationOptics | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | PolarizationOptics | ID | 2 |  |  | R | PolarizationOpticsID
| 2 | PolarizationOptics | Name | 2 |  | A user defined name for this Polarization Optics element. | R | string
| 2 | PolarizationOptics | Function | 2 | PolarizationOpticsFunction | This records whether this polarizer optics is used to generate polarized light or to analyze polarized light | R | enum | Polarizer, Analyzer
| 3 | PolarizationOptics | Construction | 3 | PolarizationOpticsConstruction | The modality in with which this Polarization Optics is constructed |  | enum | Achromatic, Zero Order, Multiple Order, Other
| 5 | PolarizationOptics | Fabrication | 5 | PolarizationOpticsFabrication | The way in with which this Polarization Optics is fabricated |  | enum | Wire Grid, Film, Polarizing Beamsplitter, Polarizing BandPass, Variable, Birefringent, Polymer, Quarz, Other
| 3 | PolarizationOptics | Retardation | 3 | PolarizationOpticsRetardation | Retardation describes the phase shift between the polarization component projected along the fast axis and the component projected along the slow axis |  | enum | Half Wave, Quarter Wave, Full Wave, Other
| 5 | PolarizationOptics | MaterialName | 5 | PolarizationOpticsMaterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | alpha-Barium Borate, Calcite, Calcium Fluoride, N-BK7 (Borosilicate Crown), Rutile (TiO2), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | PolarizationOptics | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | PolarizationOptics | WavelengthRange | 5 |  | This records a range of wavelengths associated with the associated microscope component. Multiple of this elements may be used. | 1 | Element
| 2 | PolarizationOptics | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | FilterSet |  | 2 |  | A filter set is typically composed of one excitation filter, one emission filter and a dichroic. Often they are pre-assebled by the manufacturer and pruchased as a whole | 0 … ∞
| 2 | FilterSet | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | FilterSet | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | FilterSet | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | FilterSet | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | FilterSet | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | FilterSet | ID | 2 |  |  | R | FilterSetID
| 2 | FilterSet | Name | 2 |  | A user defined name for this FilterSet. | R | string
| 3 | FilterSet | ExcitationFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Excitation light path | 1 | Extension of Reference
| 3 | FilterSet | DichroicRef | 3 | Reference | This empty element has a reference (the Mirror ID attribute) to the Mirror/Dichroic placed in the light path | 1 | Extension of Reference
| 3 | FilterSet | EmissionFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Emission light path | 1 | Extension of Reference
| 2 | FilterSet | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 2 | Filter-APE |  | 2 |  | An optical filter is a device that selectively transmits light of different wavelengths. A filter maybe either an excitation, an emission filter or an additional filter positioned in the light path. For each fluorescence Image there should be at least one filter element specified per Channel. This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet. | 1 … ∞
| 2 | Filter-APE | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Filter-APE | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Filter-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Filter-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Filter-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Filter-APE | ID | 2 |  |  | R | FilterID
| 2 | Filter-APE | Name | 2 |  | A user defined name for this Filter. | R | string
| 3 | Filter-APE | Type | 3 | FilterType | The filter type. | R | enum | Bandpass, Multipass, Longpass, Shortpass, NeutralDensity, HeatFilter, Tunable, Other
| 5 | Filter-APE | ReflectionCoating | 5 |  | An optical coating is one or more thin layers of material deposited on an optical component such as a lens or mirror, which alters the way in which the optic reflects and transmits light. | R | string
| 4 | Filter-APE | FilterWheel | 4 |  | A filter 'wheel' in OME can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. |  | string
| 5 | Filter-APE | TransmittanceProfileFie | 5 |  | This contains the location of file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 2 | Filter-APE | TransmittanceRange | 3 |  | This records the range of wavelengths that are transmitted by the filter. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | 1 … ∞ | Element
| 2 | Filter-APE | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 2 | TransmittanceRange | CutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 2 | TransmittanceRange | CutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 3 | TransmittanceRange | Transmittance | 4 |  | The amount of light the filter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. | R | percentFraction with unit:none
| 2 | ExcitationFilter |  | 2 |  | An optical filter is a device that selectively transmits light of different wavelengths.This element refers to an excitation filter. This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | EmissionFilter |  | 2 |  | An optical filter is a device that selectively transmits light of different wavelengths.This element refers to an emission filter. For each fluorescence Image there should be at least one emission filter element specified per Channel.This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | AdditionalFilter |  | 2 |  | An optical filter is a device that selectively transmits light of different wavelengths.This element refers to a filter positioned along the light path in addition to the excitation and the emission filters.This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | Mirror-APE |  | 2 | Substitution : choose one child from the list below | This element describes mirror component(s) available in this microscope. | 1 … ∞ |  | Dichroic, Beamsplitter, Reflector
| 2 | Mirror-APE | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | Mirror-APE | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | Mirror-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Mirror-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Mirror-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Mirror-APE | ID | 2 |  |  | R | MirrorID
| 2 | Mirror-APE | Name | 2 |  | A user defined name for this Mirror. | R | string
| 4 | Mirror-APE | Reflectance | 4 |  | The amount of light the mirror reflects at a maximum. It is represented as a fraction, as a value from 0.0 to 1.0. | R | float with unit:none
| 5 | Mirror-APE | MirrorType | 5 | MirrorType | The general category this lens belongs to |  | enum | Convex, Concave, Planar, Other
| 5 | Mirror-APE | Geometry | 5 | MirrorGeometry | The geometrical shape of the surface of this mirror |  | enum | Spherical, Aspherical, Cylindrical, Ellipsoidal, Hyeprboloidal, Paraboloidal, Other
| 5 | Mirror-APE | AngleOfIncidence | 5 |  | This field records the optimal angle of incidence (AOI) for this mirror. TheAOI is the angle between the optical axis of the incident light and the axis normal to the surface of the mirror. In specular reflection AOI is by definition equal to the Angle Of Reflection (AOR). | R | float with unit:AngleOfIncidenceUnit
| 5 | Mirror-APE | RadiusOfCurvature | 5 |  | A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusOfCurvatureUnit
| 5 | Mirror-APE | SubstrateType | 5 | SubstrateType | This field records the type of substrate that is used for the beamsplitter coating. It is typically a ground and polished piece of optical glass. |  | enum | Metallic, Dielectric, Other
| 5 | Mirror-APE | SubstrateMaterial | 5 | SubstrateMaterialName | This records the chemical composition or the nonproprietary trade designation of the mirror's optical substrate. For dichoric mirrors the substrate is typically composed of a glass material with low autofluorescence, such as UV-grade fused silica. |  | enum | Calcium Fluoride, N-BK7 (Borosilicate Crown), UV-Grade Fused Silica (UVFS), Zerodur (Lithium Aluminum Ailicon Oxide), Zinc Selenide (ZnSe), Other
| 5 | Mirror-APE | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 2 | Mirror-APE | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 2 | Dichroic |  | 2 |  | A dichroic beamsplitter or dichroic mirror is a mirror that spectrally separate light by transmitting and reflecting light as a function of wavelength. In the context of fluorescence microscopy, they are typically used to reflect the excitation light, but transmit the emitted fluorescence. Dichroic mirrors can be categorized in different manners. However, the most important criteria for choosing a dichoric mirror are the transmittance and the reflectance wavelenght ranges.
| 3 | Dichroic | DichroicType | 3 |  | The type of dichroic mirror used in this case | R | enum | Bandpass, Multipass, Longpass, Shortpass, Other
| 3 | Dichroic | HardSputtered | 3 |  | This field records whether this coating of this dichroic is hard-sputtered or traditional. | R | boolean
| 3 | Dichroic | AttenuationLevel | 3 |  | The attenuation level, aka blocking level, is used as a measure of the out-of-band attenuation of an optical filter or dichroic mirror over an extended range of the spectrum. The attenuation level is often defined in units of optical density (OD), wherease OD= -log10(T), and T=Transmittance (0≤T≤1). | R | float with unit:none
| 5 | Dichroic | AttenuationMethod | 5 |  | The mean by which this dichroic filter achieves out-of-band transmission attenuation. |  | enum | Absorption, Reflection, Rifraction, Other
| 5 | Dichroic | Technology | 5 |  | The main types of filter technology used in fluorescence analysis are absorption (aka colored) glass, thin-film coating, and acousto-optical tunable, liquid-crystal tunable filters. |  | enum | Absorption Glass, Metallic-coating, Thin-film interference-coating, Acousto-optical, Liquid-crystal, Other
| 5 | Dichroic | Polarization | 5 |  | Dichroic beamsplitters (and, in fact, any thin-film interference coating that is used at non-normal angles-of-incidence) will cause some amount of polarization. ), the precise effect varying greatly with wavelength and with the particular coating design. This field records the polarization effect of this dichroic. |  | enum | Planar, Circular, Elliptical, Other
| 4 | Dichroic | FilterWheel | 4 |  | A filter 'wheel' in OME can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. [plain text string] |  | string
| 5 | Dichroic | TransmittanceProfileFile | 5 |  | This contains the location of file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 5 | Dichroic | ReflectanceProfileFile | 5 | WavelengthBounds | This contains the location of file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 2 | Dichroic | TransmittanceRange | 3 |  | This records the range of wavelengths that are transmitted by the dichroic mirror. It also records the maximum amount of light transmitted | 1 … ∞ | Element
| 3 | Dichroic | ReflectanceRange | 3 |  | This records the range of wavelengths that are reflected by the mirror. | 0 … ∞ | Element
| 2 | TransmittanceRange | CutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 2 | TransmittanceRange | CutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 3 | TransmittanceRange | Transmittance | 4 |  | The amount of light the filter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. | R | percentFraction with unit:none
| 3 | ReflectanceRange | CutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% | R | positivefloat with unit:CutInUnit
| 3 | ReflectanceRange | CutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% | R | positivefloat with unit:CutOutUnit
| 2 | Beamsplitter |  | 2 |  | This is a mirror that reflects part of the light, regardless of wavelenght
| 4 | Beamsplitter | Transmittance | 4 |  | The amount of light the beamsplitter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. | R | float with unit:none
| 5 | Beamsplitter | TransmissionAngle | 5 |  | The angle at which light is transmitted by this beamsplitter | R | float with unit:none
| 2 | Reflector |  | 2 |  | This is a total reflecion mirror
| 2 | Detector-APE |  | 2 | Substitution : choose one child from the list below | This element describe the detector(s) available in this microscope. The Detector ID can be used as a reference within the Channel element in the Image element. The values stored in Detector represent the fixed values, variable values modified during the acquisition go in DetectorSettings. | 1 … ∞ |  | Camera, Photomultiplier
| 2 | Detector-APE | Manufacturer | 2 | Manufacturer Spec | The manufacturer of the component. | R | string
| 2 | Detector-APE | Model | 2 | Manufacturer Spec | The model of the component. | R | string
| 5 | Detector-APE | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | Detector-APE | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | Detector-APE | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Detector-APE | ID | 2 |  |  | R
| 2 | Detector-APE | Name | 2 |  | A user defined name for this Detector. | R | string
| 3 | Detector-APE | Gain | 3 |  | The analog gain is the conversion unit for this detector. It specifies how many electrons correspond to 1 signal unit. It is expressed in e-/Analog-to-Digit-Unit (ADU). | R | float with unit:none
| 5 | Detector-APE | Mount | 5 |  | The type of mount used for the camera. | R | enum | C-mount, T-mount, SLR adapter, D-SLR adapter, Other
| 5 | Detector-APE | ReadNoise | 5 |  | Read noise includes thermal noise build-up and clock induced charge. This is generally expressed in e- rms. | R | float with unit:ReadNoiseUnit
| 5 | Detector-APE | DetectorNoiseModel | 5 | DetectorNoiseModel | The noise contribution that comes from the detector | R | enum | Poisson, Gaussian, Other
| 5 | Detector-APE | DarkCurrentRate | 5 |  | Median current detected in the absence of light. Generally expressed in e-/pixel/sec | R | float with unit:DarkCurrentRateUnit
| 5 | Detector-APE | PeakWavelength | 5 |  | The wavelength associated with maximum light intensity detected by this device. | R | float with unit:PeakWavelengthUnit
| 5 | Detector-APE | QuantumEfficiencyatPeakWavelength | 5 |  | The Quantum Efficiency observed at peak wavelength. Quantum Efficiency is generally defined as is the ratio of incident photon to converted electrons. Typically expressed as a number between 0 and 1. | R | float with unit:none
| 5 | Detector-APE | AmbientOperatingTemperature | 5 |  | The ambient temperature at which this detector is designed to operate. |  | float with unit:TemperatureUnit; default: Celsius[ºC]
| 5 | Detector-APE | AbientOperatingHumidity | 5 |  | The ambient humidity at which this detector is designed to operate. |  | float with unit:none
| 5 | Detector-APE | WavelengthRange | 5 |  | This records a range of wavelengths associated with the associated microscope component. Multiple of this elements may be used. | 0 … 1 | Element
| 3 | Detector-APE | RelayLens | 3 |  | This is a reference to a lens positioned at the mount connecting this detector to the microscope body. | 0 … 1 | Extension of Reference
| 2 | Detector-APE | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelenghtProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | Camera-APE |  | 2 | Substitution : choose one child from the list below | This element describes the camera(s) devices available in this microscope. A camera is device that converts optical brightness into an electrical signal using a CCD(a semiconductor chip, which converts the intensity of light falling onto its surface at different points into a charge) and then reproduces an image of the specimen. |  |  | CCD, Intensified CCD, AnalogVideo, CMOS
| 3 | Camera-APE | Illumination | 3 | CameraIllumination | The type of illumination the camera chip receives | R | enum | Front, Back, Other
| 3 | Camera-APE | ArrayWidth | 3 |  | The width of this camera chip expressed in pixel. | R | integer with unit:pixel
| 3 | Camera-APE | ArrayHeight | 3 |  | The height of this camera chip expressed in pixel. | R | integer with unit:pixel
| 3 | Camera-APE | PixelWidth | 3 |  | The physical width of each pixel | R | float with unit:PixelWidthUnit; default:µm
| 3 | Camera-APE | PixelHeight | 3 |  | The physical height of each pixel | R | float with unit:PixelHeightUnit; default:µm
| 3 | Camera-APE | ImageWidth | 3 |  | The width of the image generated by the camera |  | float with unit:ImageWidthUnit
| 3 | Camera-APE | ImageHeight | 3 |  | The width of the image generated by the camera |  | float with unit:ImageHeightUnit
| 5 | Camera-APE | PixelWellCapacity | 5 |  | The maximum amount of charge each pixel well can hold, measured in electron counts | R | integer with unit:PixelWellCapacity
| 5 | Camera-APE | RegisterWellCapacity | 5 |  | The maximum amount of charge each register well can hold, measured in electron counts |  | integer with unit:RegisterWellCapacity
| 5 | Camera-APE | VerticalClockSpeed | 5 |  | The maximal speed each vertical colunm of pixels in the camera chip, can be shifted down and read |  | float with unit:VerticalClockSpeedUnit
| 5 | Camera-APE | MaximumFrameRate | 5 |  | The maximal rate this camera can acquire a single frame | R | float with unit:MaximumFrameRateUnit [default fps]
| 5 | Camera-APE | MaximumReadoutRate | 5 |  | The maximal pixel read out rate allowed by this camera | R | float with unit:MaximalReadoutRateUnit[default pixel per second]
| 2 | CCD |  | 2 |  | A charged coupled device camera
| 3 | CCD | Type | 3 | CCDType | The type of CCD camera used. | R | enum | Conventional, EMCCD, Interline, Other
| 2 | IntensifiedCCD |  | 2 |  | An intensified charged coupled device camera
| 3 | IntensifiedCCD | Type | 3 | IntensifiedCCDType | The type of Intensified CCD camera used. | R | enum | Conventional, EBCCD, Other
| 2 | AnalogVideo |  | 2 |  | Andevice used to capture analog videos
| 2 | CMOS |  | 2 |  | A complementary metal–oxide–semiconductor camera
| 3 | CMOS | Type | 3 | CMOSType | The type of CMOS camera used. | R | enum | Conventional, Scientific, Other
| 2 | PhotoMultiplier |  | 2 |  | A photomultiplier tube, useful for light detection of very weak signals, is a photoemissive device in which the absorption of a photon results in the emission of an electron. These detectors work by amplifying the electrons generated by a photocathode exposed to a photon flux.
| 3 | PhotoMultiplier | Type | 3 | PhotoMultiplierType | The type of photomultiplier used. | R | enum | PMT, Photodiode, APD, Other
| 5 | PhotoMultiplier | RadiantSensitivity | 5 |  | The radiant sensitivity observed with this PhotoMultiplier at peak wavelength | R | float with unit:RadiantSensitivityAtPeakWavelengthUnit
| 3 | IntensityCalibrationTool |  | 3 |  | A tool, such as the Meta-Max calibration toold developed by the 4D Nucleome Imaging Standards Working Group, used to perform an intensity calibration procedure | 0 … ∞
| 3 | IntensityCalibrationTool | Manufacturer | 3 | Manufacturer Spec | The manufacturer of the component. | R | string
| 3 | IntensityCalibrationTool | Model | 3 | Manufacturer Spec | The model of the component. | R | string
| 5 | IntensityCalibrationTool | SerialNumber | 5 | Manufacturer Spec | The serial number of the component. | R | string
| 5 | IntensityCalibrationTool | LotNumber | 5 | Manufacturer Spec | The lot number of the component. |  | string
| 5 | IntensityCalibrationTool | SpecsFile | 5 | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | IntensityCalibrationTool | ID | 3 |  |  | R
| 3 | IntensityCalibrationTool | Name | 3 |  | A user defined name for this Intensity Calibration Tool | R | string
| 3 | IntensityCalibrationTool | Type | 3 |  | The type of calibration proceudure this tool is intended to perform | R | enum | Detector, Excitation, Other
| 3 | IntensityCalibrationTool | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 … 1 | Extension of Reference
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which 1): BOXES represent real-life entity-sets, about which information need to be collected. One way to think about it is that ENTITIES correspond to NOUNS. 2) LINES represent RELATIONSHIPS that connect entities with each other. Namely, RELATIONSHIPS correspond to VERBS (i.e. actions/states/occurrences). 3) 3) ATTRIBUTES corresponds to ADJECTIVES, which capture the actual information that needs to be collected about each entity.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on C18: We do not have any place in which the microscope acquisition software is described -Caterina Strambio De Castillia
Note on C43: APE = Abstract Parent Entity In the .xsd file, "Group" is used instead of " -APE"
Note on C44: Duplicate LightSource Transmitted vs Fluorescence
Note on D295: Change to CutOn CutOff

## 4DN proposal_Settings_v01-02
4DN-OME model
| Tier | Element | Attribute | Tier | Complex type | Description | Cardinality | Data type | Allowed values
IMAGE ACQUISITION Settings | 1 | Experiment |  | 1 |  | This element describes the type of experiment. The required Type attribute must contain one or more entries from the following list: FRET, Time-lapse, 4D+, HightThroughputScreen, FISH, Electrophysiology, Ion-Imaging, Co-localization, PGI/Documentation FRAP, Photo-ablation, Optical-Trapping, Photo-activation, Fluorescence-Lifetime, Spectral-Imaging, Other The required LabellingMethod attribute must contain one or more entries from the following list: FluorescentProtein, CASFISH, FISH, Single-Molecule-FISH, Indirect-Immunofluorescene, Immunocytochemistry, None, Other PGI/Documentation is not a 'data' image. The optional Description element may contain free text to further describe the experiment | 1
| 1 | Experiment | ID | 1
| 1 | Experiment | Name | 1 |  | A user assigned name for this Experiment
| 2 | Experiment | Type | 2 | ExperimentType | A term to describe the type of experiment. |  | enum | FRET, TimeLapse, 4D+, HightThroughputScreen, Electrophysiology, IonImaging, Co-localization, PGI-Documentation, FRAP, PhotoAblation, OpticalTrapping, PhotoActivation, FLIM, SpectralImaging, SPIM, Other
| 1 | Experiment | LabellingMethod | 1 | ExperimentLabellingMethod | The method used for labelling the molecular component under study in this experiment |  | enum | FluorescentProtein, CASFISH, FISH, Single-Molecule-FISH, Immunofluorescene, Immunocytochemistry, None, Other
| 1 | Experiment | ExperimenterRef | 1 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME. | 1 … ∞ | Extension of Reference
| 1 | Experiment/Description | Experiment/Description | 1 |  | A description for the Experiment | 0 … 1 | Element
| 2 | Experiment | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 1 | Sample |  | 1 |  | A description of the sample that was imaged in this experiment | 1
| 1 | Sample | ID | 1
| 1 | Sample | Name | 1 |  | A user assigned name for this Sample |  | string
| 2 | Sample | MountingMedium | 2 |  | The mounting medium the sample is embedded in during acquisition |  | string
| 4 | Sample | MountingMediumRefractiveIndex | 4 |  | The measured refractive index of the medium the sample is embedded in during imaging |  | float with unit:MountingMediumRefractiveIndexUnit
| 1 | Sample | Sample/Description | 1 |  | A description for the Sample | 0 ... 1 | Eleement
| 2 | Sample | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | MicrobeamManipulation |  | 3 |  | Defines a microbeam operation type and the region of the image it was applied to. The LightSourceSettingsRef element is a reference to a LightSourceSettings and in turn to the LightSource specified in the Instrument element which was used for this micromanipulation experiment. For example, a laser used for photobleaching. | 0 … ∞
| 3 | MicrobeamManipulation | ID | 3
| 3 | MicrobeamManipulation | Type | 3 | MicrobeamManipulationType | Defines the type of micromanipulation performed. |  | enum | Fluorescence Recovery After Photobleaching, inverse FRAP, Fluorescence Loss In Photobleaching, Photo-Ablation, Photo-Activation, Uncaging, OpticalTrapping, Other
| 3 | MicrobeamManipulation | ExperimenterRef | 3 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME | 0 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | ROIRef | 3 |  | This empty element has a reference (the ROIID attribute) to a four dimensional ROI ( 'Region of Interest') that was used in this micromanipulation experiment | 1 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | LightSourceSettings | 3 |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 1 … ∞ | Element
| 3 | MicrobeamManipulation | MicrobeamManipulation/Description | 3 |  | A description for the Microbeam Manipulation experiment | 0 … 1 | plain-text multi-line string
| 3 | MicrobeamManipulation | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 1 | Image |  | 1 |  | This element describes the actual image and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. The required attributes and elements are ID and Pixels. ExperimenterRef is required for all Images with well formed LSIDs. ImageType is a vendor-specific designation of the type of image this is. Examples of ImageType include 'STK', 'SoftWorx', etc. The Name attributes are in all cases the name of the element instance. In this case, the name of the image, not necessarily the filename. Physical size of pixels are microns[µm].
| 1 | Image |  | 1
| 1 | Image | Name | 1 |  | A user assigned name for this Image |  | string
| 1 | Image | AcquisitionDate | 1 |  | The acquisition date of the Image. |  | date
| 1 | Image (was Pixels) | SizeX | 1 |  | X size of pixel data array |  | integer with unit:none
| 1 | Image (was Pixels) | SizeY | 1 |  | Y size of pixel data array |  | integer with unit:none
| 1 | Image (was Pixels) | SizeZ | 1 |  | Number of Z-planes |  | integer with unit:none
| 1 | Image (was Pixels) | SizeC | 1 |  | Number of Channels used in this Image |  | integer with unit:none
| 4 | Image (was Pixels) | SizeT | 4 |  | Number of Timepoints used in this Image |  | integer with unit:none
| 1 | Image (was Pixels) | PhysicalSizeX | 1 |  | Physical width of a pixel |  | float with unitis:PhysicalSizeXUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeY | 1 |  | Physical height of a pixel |  | float with unitis:PhysicalSizeYUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeZ | 1 |  | Physical distance between optical planes |  | float with unitis:PhysicalSizeZUnit; default µm
| 4 | Image (was Pixels) | TimeIncrement | 4 |  | TimeIncrement is used for time series that have a global timing specification instead of per-timepoint timing info. For example in a video stream. Units are set by TimeIncrementUnit. |  | float with unit:TimeIncrementUnit
| 2 | Image (was Pixels) | DimensionOrder | 2 | PixelsDimensionOrder | The order in which the individual planes of data are interleaved in the Image |  | enum | XYZCT, XYZTC, XYCTZ, XYCZT, XYTCZ, XYTZC
| 2 | Image (was Pixels) | ChannelOrder | 2 |  | The order in which the channels were acquired and are stored |  | string
| 3 | Image (was Pixels) | PixelType | 3 | PixelType | The variable type used to represent the pixel intensity in the image |  | enum | int8, int16, int32, uint8, uint16, uint32, float, double, complex, double-complex, bit, Other
| 5 | Image (was Pixels) | SignificantBit | 5 |  | The number of bits within the type storing each pixel that are significant. e.g. you can store 12 bit data within a 16 bit type. This does not reduce the storage requirements but can be a useful indicator when processing or viewing the image data. |  | PositiveInteger
| 3 | Image (was Pixels) | Interleaved | 3 |  | How the channels are arranged within the data block |  | boolean | TRUE: if channels are stored RGBRGBRGB...; FALSE: if channels are stored RRR...GGG...BBB…
| 5 | Image (was Pixels) | BigEndian | 5 |  | This is true if the pixels data was written in BigEndian order. If this value is present it should match the value used in BinData or TiffData. If it does not a reader should honour the value used in the BinData or TiffData. This values is useful for MetadataOnly files and is to allow for future storage solutions. |  | boolean | TRUE: if pixel data was written in BigEndian mode
| 1 | Image | InstrumentRef | 1 |  | This empty element can be used (via the required Instrument ID attribute) to refer to an Instrument defined within OME | 1 | Extension of Reference
| 1 | Image | ExperimentRef | 1 |  | This empty element has a reference (the Experiment ID attribute) to a Experiment defined within OME. | 1 | Extension of Reference
| 1 | Image | ExperimenterRef | 1 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME. | 0 … ∞ | Extension of Reference
| 1 | Image | ExperimentGroupRef | 1 |  | This empty element has a reference (the ExperimenterGroup ID attribute) to a ExperimenterGroup defined within OME. | 0 … ∞ | Extension of Reference
| 3 | Image | MicroBeamManipulationRef | 3 |  | This empty element has a reference (the MicroBeamManipulationID attribute) to a Microbeam Manipulation defined into OME. E.g. a microbeam operation, such as a laser used for photobleaching. | 0 … ∞ | Extension of Reference
| 2 | Image | ROIRef | 2 |  | This empty element has a reference (the ROIID attribute) to a four dimensional ROI ( 'Region of Interest'). If they are not used, and the Image has more than one plane, the entire set of planes is assumed to be included in the ROI. Multiple ROIs may be specified. | 0 … ∞ | Extension of Reference
| 3 | Image | TIRFSettings | 3 | Extension of Settings | Defines the TIRF settings used for this experiment | 0 … ∞ | Element
| 4 | Image | ImagingEnvironment | 4 |  | This describes the environment that the biological sample was in, during the experiment. | 0 … ∞ | Element
| 1 | Image | MicroscopeSettings | 1 | Extension of Settings | This holds the setting applied to a microscope as well as a reference to the microscope. The ID is the microscope used in this case. | 1 … ∞ | Element
| 2 | Image | ObjectiveSettings | 2 |  | This holds the setting applied to an objective as well as a reference to the objective. The ID is the objective used in this case. | 1 … ∞ | Element
| 3 | Image | OpticalCalibration | 3 |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | 0 … ∞ | Element
| 3 | Image | DetectorCalibration | 3 |  | This procedure measures the performance of the dectector | 0 … ∞ | Element
| 1 | Image | Plane | 1 |  | The Plane object holds microscope stage and image timing data for a given channel/z-section/timepoint. | 1 … ∞ | Element
| 1 | Image | Channel | 1 |  | There must be one per Channel in the Image, even for a single-plane image. Information about how each of the Channel was acquired is stored in the various optional *Ref elements. One useful feature of the model, is that this design allows to create LogicalChannels which are composed of one or more ChannelComponents. For example, an entire spectrum in an FTIR experiment may be stored in a single Logical Channel with each discrete wavenumber of the spectrum constituting a ChannelComponent of the FTIR Logical Channel. An RGB image where the Red, Green and Blue components do not reflect discrete probes but are instead the output of a color camera would be treated similarly - one Logical channel with three ChannelComponents in this case. The total number of ChannelComponents for a set of pixels must equal SizeC. The IlluminationType attribute is a string enumeration which may be set to 'Transmitted', 'Epifluorescence', 'Oblique', or 'NonLinear'. The user interface logic for labeling a given channel for the user should use the first existing attribute in the following sequence: Name -> Fluor -> EmissionWavelength -> ChannelComponent/Index. | 1 … ∞ | Element
| 1 | Image | TiffData (or BidData or MetadataOnly) | 1 |  | This described the location of the pixel data in a tiff file (or The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream; or This place holder means there is on pixel data in this file) | 1 … ∞ | Element
| 1 | Image | Image/Description | 1 |  | A description for the image | 0 ... 1 | Element
| 2 | Image (was Pixels) | AnnotationRef | 2 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | TIRFSettings |  | 3 |  | Defines the TIRF settings used for this experiment | 0 … ∞
| 3 | TIRFSettings | ID | 3
| 3 | TIRFSettings | ThroughLens | 3 |  | The TIRF modality used in this case. True=Through the lens TIRF modality. False=TIRF modality in which the laser light is not travelling through the objective lens. |  | boolean
| 3 | TIRFSettings | TIRFAngle | 3 |  | The incidence angle used in this case |  | float with unit:TIRFAngleUnit
| 3 | TIRFSettings | Geometry | 3 |  | The TIRF geometry used in this case |  | enum | Spot, Ring, Other
| 3 | TIRFSettings/Description | TIRFSettings/Description | 3 |  | A description for the TIRF Settings | 0 ... 1 | plain-text multi-line string
| 3 | TIRFSettings | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 4 | ImagingEnvironment |  | 4 |  | This describes the environment that the biological sample was in during the experiment | 0 … ∞
| 4 | ImagingEnvironment | Temperature | 4 |  | The temperature the acquisition was performed at |  | float with unit:TemperatureUnit
| 4 | ImagingEnvironment | AirPressure | 4 |  | The airpressure the acquisition was performed at |  | float with unit:AirPressureUnit
| 4 | ImagingEnvironment | Humidity | 4 |  | The humidity the acqusition was performed at |  | float with unit:HumidityUnit
| 4 | ImagingEnvironment | CO2Percent | 4 |  | The CO2 concentration, expressed in percentage, the acquisition was performed at |  | float with unit:CO2PercentUnit
| 1 | MicroscopeSettings |  | 1 |  | This holds the setting applied to a Microscope as well as a reference to the Microscope (the ID attribute) in question. The ID referes to the microscope used in this case. | 0 … ∞
| 1 | MicroscopeSettings | ID (MicroscopeID) | 1
| 3 | MicroscopeSettings | FieldOfView | 3 |  | The Field of View defines the area of the specimen that is viewable via the eyepiece. This attribute caputers the measures size of the Field of View. |  | float with unit:FieldOfViewUnit[default mm]
| 1 | MicroscopeSettings | OptovarMagnification | 1 |  | This is the magnification associated with the Optovar used in this case |  | float with unit:none
| 1 | MicroscopeSettings | TotalEffectiveMagnification | 1 |  | This is the total effective magnification used in this case. This number can be obtained by multiplying the OptovarMagnification times the Objective/NominalMagnification or even better the Objective/CalibratedMagnification if available |  | float with unit:none
| 2 | ObjectiveSettings |  | 2 |  | This holds the setting applied to an Objective e as well as a reference to the Objective (the ID attribute) in question. The ID referes to the Objective used in this case. | 0 … ∞
| 2 | ObjectiveSettings | ID (ObjectiveID) | 2
| 5 | ObjectiveSettings | CorrectionCollar | 5 |  | The CorrectionCollar is normally an adjustable ring on the objective. Each has an arbitrary scale on it so the values is unit-less.[units:none] |  | float with unit:none
| 2 | ObjectiveSettings | ImmersiumMedium | 2 | ObjectiveSettingsImmersionMedium | The actual immersion medium that was used in this case |  | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 2 | ObjectiveSettings | ImmersionRefractiveIndex | 2 |  | The nominal refractive index of the immersion medium utilized in this case. This is a ratio so it also unit-less. [units:none] |  | float with unit:none
| 4 | ObjectiveSettings | MeasuredRefractiveIndex | 4 |  | The temperature dependenent measured RefractiveIndex of ther immersion medium utilized in this case. This is a ratio so it also unit-less. [units:none] |  | float with unit:none
| 4 | ObjectiveSettings | Temperature | 4 |  | The temperature of the objective if different from ImagingEnvironment/Temperature |  | float with unit:ObjectiveTemperatureUnit
| 1 | Plane |  | 1 |  | The Plane object holds microscope stage and image timing data for a given channel/z-section/timepoint. | 1 … ∞
| 1 | Plane | ID | 1
| 1 | Plane | TheZ | 1 |  | The Z-section to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 1 | Plane | TheT | 1 |  | The timepoint to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 1 | Plane | TheC | 1 |  | The channel to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 2 | Plane | Timestamp | 2 |  | This records the total time intercoursed since the beginning of the experiment |  | float with unit:TimestampUnit
| 2 | Plane | ExposureTime | 2 |  | The length of the exposure. |  | float with unit:ExposureTimeUnit.
| 4 | Plane | SamplePosition | 4 |  | This is used to specify a name and position for a sample position in the microscope's reference frame. | 0 … 1 | Element
| 3 | Plane | FocalPosition | 3 |  | This element is used to record the Z-axis position of a Plane with respect to the microscope's reference frame | 1 | Element
| 3 | Plane | MechanicalCalibration | 3 |  | This describes the procedures that were used for mechanical calibration and the resulting calibration results. | 0 … ∞ | Element
| 5 | Plane | HashSHA1 | 5 |  | This optional element is a hash of the plane's image data. It is a choice between all the supported hash types. Currently the only method supported is SHA1 | 0 ... 1 | Element
| 3 | Plane | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | FocalPosition |  | 3 |  | This element is used to record the Z-axis position of a Plane with respect to the microscope's reference frame | 0 ... 1
| 3 | FocalPosition | ID | 3
| 3 | FocalPosition | Name | 3 |  | A user assigned name for this Z-axis focal position.
| 3 | FocalPosition | StagePositionZ | 3 |  | The Z-axis focal position.
| 3 | FocalPosition | FocusingRef | 3 |  | This empty element has a reference (the FocusingID attribute) to a Focusing element defined within OME. | 0 ... 1 | Extension of Reference
| 3 | FocalPosition | AutoFocusRef | 3 |  | This empty element has a reference (the AutofocusID attribute) to a Autofocus element defined within OME. | 0 ... 1 | Extension of Reference
| 4 | SamplePosition |  | 4 |  | SamplePosition is used to specify a name and location for a specific sample holder position in the microscope's reference frame. This element subsitutes a deprecated element called StageLabel | 0 … ∞
| 4 | SamplePosition | ID | 4
| 4 | SamplePosition | Name | 4 |  | A user assigned label for a specific sample holder position
| 4 | SamplePosition | SamplePositionX | 4 |  | The X position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionXUint
| 4 | SamplePosition | SamplePositionY | 4 |  | The Y position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionYUint
| 4 | SamplePosition | SamplePositionZ | 4 |  | The Z position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionZUint
| 4 | SamplePosition | SampleHolderRef | 4 |  | This empty element has a reference (the SampleHolderID attribute) to a SampleHolder element defined within OME. | 0 … 1 | Extension of Reference
| 1 | Channel |  | 1 |  | There must be one per Channel in the Image, even for a single-plane image. Information about how each of the Channel was acquired is stored in the various optional *Ref elements. One useful feature of the model, is that this design allows to create LogicalChannels which are composed of one or more ChannelComponents. For example, an entire spectrum in an FTIR experiment may be stored in a single Logical Channel with each discrete wavenumber of the spectrum constituting a ChannelComponent of the FTIR Logical Channel. An RGB image where the Red, Green and Blue components do not reflect discrete probes but are instead the output of a color camera would be treated similarly - one Logical channel with three ChannelComponents in this case. The total number of ChannelComponents for a set of pixels must equal SizeC. The IlluminationType attribute is a string enumeration which may be set to 'Transmitted', 'Epifluorescence', 'Oblique', or 'NonLinear'. The user interface logic for labeling a given channel for the user should use the first existing attribute in the following sequence: Name -> Fluor -> EmissionWavelength -> ChannelComponent/Index. | 1 … ∞
| 1 | Channel | ID | 1
| 1 | Channel | Name | 1 |  | A name for the channel that is suitable for presentation to the user. |  | string
| 5 | Channel | SamplesPerPixel | 5 |  | The number of samples the detector takes to form each pixel value. [units:none] Note: This is not the same as "Frame Averaging" - see Integration in DetectorSettings. |  | PositiveInteger
| 2 | Channel | IlluminationType | 2 | ChannelIlluminationType | The method of illumination used to capture the channel. |  | enum | Transmitted, Epifluorescence, Confocal, LightSheet, TIRF, NonLinear, Oblique, Other
| 3 | Channel | PinholeSize | 3 |  | The Pinhole Size utilized for this channel. |  | float with unit:PinholeSizeUnit.
| 1 | Channel | AcquisitionMode | 1 | ChannelAcquistionMode | AcquisitionMode describes the type of microscopy performed for each channel. |  | enum | WideField, LaserScanningConfocalMicroscopy, SpinningDiskConfocal, SlitScanConfocal, MultiPhotonMicroscopy, StructuredIllumination, SingleMoleculeImaging, TotalInternalReflection, FluorescenceLifetime, SpectralImaging, FluorescenceCorrelationSpectroscopy, NearFieldScanningOpticalMicroscopy, SecondHarmonicGenerationImaging, PALM, STORM, STED, TIRF, FSM, LCM, Other, BrightField, SweptFieldConfocal, SPIM
| 1 | Channel | ContrastMethod | 1 | ChannelContrastMethod | ContrastMethod describes the technique used to achieve contrast for each channel. |  | enum | Brightfield, Phase, DIC, HoffmanModulation, ObliqueIllumination, PolarizedLight, Darkfield, Fluorescence, Other
| 3 | Channel | ExcitationWavelength | 3 |  | The excitation wavelength used for this channel. |  | float with unit:ExcitationWavelengthUnit
| 3 | Channel | EmissionWavelength | 3 |  | The emission wavelength used for this channel. |  | float with unit:EmissionWavelengthUnit
| 2 | Channel | Fluorophore | 2 |  | This is the name of the fluorophore used to produce this channel [plain text string]. |  | string
| 5 | Channel | NDFilter | 5 |  | The NDfilter attribute is used to specify the combined effect of any neutral density filters used. The amount of light the filter transmits at a maximum expressed as a fraction with value from 0.0 to 1.0. NOTE: This was formerly described as "units optical density expressed as a PercentFraction". This was how the field had been described in the schema from the beginning but all the use of it has been in the opposite direction, i.e. as a amount transmitted, not the amount blocked. This change has been made to make the model reflect this usage. |  | float with unit:none
| 5 | Channel | PockelCellSetting | 5 |  | The PockelCellSetting used for this channel. This is the amount the polarization of the beam is rotated by. |  | integer with unit:none
| 1 | Channel | Color | 1 |  | A color used to render this channel - endocded as RGBA. The default value is "-1" corresponding #FFFFFFFF, so solid white. |  | string
| 3 | Channel | LightSourceSettings | 3 |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 1 | Element
| 2 | Channel | LightPath | 2 |  | A description of the light path | 1 | Element
| 3 | Channel | DetectorSettings-APE | 3 |  | This element holds the setting applied to a Detector e as well as a reference to the Detector (the ID attribute) in question. The ID referes to the Detector used in this case. | 1 | Element
| 3 | Channel | OpticalCalibration | 3 |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | 0 … ∞ | Element
| 3 | Channel | FieldCalibration | 3 |  | This describes the procedure used to measure the field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 0 … ∞ | Element
| 3 | Channel | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation.. | 0 ... 1 | Extension of Reference
| 3 | LightSourceSettings |  | 3 |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 1
| 3 | LightSourceSettings | ID (LightSourceID) | 3
| 3 | LightSourceSettings | Attenuation | 3 |  | The Attenuation of the light source. A fraction, as a value from 0.0 to 1.0 |  | float with unit:none
| 3 | LightSourceSettings | Wavelength | 3 |  | The Wavelength of the light source |  | float with unit:WavelengthUnit
| 3 | LightSourceSettings | OutputPower | 3 |  | Light power measured in the back of the objective |  | float with unit:OutputPowerUnit
| 3 | LightSourceSettings | PowerAtSample | 3 |  | Light power measured at the sample interface |  | float with unit:PowerAtSampleUnit
| 2 | LightPath |  | 2 |  | A description of the light path | 1
| 2 | LightPath | ID | 2 |  |  |  | LSID
| 2 | LightPath | Name | 2 |  | Any user defined name/description for this element |  | string
| 3 | LightPath | LightSourceCouplingRef | 3 |  | This empty element can be used (via the required LightSourceCouplingID attribute) to refer to a LightSourceCoupling element defined within OME | 0 … ∞ | Extension of Reference
| 2 | LightPath | FilterSetRef | 2 | Reference | This empty element has a reference (the FilterSet ID attribute) to the FilterSet placed in the light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | ExcitationFilterRef | 2 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Excitation light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | DichroicRef | 2 | Reference | This empty element has a reference (the Mirror ID attribute) to the Mirror/Dichroic placed in the light path | 1 … ∞ | Extension of Reference
| 2 | LightPath | EmissionFilterRef | 2 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Emission light path | 1 … ∞ | Extension of Reference
| 2 | LightPath | AdditionalOptics | 2 |  | A description of additional optics placed in the light path | 0 … ∞ | Element
| 3 | LightPath | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 2 | AdditionalOptics |  | 2 |  | A description of additional optics placed in the light path | 0 … ∞
| 2 | AdditionalOptics | ID | 2
| 2 | AdditionalOptics | Name | 2 |  | Any user defined name/description for this element |  | string
| 3 | AdditionalOptics | AdditionalLensRef | 3 | Reference | This empty element has a reference (the LensID attribute) to the an additional Lens placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | PrimsRef | 3 | Reference | This empty element has a reference (the PrismID attribute) to the a Prism placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | PolarizationOpticsRef | 3 | Reference | This empty element has a reference (the PorarizationOpticsID attribute) to the a PolarizationOptics placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AdditionalFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the an additional Filter placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AdditionalMirrorRef | 3 | Reference | This empty element has a reference (the MirrorID attribute) to the an additional Mirror placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | DetectorSettings-APE |  | 3 |  | This element holds the setting applied to a Detector e as well as a reference to the Detector (the ID attribute) in question. The ID referes to the Detector used in this case. | 1
| 3 | DetectorSettings-APE | ID (DetetectorID) | 3 |  |  |  | DetectorID
| 3 | DetectorSettings-APE | EffectiveOffset | 3 |  | This field reports the actual offset utilized for the acquisition of this image. The offset is an electronic baseline added to the output signal from the sensor to ensure that the displayed signal level is always a positive number of counts. |  | float with unit:none
| 3 | DetectorSettings-APE | Integration | 3 |  | This is the number of sequential frames that have been averaged for this image, to improve the signal-to-noise ratio |  | integer with unit:none
| 3 | DetectorSettings-APE | CameraFieldOfView | 3 |  | The Camera Field of View is the diameter of the field visible by the camera as measured in the specimen plane. It is expressed in mm and it is defined as the Cameara Field Number divided by the Total Effective Magnification of the system. In turn the Field Number is defined as the diameter of the view field in mm measured at the intermediate image plane. |  | float with unit:CameraFieldOfViewUnit[default mm]
| 3 | CameraSettings |  | 3 |  | This holds the setting applied to a Camera e as well as a reference to the Camera (the ID attribute) in question. The ID referes to the Camera used in this case.
| 3 | CameraSettings | Binning | 3 | Binning | Represents the number of pixels that are combined to improve readout rates, signal to noise ratios albeit at the expense of reduced spatial resolution |  | enum | 1x1, 2x2, 4x4, 8x8, Other
| 3 | CameraSettings | AnalogToDigitalConverter | 3 | PixelType | The bit depth used to report the light signal in digital units. |  | enum | int8, int16, int32, uint8, unit16, unit32, float, double, complex, double-complex, bit
| 3 | CameraSettings | EMGain | 3 |  | This is the electron multiplication gain (as apposed to the inherent gain) that was applied to the camera for the acquisition of this image. |  | float with unit:none
| 5 | CameraSettings | ReadoutRate | 5 |  | Effective pixel read out rate used for the acquisition of this image. This is the bytes per second that can be read from the detector (like a baud rate). |  | float with unit:ReadOutRateUnit
| 3 | CameraSettings | FrameRate | 3 |  | Effective frame rate used for the acquisition of this image. |  | float with unit:FrameRateUnit
| 3 | CameraSettings | OperatingTemperature | 3 |  | Effective cooling temperature used duringh the acquisition of this image. |  | float with unit:TemperatureUnit; default:Celsius[°C]
| 3 | PhotoMultiplierSettings |  | 3 |  | This holds the setting applied to a PhotoMultiplier e as well as a reference to the PhotoMultiplier (the ID attribute) in question. The ID referes to the PhotoMultiplier used in this case.
| 3 | PhotoMultiplierSettings | PMTGain | 3 |  | In general, the gain is a measure of the ability a two-port circuit to add energy to an input signal and therefore increase its power or amplitude. Specifically, PMTGain refers to the ratio of output to input current set for the acquisition of this image |  | float with unit:none
| 3 | PhotoMultiplierSettings | PMTVoltage | 3 |  | This refers to the voltage applied to the photomultiplier for the acquisition of this image |  | float with unit:PMTVoltageUnit
| 3 | PhotoMultiplierSettings | EffectiveZoom | 3 |  | The Zoom factor, aka "Confocal Zoom" or "Scan Zoom", utilized for the acquisition of this image. |  | foat with unit:none
| 3 | OpticalCalibration |  | 3 |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | 0 … ∞ |  |  | Alternative name: PSFCalibration
| 3 | OpticalCalibration | ID | 3 |  |  | R
| 3 | OpticalCalibration | Date | 3 |  | The date in which this calibration was performed. | R
| 3 | OpticalCalibration | MedianSpotSize_FWHM | 3 |  | Observed median lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MedianSpotSize
| 3 | OpticalCalibration | MedianTheoretical_FWHM | 3 |  | Theoretical (expected) medians lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MedianTheoretical
| 3 | OpticalCalibration | MinSpotSize_FWHM | 3 |  | Observed minimum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MinSpotSize
| 3 | OpticalCalibration | MinTheoretical_FWHM | 3 |  | Theoretical (expected) minimum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MinTheoretical
| 3 | OpticalCalibration | MaxSpotSize_FWHM | 3 |  | Observed maximum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MaxSpotSize
| 3 | OpticalCalibration | MaxTheoretical_FWHM | 3 |  | Theoretical (expected) maximum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MaxTheoretical
| 3 | OpticalCalibration | ZSpotSize_FWHM | 3 |  | Observed z-axial spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the z-plane. |  | float with unit:FWHM_ZSpotSize
| 3 | OpticalCalibration | ZTheoretical_FWHM | 3 |  | Theoretical (expected) z-axial spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the z-plane. |  | float with unit:FWHM_Ztheoretical
| 3 | OpticalCalibration | Planarity | 3 |  | This is a measure of the flatness of the focal plane across the Field of View (FOV). The planarity of the focal plane is calculated from the observed z-position of the best focus of each spot. It indicates focal aberrations or tilt. |  | float with unit:Planarity
| 3 | OpticalCalibration | Asymmetry | 3 |  | This field reports the mean +/- standard deviation over the Field of View (FOV) of the observed Point Spread Function (PSF) Asymmetry (i.e. the ratio of the minimum lateral FWMH over the maximum lateral FWHM). It indicates the presence of optical aberrations. |  | float with unit:none
| 3 | OpticalCalibration | LateralChromaticShift | 3 |  | The lateral offset of the observed spot position between two color channels. |  | float with unit:LateralChromaticShift
| 3 | OpticalCalibration | AxialChromaticShift | 3 |  | The axial offset of the observed spot position between two color channels. |  | float with unit:AxialChromaticShift
| 3 | OpticalCalibration | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | OpticalCalibrationImageRef |  | 3 |  | An image obtained by observing the optical standard used in this instance and acquired using the same instrument and settings employed to acquire the image this calibration refers to. | 1 … ∞ | Image
| 3 | OpticalCalibrationImageRef | ID (ImageID) | 3
| 3 | OpticalCalibrationReportRef |  | 3 |  | A report generated by the OpticalCalibrationSoftware referred to in OpticalCalibration. | 1 … ∞ | FileURI
| 3 | OpticalCalibrationReportRef | ID (OpticalCalibrationReportID) | 3
| 3 | OpticalCalibrationReportRef | OpticalCalibrationSoftware | 3 |  | The name of the software tool utilized to generate the OpticalCalibration report. As an example, such a tool might analyzes microscopy images of fluorescent beads in order to determine quantitatively the resolution and chromatic aberrations of a fluorescence microsocope over the full field of view. An example of such a tool is PSFj (http://www.knoplab.de/psfj/sample-page/).
| 3 | OpticalStandard-APE |  | 3 |  | The optical standard used in this case. This might be a pre-manufactured slide, such as those produced by Argolight (http://argolight.com/), Geller MicroAnalytical Laboratory (Magnification Reference Standard MRS-4 and others; http://www.gellermicro.com/mag_standards/mrs.html), Zeiss and others. Alternatively, the standard might be a slide containing multi-colored beads of known size and emission characteristics. Finally, slides coated with specific DNA origami patterns are increasingly being used for this purpose as well. | 1 … ∞
| 3 | OpticalStandard-APE | Manufacturer | 3 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | OpticalStandard-APE | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 3 | OpticalStandard-APE | SerialNumber | 3 | Manufacturer Spect | The serial number of the component |  | string
| 3 | OpticalStandard-APE | LotNumber | 3 | Manufacturer Spect | The lot number of the component |  | string
| 3 | OpticalStandard-APE | SpecsFile | 3 |  | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. |  | file
| 3 | OpticalStandard-APE | ID | 3
| 3 | OpticalStandard-APE | Name | 3 |  | A user assigned name for this optical standard. |  | Denomination
| 3 | OpticalStandard-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | ColoredBeadsSlide |  | 3 |  | This element describes a slide with colored beads that were used for this calibration procedure. |  | Element
| 3 | ColoredBeadsSlide | Name | 3 |  | A user assigned name for this beads preparation |  | Denomination
| 3 | ColoredBeadsSlide | BeadSize | 3 |  | The size of the colred beads utilized to make this slide. |  | float with unit:BeadSizeUnit
| 3 | ColoredBeadsSlide | BeadConcentration | 3 |  | The concentration of the colored beads preparation utilized to make this slide. |  | float with unit:BeadConcentrationUnit
| 3 | ColoredBeadsSlide | BeadVolume | 3 |  | The volume of the colored beads preparation utilized to prepare this slide |  | float with unit:BeadVolumeUnit
| 3 | ColoredBeadsSlide | BeadEmissionWavelenght | 3 |  | The peak wavelenght of the light emitted by these beads. In case of multi-color beads list different peak weavelength separated by commas. |  | string
| 3 | ColoredBeadsSlide | Type | 3 | BeadType | The type of colred beads used in this case |  | enum | Single color, Multi-color, Other
| 3 | CalibrationSlide |  | 3 |  | This element describes a slide used to measure the different optical properties such as magnification, distance, and different kinds of optical distorsions
| 3 | CalibrationSlide | Type | 3 | CalibrationSlideType | The type of calibration slide used in this case. Examples include slides produced by Argolight, Geller MicroAnalytical Laboratory, Zeiss and others. |  | enum | ARGO-HM, ARGO-LM, ARGO-SIM, Geller-MRS4, Zeiss universal D=0, Other
| 3 | CalibrationSlide | CalibratedMeasure | 3 | CalibratedMeasure | The optical characteristic this slide is used to calibrate, examples include magnification, distance, optical distorsions and others. |  | enum | Magnification, Distance, Multi, Distorsion, Other
| 3 | DNAOrigami |  | 3 |  | This element describes an optical standard constructed using the DNA origani technique.
| 3 | DNAOrigami | Type | 3 |  | This field describes the type of DNA origami-based optical standard used in this case. Currently this is fild is a simple string. However it might be substituted by an enumeration at a later stage. |  | string
| 3 | FieldCalibration |  | 3 |  | This describes the procedure used to measure the field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 0 … ∞
| 3 | FieldCalibration | ID | 3 |  |  | R
| 3 | FieldCalibration | Date | 3 |  | The date in which this calibration procedure was performed. | R
| 3 | FieldCalibration | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | FlatFieldImageRef |  | 3 |  | An image describing field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 1 … ∞ | Image
| 3 | FlatFieldImageRef | ID (ImageID) | 3
| 3 | IntensityCalibration-APE |  | 3 |  | This describes the procedures that were used for excitation and detector calibration and the resulting calibration results. | 0 … ∞ |  |  | This is for Field, Excitation and Detector calibration
| 3 | IntenstiyCalibration-APE | ID | 3 |  |  | R
| 3 | IntenstiyCalibration-APE | Date | 3 |  | The date in which this calibration procedure was performed. | R
| 3 | IntenstiyCalibration-APE | IntensityCalibrationToolRef | 3 |  | This element is a reference to the tool, if any, used to perform this calibratin procdure | 0 … ∞ | Extension of Reference
| 3 | IntenstiyCalibration-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
| 3 | ExcitationCalibration |  | 3 |  | This procedure measures the performance and allows calibration of the wavelength and irradiance of the excitation light.
| 3 | ExcitationCalibration | ObservedExcitationWavelength | 3 |  | Observed peak excitation wavelength. |  | float with unit:ObservedExcitationWavelengthUnit
| 3 | ExcitationCalibration | ObservedExcitationPower | 3 |  | Observed mean excitation light irradiance as measured at the sample plane over the Field of View (FOV). |  | float with unit:ObservedExcitationPowerUnit
| 3 | ExcitationCalibration | ObservedExcitationVariance | 3 |  | Observed irradiance variation over the Field of View (FOV). |  | float with unit:none
| 3 | DetectorCalibration |  | 3 |  | This procedure measures the performance of the dectector
| 3 | DetectorCalibration | Method | 3 |  | This is the method used for this calibration procedure |  | string
| 3 | CalibrationMap |  | 3 |  | This is a 2D matrix containing values computed as a resulut of the calibration procedure. This matrix is the same size of the image this calibration refers to. | 1 … ∞
| 3 | CalibrationMap | MetricType | 3 |  | This field captures the type of metric that was computed as part of the calibratin procedure in this instance. |  | enum | Gain, LinearOffset, DetectionLimit, LinearRangeMax, Variance, Mean, Other
| 3 | CalibrationMap | IlluminationIntensity | 3 |  | This attribute only applies to matrices of type Variance and Mean. It represents the illumination intensity that was used to acquire the images used to contruct this matrix. The value can range from 0 (i.e. dark image) to any intenisty value experssed either as percentage or as absolute power. |  | float with unit:IlluminationIntensityUnit
| 3 | CalibrationMap | NrOfImages | 3 |  | This field represents the number of images that were employed to generate this calibration matrix. |  | integer with unit:none
| 3 | Row |  | 3 |  |  | 1 … ∞
| 3 | Row | RowNr | 3 |  | The index of this row | integer with unit:none
| 3 | Cell |  | 3 |  |  | 1 … ∞
| 3 | Cell | CellNr | 3 |  | The index of this cell | integer with unit:none
| 3 | Cell | CellValue | 3 |  | The value of the metric contained in this cell | float with unit:CellValueUnit
| 4 | MechanicalCalibration |  | 4 |  | This describes the procedures that were used for mechanical calibration and the resulting calibration results. | 0 … ∞
| 4 | MechanicalCalibration | ID | 4 |  |  | R
| 4 | MechanicalCalibration | Date | 4 |  | The date in which this calibration was performed. | R
| 4 | MechanicalCalibration | MeasuredLateralDrift | 4 |  | Measured shift in (x, y) position over time. |  | float with unit:MeasureLaterlDriftUnit
| 4 | MechanicalCalibration | MeasuredFocalDrift | 4 |  | Measured shift in z position over time. |  | float with unit:MeasureFocalDriftUnit
| 4 | MechanicalCalibration | MeasuredRepeatability | 4 |  | Measured precision with which this device is capable of returning to a previously visited position (useful for tiling). |  | float with unit:MeasuredRepeatablityUnit
| 4 | MechanicalCalibration | MeasuredSettlingTime | 4 |  | Measured time it takes for this device to stabilize at the target position. |  | float with unit:MeasuredSettlingTimeUnit
| 4 | MechanicalCalibration | SampleHolderRef | 4 |  | This empty element has a reference (the SampleHolderID attribute) to a SampleHolder element defined within OME. | 0 … 1 | Extension of Reference
| 4 | MechanicalCalibration | FocusingRef | 4 |  | This empty element has a reference (the FocusingID attribute) to a Focusing element defined within OME. | 0 ... 1 | Extension of Reference
| 4 | MechanicalCalibration | AutoFocusRef | 4 |  | This empty element has a reference (the AutofocusID attribute) to a Autofocus element defined within OME. | 0 ... 1 | Extension of Reference
| 4 | MechanicalCalibration | AnnotationRef | 4 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommentAnnotation element, which consists of a simple comment multi-line annotation. | 0 ... 1 | Extension of Reference
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on E2: What is the meaning of Tier5 as far as Settings? It seems to me that Tier5 as it is defined now only applies to HARDWARE but not as much to SETTINGS… -Caterina Strambio De Castillia
Note on C64: This seems too specific. Should be generalized to include other similar experiments
Note on C154: Need to enter a Ref to a optical diagram describing how the different components are arranged
Note on D168: HOW DOES THIS GET MEASURED? -Caterina Strambio De Castillia