## Tier system_v00-03
-- required metadata --
Old ID | New ID | Name | Description | Example | Example Labelling | PSFj calibration | detector calibration | experiment/sample | specifications | settings
D1 | 1 | Documentation | Reporting qualitative effects without quantification | Transfection control, viability assay | FISH, Immuno Fluorescence | recommended annually (not required) | not recommended | - experiment, sample and image name, date and description - experimenter name | -microscope manufacturer, name and type -light source type -objective magnification and NA | - dimension order - physical pixel size x and y, physical z distance - time increment - optovar magnification - channel name, color and acquisition mode
D2 | 2 | Simple Quantification | Identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Counting of cells and nuclei, expression level measurements | FISH, Immuno Fluorescence | recommended quarterly (not required) | not recommended | - mounting medium | -microscope model -light source name, description, manufacturer, model, wavelenght -objective manufacturer, model, and correction -detector type, manufacturer, model, and relay magnification | - exposure time - immersion medium name and nominal refractive index - image X, Y, Z, C and T size - channel illumination type, contrast method, fluorophore, - filter manufacturer and name
A1 | 3 | Advanced Quantification | Identification and localization of refraction-limited particles, super-resolution microscopy | Measurement of distances, co-localization studies | Single Molecule FISH or Immuno Fluorescene | required quarterly (submit report) | recommended | - imaging temperature | - detailed light source, light source coupling, light path, channel and detector specification & settings - detailed light path specifications & settings - light source power output and detector calibration results
A2 | 4 | Life Cells Imaging | Tracking of intracellular dynamics | Particle tracking, live cell expression level measurements, limited signal, advanced processing, | CasFISH, CRISPRainbow, fluoresent protein labelling | required monthly (submit report) | highly-recommended | - imaging air pressure, humidity and CO2 percentage - mounting medium measured refractive index | - imaging standard type - light source coupling, foucusing and autofocus specification & settings - objective iris specifications - dichroic mirror and additional optics transmittance, reflectance and angles | - objective temperature - immersium medium measured refractive index
P | 5 | Virtuoso | Full reproducibility of microscopic set up and image acquisition | Development of novel unproved technology, new gold-standard, need to make sure microscopy set up can be reproduced in another laboratory | All of the above | required weekly (submit report) | highly-recommended | all the metadata specified by the data model (Figures 2 and 3) - including any novel technology-specific metrics
Note on C7: Alternative designation: Development Wizard Unicorn

## New OME proposal_Specs_v00-11
4DN-OME model
| Tier | Element | Attribute or Subelement | Tier | Complex type | Description | Cardinality | Data type | Allowed values
MICROSCOPE Speecifications | 1 | Instrument |  | 1 |  | This element describes the instrument used to capture the Image. The Filters elements are described in the Light Path and can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. In addition the Light Path might contain additional optics such as Lens, Prism, Mirrors, Polarization Optics and other Filters. Within the Image itself, a reference is made to the Instrument elements and, if appropriate, their settings. There may be multiple light sources, detectors, objectives and filters on a microscope. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the light path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. | 1
| 1 | Instrument | ID | 1 |  |  |  | InstrumentID
| 1 | Instrument | Name | 1 |  | A user defined name for this instrument. |  | string
| 3 | Instrument | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| NA | AnnotationRef |  | NA | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element
| NA | AnnotationRef | ID | NA |  |  |  | AnnotationID
| NA | CommentAnnotation |  | NA |  | A simple comment annotation
| NA | CommentAnnotation | ID | NA |  |  |  | AnnotationID
| NA | CommentAnnotation | Namespace | NA |  | We recommend the inclusion of a namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type.
| NA | CommentAnnotation | Annotator | NA |  | The Annotator is the person who attached this annotation. e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator.
| NA | CommentAnnotation/AnnotationRef | AnnotationRef | NA | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element |  | Extension of Reference
| NA | CommentAnnotation/Value | Value | NA |  | This element contains the actual annotation |  | string
| NA | CommentAnnotation/Description | Descrption | NA |  | This element contains a description for the annotation |  | string
| 1 | MicroscopeBody |  | 1 |  | A description of the body of this microscope. | 1 … ∞
| 1 | MicroscopeBody | Manufacturer | 1 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | MicroscopeBody | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | MicroscopeBody | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | MicroscopeBody | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | MicroscopeBody | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 1 | MicroscopeBody | ID | NA |  |  |  | MicroscopeID
| 1 | MicroscopeBody | Name | 1 |  | A user defined name for this element |  | string
| 1 | MicroscopeBody | Type | 1 | MicroscopeType | This element describes the type of this microscope |  | enum | Upright, Inverted, Dissection, Stereomicroscope, Electrophysiology, Custom, Other
| 3 | MicroscopeBody | EyepieceFieldNumber | 3 |  | The Field Number represents the diaphragm size of the microscope eyepiece expressed in mm. |  | float with unit:none
| 3 | MicroscopeBody | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 4 | MicroscopeTable |  | 4 |  | This element describes the table used to support this microscope | 0 … 1
| 4 | MicroscopeTable | Manufacturer | 4 | Manufacturer Spect | The manufacturer of the component |  | string
| 4 | MicroscopeTable | Model | 4 | Manufacturer Spect | The model of the component |  | string
| 5 | MicroscopeTable | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | MicroscopeTable | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | MicroscopeTable | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 4 | MicroscopeTable | ID | 4 |  |  |  | LSID
| 4 | MicroscopeTable | Name | 4 |  | A user defined name for this element |  | string
| 4 | MicroscopeTable | Type | 4 | MicroscopeTableTYpe | This element describes the type of this microscope table |  | enum | Simple Bench, Non-isolating Frame, Active Air, Passive Air, Stone, Other
| 5 | MicroscopeTable | TableWidth | 5 |  | The width of this microscope table |  | float with unit:TableDepthUnit
| 5 | MicroscopeTable | TableDepth | 5 |  | The depth of this microscope table |  | float with unit:TableWidthUnit
| 1 | LightSource-APE |  | 1 | LightSource | This element describes the light source(s) available in this microscope | 1 … ∞ | Substitution : choose one of the children elements from the list below | Laser, Filament, Arc, LightEmittingDiode, GenericLightSource, Other
| 1 | LightSource-APE | Manufacturer | 1 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | LightSource-APE | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | LightSource-APE | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | LightSource-APE | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | LightSource-APE | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 1 | LightSource-APE | ID | 1 |  | A LightSource ID must be specified for each light source, and the individual light sources can be referred to by their LightSource IDs (eg from Channel). |  | LightSourceID
| 1 | LightSource-APE | Name | 1 |  | A user defined name for this element |  | string
| 3 | LightSource-APE | NominalPower | 3 |  | The nominal power of the light-source |  | float with unit:PowerUnit
| 3 | LightSource-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 1 | Laser |  | 1 |  | A laser ("light amplification by stimulated emission of radiation") is a device that emits light through a process of optical amplification based on the stimulated emission of electromagnetic radiation.
| 5 | Laser | Type | 5 | LaserType | The general category of this laser |  | enum | Gas, MetalVapor, SolidState, Dye, Semiconductor, FreeElectron, Other
| 5 | Laser | LaserMedium | 5 | :aserMedium | This is the material this laser consists of |  | enum | Cu, Ag, ArFl, ArCl, KrFl, KrCl, XeFl, XeCl, XeBr, N, Ar, Kr, Xe, HeNe, HeCd, CO, CO2, H2O, HFl, NdGlass, NdYAG, ErGlass, ErYAG, HoYLF, HoYAG, Ruby, TiSapphire, Alexandrite, Rhodamine6G, CoumarinC30, GaAs, GaAlAs, EMinus, Other
| 2 | Laser | Wavelength | 2 |  | The center Wavelength of the light emitted by this laser |  | float with unit:Wavelength
| 5 | Laser | FrequencyMultiplication | 5 |  | FrequencyMultiplication that may be specified |  | integer with unit:none
| 5 | Laser | Tuneable | 5 |  | Whether or not the laser is Tuneable |  | boolean
| 5 | Laser | Pulse | 5 | LaserPulse | The Pulse mode of the laser |  | enum | CW, Single, QSwitched, Repetitive, ModeLocked, Other
| 5 | Laser | PockelCell | 5 |  | If true the laser has a PockelCell to rotate the polarization of the beam |  | boolean
| 5 | Laser | RepetitionRate | 5 |  | The rate in Hz at which the laser pulses if the Pulse type is 'Repetitive' |  | foat with unit:RepetitionRateUnit
| 5 | Pump |  | 5 |  | The pump element is a reference to a LightSource. It is used within the Laser element to specify the light source for the laser's pump (if any).
| 5 | Pump | ID (LightSourceID) | 5
| 1 | Filament |  | 1 |  | The Filament element is used to describe various kinds of filament bulbs such as Incadescent or Halogen.
| 3 | Filament | Type | 3 | FilamentType | The type of Filament lamp |  | enum | Incandescent, Halogen, Other
| 1 | Arc |  | 1 |  | This element describes various kinds of Arc lamps
| 3 | Arc | Type | 3 | ArcType | The type of Arc lamp |  | enum | Hg, Xe, HgXe, Other
| 1 | LightEmittingDiode |  | 1 |  | The LightEmittingDiode element is used to describe various kinds of LED lamps.
| 3 | LightEmittingDiode | Type | 3 | LightEmittingDiodeType | The type of LightEmittingDiode light source |  | enum | Unmounted, Mounted, Collimate, Other
| 2 | LightEmittingDiode | PeakWavelenght | 2 |  | The wavelength at which there is maximal response rate |  | float with unit: PeakWavelengthUnit
| 3 | LightEmittingDiode | LumenPower | 3 |  | The amount of light this light source generate at maximum expressed in lumen. The lumen is the SI derived unit of luminous flux, a measure of the total quantity of visible light emitted by a source |  | float with unit:lumen
| 2 | LightEmittingDiode | LEDWavelengthRange | 3 |  | This records a range of wavelengths. Multiple of this elements can be used | 1 … ∞ | Element
| 5 | LightEmittingDiode | ProjectionAngle | 5 |  | The angle at which this light source projects light |  | integer with unit:ProjectionAngleUnit
| 2 | LEDWavelengthRange | WavelengthCutIn | 2 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 2 | LEDWavelengthRange | WavelengthCutOut | 2 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutUnit
| 5 | LEDWavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 1 | GenericExcitationSource |  | 1 |  | The GenericExcitationSource element is used to represent a source as a collection of key/value pairs, stored in a Map. The other lightsource objects should always be used in preference to this if possible.
| 4 | LightSourceCoupling-APE |  | 4 | LightSourceCoupling | This element describes the method(s) used to couple the light source with the light path moving through speciment to the detector. | 1 … ∞ | Substitution : choose one child from the list below | LighthGuide, Direct
| 4 | LightSourceCoupling-APE | Manufacturer | 4 | Manufacturer Spect | The manufacturer of the component |  | string
| 4 | LightSourceCoupling-APE | Model | 4 | Manufacturer Spect | The model of the component |  | string
| 5 | LightSourceCoupling-APE | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | LightSourceCoupling-APE | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | LightSourceCoupling-APE | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 4 | LightSourceCoupling-APE | ID | 4 |  |  |  | LightSourceCouplingID
| 4 | LightSourceCoupling-APE | Name | 4 |  | A user defined name for this element |  | string
| 5 | LightSourceCoupling-APE | Transmittance | 5 |  | The amount of light this LightSourceCoupling apparatus transmits at a maximum. A fraction, as a value from 0.0 to 1.0 |  | float with unit:none
| 5 | LightSourceCoupling-APE | CouplingLensRef | 5 |  | This element is a reference to a Lens element |  | LensRef
| 3 | LightSourceCoupling-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 4 | LightGuide |  | 4 |  | Light Guides interface with illuminators to transfer light to the speciment in a usable manner
| 5 | LightGuide | Type | 5 | LightGuideType | The type of light guide used in this case |  | enum | Fiber optics, Liquid, Other
| 5 | LightGuide | Geometry | 5 | LightGuideGeometry | The shape of this light guide |  | enum | Spot Light, Line Light, Ring Light, Backlight, Diffuse Axial, Other
| 5 | LightGuide | MaterialName | 5 | LightGuideMaterialName | This records the chemical composition or the nonproprietary trade designation of the material used for this light guide. |  | enum | Quartz, Silica, Borosilicate, Other
| 5 | LightGuide | PeakWavelength | 5 |  | The wavelength associated with maximum light intensity transmitted by this light guide |  | float with unit:PeakWavelengthUnit
| 5 | LightGuide | LightGuideWavelengthRange | 5 |  | This records a range of wavelengths. Multiple of this elements can be used | 1 … ∞ | Element
| 5 | LightGuide | SingleMode | 5 |  | Records whether or not the light guide is single-mode. If not, then the guide is assumend to be multi-modal. [flag] |  | boolean
| 5 | LightGuide | WaveguideMode | 5 | LightGuideWaveguideMode | When boundary conditions are imposed by a physical structure, a wave of a particular frequency can be described in terms of a transverse mode (or superposition of such modes). These modes generally follow different propagation constants. |  | enum | Transverse electromagnetic, Transverse electric, Transverse magnetic, Hybrid, Other
| 5 | LightGuide | RefractiveIndex | 5 |  | The refractive index of the material this light guide is made of. |  | float with unit:none
| 5 | LightGuide | NumericalAperture | 5 |  | The numerical aperture of the light guide expressed as a floating point (real) number. Expected range 0.02 - 1.5. |  | float with unit:none
| 5 | LightGuide | AcceptanceAngle | 5 |  | The angle within which light beam may be coupled into the bound modes of an optical fiber. |  | float with unit:AcceptanceAngleUnit
| 5 | LightGuide | Diameter | 5 |  | The outer diameter of the light guide |  | float with unit:DiameterUnit
| 5 | LightGuideWavelengthRange | WavelengthCutIn | 2 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 5 | LightGuideWavelengthRange | WavelengthCutOut | 2 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutUnit
| 5 | LightGuideWavelengthRange | WavelengthProfile | 5 | WavelengthBounds | A file containing the wavelenght profile of this component. |  | anyURI
| 4 | Direct |  | 4 |  | Direct coupling consists of either direct mounting or a optical light path that connects the Light Source with the microscope.
| 5 | Direct | Type | 5 | DirectType | The type of direct coupling available for this Instrument. |  | enum | Mounted, Optical light path, Other
| 5 | Direct | Mount | 5 | DirectMount | The mount utilized to attach a LightSource to the microscope. |  | string
| 3 | SampleHolder |  | 3 |  | This element describes an optional sample holder mounted on the microscope stage | 0 … ∞
| 3 | SampleHolder | Manufacturer | 3 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | SampleHolder | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | SampleHolder | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | SampleHolder | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | SampleHolder | SpecsFile | 5 | Manufacturer Spect | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 3 | SampleHolder | ID | 3 | Stage |  |  | SampleHolderID
| 3 | SampleHolder | Name | 3 | Stage | A user defined name for this element |  | string
| 5 | SampleHolder | MovementAxes | 5 | Stage | The axis along which this device moves |  | enum | X, Y, Z, XY, XYZ, Other
| 5 | SampleHolder | XTravelRange | 5 | Stage | The X axis range this device is capable of moving in |  | Float with unit:XAxisTravelRangeUnit
| 5 | SampleHolder | YTravelRange | 5 | State | The Y axis range this device is capable of moving in |  | Float with unit:YAxisTravelRangeUnit
| 5 | SampleHolder | ZTravelRange | 5 | Stage | The Z axis range this device is capable of moving in |  | Float with unit:ZAxisTravelRangeUnit
| 5 | SampleHolder | MaxVelocity | 5 | Stage | The maximum velocity this device is capable of moving at |  | Float with unit:MaxVelocityUnit
| 5 | SampleHolder | Type | 5 |  | The type of sample holder |  | enum | Rigid Stand, Translational, Other
| 5 | SampleHolder | RotationAngle | 5 |  | The angle by which this device is capable of rotating |  | Float with unit:RotationAngleUnit
| 3 | SampleHolder | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 3 | Focusing-APE |  | 3 | Substitution : choose one child from the list below | This element describes the objective lens and stage positioning device(s) available in this microscope. An instrument may have more than one of these devices. The type of focusing device is specified by one of the child-element which are 'MechanicalStage', 'PiezoElectricSate', or 'PiezoElectricObjective'. | 1 … ∞ |  | MechanicStage, PiezoelectricStage, PiezoelectricObjective, Other
| 3 | Focusing-APE | Manufacturer | 3 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Focusing-APE | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Focusing-APE | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Focusing-APE | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Focusing-APE | SpecsFile | 5 | Manufacturer Spect | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 3 | Focusing-APE | ID | 3 | Stage |  |  | FocusingID
| 3 | Focusing-APE | Name | 3 | Stage | A user defined name for this element |  | string
| 5 | Focusing-APE | MovementAxes | 5 | Stage | The axis along which this device moves |  | enum | X, Y, Z, XY, XYZ, Other
| 5 | Focusing-APE | XTravelRange | 5 | Stage | The X axis range this device is capable of moving in. |  | float with unit:XTravelRangeUnit
| 5 | Focusing-APE | YTravelRange | 5 | Stage | The Y axis range this device is capable of moving in |  | float with unit:YTravelRangeUnit
| 5 | Focusing-APE | ZTravelRange | 5 | State | The Z axis range this device is capable of moving in |  | float with unit:ZTravelRangeUnit
| 5 | Focusing-APE | MaxVelocity | 5 | Stage | The maximum velocity this device is capable of moving at |  | float with unit:MaximumVelocity
| 5 | Focusing-APE | Resolution | 5 |  | The motion resolution of this device |  | float with unit:ResolutionUnit
| 5 | Focusing-APE | PositionLinearityError | 5 |  | The positioning error associated with this device |  | float with unit:PositionLinearityErrorUnit
| 5 | Focusing-APE | SettlingTime | 5 |  | Measures the time it takes for this device to stabilize at the target position |  | float with unit:SettlingTimeUnit
| 5 | Focusing-APE | Repetability | 5 |  | Measures the precision with which this device is capable of returning to a previously visited position |  | float with unit:RepeatablityUnit
| 3 | Focusing-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 3 | MechanicalStage |  | 3 |  | This is a mechanical stage positioning device. It is not very common but can be found in simple surveying microscopes.
| 3 | PiezoElectricStage |  | 3 |  | This is a piezoelectric stage positioning device.
| 3 | ObjectiveTurret |  | 3 |  | This is a objective turret positoning device
| 3 | PiezoElectricObjective |  | 3 |  | This is a piezoelectric objective-lens positioning device.
| 3 | AutoFocus |  | 3 |  | This element describes the focus stabilization device(s) available for this microscope | 0 … ∞
| 3 | AutoFocus | Manufacturer | 3 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | AutoFocus | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | AutoFocus | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | AutoFocus | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | AutoFocus | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 3 | AutoFocus | ID | 3 |  |  |  | AutoFocusID
| 3 | AutoFocus | Name | 3 |  | A user defined name for this element |  | string
| 3 | AutoFocus | Type | 3 | AutoFocusType | The type of this autofocus device |  | enum | Image based, Optical, Other
| 5 | AutoFocus | AutoFocus/Description | 5 |  | A description for this autofocus device |  | multi-line string
| 3 | AutoFocus | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 1 | Optovar |  | 1 |  | Optovars are magnification changing devices originally developed by Zeiss, which provide additional magnification that can be beneficial when a desired field of view is required for the imaging sensor. Magnification changer components are placed before the tube lens within the infinity space of the optical system. Typical magnification increments are: 1.25x, 1.6x, and 2x. | 0 … ∞
| 1 | Optovar | Manufacturer | 1 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Optovar | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Optovar | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Optovar | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Optovar | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 1 | Optovar | ID | 1 |  |  |  | ID
| 1 | Optovar | Name | 1 |  | A user defined name for this element |  | Denomination
| 3 | Optovar | LensType | 3 |  | The type of lens used by the Magnification Changer available with this microscope. |  | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 1 | Objective |  | 1 |  | The microscope's objective lens is a compound lens consisting of several simple lenses (elements), usually arranged along a common axis. | 1 … ∞
| 1 | Objective | Manufacturer | 1 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Objective | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Objective | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Objective | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Objective | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 1 | Objective | ID | 1 |  |  |  | ObjectiveID
| 1 | Objective | Name | 1 |  | A user defined name for this element |  | Denomination
| 2 | Objective | Correction | 2 | ObjectiveCorrection | The correction applied to the lens |  | enum | UV, PanApo, PlanFluor, SuperFluor, VioletCorrected, Achro, Achromat, Fluor, FI, Fluar, Neofluar, Apo, PlanNeoFluar, Other
| 2 | Objective | NominalImmersionMedium | 2 | ObjectiveNominalImmmersionMedium | The immersion medium the lens is designed for |  | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 1 | Objective | LensNA | 1 |  | The numerical aperture of the lens expressed as a floating point (real) number. Expected range 0.02 - 1.5 |  | float with unit:none
| 1 | Objective | NominalMagnification | 1 |  | The magnification of the lens as specified by the manufacturer - i.e. '60' is a 60X lens |  | float with unit:none
| 5 | Objective | CalibratedMagnification | 5 |  | The magnification of the lens as measured by a calibration process- i.e. '59.987' for a 60X lens |  | float with unit:none
| 3 | Objective | WorkingDistance | 3 |  | The working distance of the lens expressed as a floating point (real) number |  | float with unit:WorkingDistanceUnit
| 4 | Objective | Iris | 4 |  | Records whether or not the objective was fitted with an Iris. [flag] |  | boolean
| 3 | Objective | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 2 | Lens |  | 2 |  | A lens is a component of glass or transparent plastic material, usually circular in diameter, which has two primary surfaces that are ground and polished in a specific manner designed to produce either a convergence or divergence of light passing through the material. This element describes a lens utilized in this microscope | 0 … ∞
| 2 | Lens | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Lens | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Lens | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Lens | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Lens | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | Lens | ID | 2 |  |  |  | LensID
| 2 | Lens | Name | 2 |  | A user defined name for this Lens element. |  | string
| 5 | Lens | Type | 5 | LensType | The general category this lens belongs to |  | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 5 | Lens | Geometry | 5 | LensGeometry | The geometrical shape of the lens surfaces |  | enum | Spherical, Aspherical, Cylindrical, Ball, Hemispheric-ball, Cone, Drum, Rod, Toric
| 5 | Lens | Construction | 5 | LensConstruction | Whether or not the lens is simple or composite |  | enum | Simple, Composite, Other
| 2 | Lens | Magnification | 2 |  | The nominal magnification provided by this lens |  | integer with unit:none
| 5 | Lens | AttenuationCoefficient | 5 |  | The attenuation coefficient is a material property that characterizes the degree of transparency of a medium. It is a fraction and therefore it has no units. |  | float with unit:none
| 5 | Lens | RadiusOfCurvature | 5 |  | A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusUnit
| 5 | Lens | FocalDistanceToObject | 5 |  | The distance between the front principal plane of the lens and the specimen |  | float with unit:FocalDistanceToObjectUnit
| 5 | Lens | FocalDistanceToImage | 5 |  | The distance from the rear principal plane to the image |  | float with unit:FocalDistanceToImageUnit
| 5 | Lens | FocalLenght | 5 |  | The distance from the rear principal plane of the lens and its focal point |  | float with unit:FocalLengthUnit
| 5 | Lens | MaterialName | 5 | LensMeterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | alpha-Barium Borate, N-BK7 (Borosilicate Crown), Calcium Fluoride , UV-Grade Fused Silica (UVFS), Germanium, Magnesium Fluoride, N-SF11 (Dense Flint), Polytetrafluoroethylene (PFTE), Silicon (Si), Zinc Selenide (ZnSe), Other
| 5 | Lens | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | Lens | RefractiveIndex | 5 |  | The refractive index of the material the prism is made of |  | float with unit:none
| 5 | Lens | AbbeNumber | 5 |  | The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion |  | float with unit:none
| 5 | Lens | Density | 5 |  | The density of the material this lens is made of |  | float with unit:DensityUnit
| 5 | Lens | Diameter | 5 |  | The physical diamter of this lens |  | float with unit:none
| 3 | Lens | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 2 | Prism |  | 2 |  | A prism is a block of glass or other transparent material with flat, polished surfaces. A prism can be employed to deflect or deviate a light beam, rotate or invert an image, separate polarization states, or disperse light into its component wavelengths. Prisms can be classified in three general categories: reflecting prisms, polarizing prisms, and refracting or dispersion prisms. Reflecting prisms are useful for redirecting light beams by total internal reflection. Refracting prisms can be employed to bend and separate light into its component colors. Polarizing prisms are birefringent crystals that divide incident non-polarized light into separate components polarized orthogonally to each other. These prisms are employed to produce polarized light for optical instruments such as microscopes and polarimeters. A prism can be viewed as an ensamble of different rigidly mounted mirrors, with each face having a permanent orientation with respect to all the others. | 0 … ∞
| 2 | Prism | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Prism | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Prism | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Prism | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Prism | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | Prism | ID | 2 |  |  |  | PrismID
| 2 | Prism | Name | 2 |  | A user defined name for this Prism element. |  | string
| 3 | Prism | Type | 3 | PrismType | The category this prism belongs to |  | enum | Reflecting, Polarizing, Refracting, Other
| 5 | Prism | Geometry | 5 | PrismGeometry | The geometrical shape of this prism |  | enum | Right-Angle, Porro, Dove, Amici, Penta, Romboid, Nicol, Glan-Thompson, Wollaston, Rochon, Senarmont, Triangular, Abbe prism, Pellin–Broca, Grim, Compound, Other
| 5 | Prism | PrismWavelengthRange | 5 |  | The range of wavelengths at which this prism operates. Multiple of this elements can be used. | 1 … ∞ | Element
| 5 | Prism | Reflectance | 5 |  | The amount of light the prism reflects at a maximum. It is represented as a fraction, as a value from 0.0 to 1.0. |  | float with unit:none
| 5 | Prism | PrismAngle | 5 |  | The angle between two surfaces of a prism |  | float with unit:PrismAngleUnit
| 5 | Prism | IncidenceAngle | 5 |  | The angle that the incident beam of light makes with the prism surface, which is typically measured relative to the perpendicular to the surface |  | float with unit:IncidenceAngleUnit
| 5 | Prism | DeviationAngle | 5 |  | A beam passing through an object like a prism is deflected twice: once entering, and again when exiting. The sum of these two deflections is called the deviation angle and is measured with respect to the incident ray |  | float with unit:DeviationUnitAngle
| 5 | Prism | MaterialName | 5 | PrismMaterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | Calcium Fluoride, F2 (Flint), N-BK7 (Borosilicate Crown), N-SF11 (Dense Flint), Rutile (TiO2), UV-Grade Fused Silica (UVFS), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | Prism | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | Prism | RefractiveIndex | 5 |  | The refractive index of the material the prism is made of |  | float with unit:none
| 5 | Prism | AbbeNumber | 5 |  | The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion |  | float with unit:none
| 5 | Prism | Density | 5 |  | The density of the material the prism is made of |  | float with unit:DensityUnit
| 3 | Prism | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 5 | PrismWavelengthRange | WavelengthCutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 5 | PrismWavelengthRange | WavelengthCutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutUnit
| 5 | PrismWavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 2 | PolarizationOptics |  | 2 |  | This is an optical filter that lets light waves of a specific polarization pass through while blocking light waves of other polarizations | 0 … ∞
| 2 | PolarizationOptics | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | PolarizationOptics | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | PolarizationOptics | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | PolarizationOptics | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | PolarizationOptics | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | PolarizationOptics | ID | 2 |  |  |  | PolarizationOpticsID
| 2 | PolarizationOptics | Name | 2 |  | A user defined name for this Polarization Optics element. |  | string
| 2 | PolarizationOptics | Function | 2 | PolarizationOpticsFunction | This records whether this polarizer optics is used to generate polarized light or to analyze polarized light |  | enum | Polarizer, Analyzer
| 3 | PolarizationOptics | Construction | 3 | PolarizationOpticsConstruction | The modality in with which this Polarization Optics is constructed |  | enum | Achromatic, Zero Order, Multiple Order, Other
| 5 | PolarizationOptics | Fabrication | 5 | PolarizationOpticsFabrication | The way in with which this Polarization Optics is fabricated |  | enum | Wire Grid, Film, Polarizing Beamsplitter, Polarizing BandPass, Variable, Birefringent, Polymer, Quarz, Other
| 3 | PolarizationOptics | PolarizerWavelengthRange | 3 |  | The range of wavelengths at which this polarizing optic operates. Multiple of this elements can be used. | 1 … ∞ | Element
| 3 | PolarizationOptics | Retardation | 3 | PolarizationOpticsRetardation | Retardation describes the phase shift between the polarization component projected along the fast axis and the component projected along the slow axis |  | enum | Half Wave, Quarter Wave, Full Wave, Other
| 5 | PolarizationOptics | MaterialName | 5 | PolarizationOpticsMaterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | alpha-Barium Borate, Calcite, Calcium Fluoride, N-BK7 (Borosilicate Crown), Rutile (TiO2), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | PolarizationOptics | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 3 | PolarizationOptics | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 3 | PolarizerWavelengthRange | WavelengthCutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 3 | PolarizerWavelengthRange | WavelengthCutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutUnit
| 5 | PolarizerWavelengthRange | WavelengthProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 2 | FilterSet |  | 2 |  | A filter set is typically composed of one excitation filter, one emission filter and a dichroic. Often they are pre-assebled by the manufacturer and pruchased as a whole | 0 … ∞
| 2 | FilterSet | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | FilterSet | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | FilterSet | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | FilterSet | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | FilterSet | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | FilterSet | ID | 2 |  |  |  | FilterSetID
| 2 | FilterSet | Name | 2 |  | A user defined name for this FilterSet. |  | string
| 3 | FilterSet | ExcitationFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Excitation light path | 0 … ∞ | Extension of Reference
| 3 | FilterSet | DichroicRef | 3 | Reference | This empty element has a reference (the Mirror ID attribute) to the Mirror/Dichroic placed in the light path | 0 … ∞ | Extension of Reference
| 3 | FilterSet | EmissionFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Emission light path | 0 … ∞ | Extension of Reference
| 3 | FilterSet | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 2 | Filter |  | 2 |  | An optical filter is a device that selectively transmits light of different wavelengths. A filter maybe either an excitation or an emission filter. For each fluorescence Image there should be at least one filter element specified per Channel. This element can be associated with an optional FilterWheel - Note: this is not the same as a FilterSet | 1 … ∞
| 2 | Filter | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Filter | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Filter | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Filter | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Filter | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | Filter | ID | 2 |  |  |  | FilterID
| 2 | Filter | Name | 2 |  | A user defined name for this Filter. |  | string
| 3 | Filter | Type | 3 | FilterType | The filter type |  | enum | LongPass, ShortPass, BandPass, MultiPass, NeutralDensity, Tunable, Other
| 4 | Filter | FilterWheel | 4 |  | A filter 'wheel' in OME can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. [plain text string] |  | string
| 3 | Filter | TransmittanceRange | 3 |  | This records the range of wavelengths that are transmitted by the filter. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | 1 … ∞ | Element
| 5 | Filter | ReflectionCoating | 5 |  | An optical coating is one or more thin layers of material deposited on an optical component such as a lens or mirror, which alters the way in which the optic reflects and transmits light. |  | string
| 3 | Filter | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 3 | TransmittanceRange | TransmittanceCutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 3 | TransmittanceRange | TransmittanceCutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutToleranceUnit
| 5 | TransmittanceRange | TransmittanceProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 3 | TransmittanceRange | Transmittance | 4 |  | The amount of light the filter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. |  | float with unit:none
| 2 | Mirror-APE |  | 2 | Substitution : choose one child from the list below | This element describes mirror component(s) available in this microscope | 1 … ∞ |  | Dichroic, Beamsplitter, Reflector, Other
| 2 | Mirror-APE | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | Mirror-APE | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 5 | Mirror-APE | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Mirror-APE | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Mirror-APE | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | Mirror-APE | ID | 2 |  |  |  | MirrorID
| 2 | Mirror-APE | Name | 2 |  | A user defined name for this Mirror. |  | string
| 5 | Mirror-APE | Type | 5 | MirrorType | The general category this lens belongs to |  | enum | Convex, Concave, Planar, Other
| 5 | Mirror-APE | Geometry | 5 | MirrorGeometry | The geometrical shape of the surface of this mirror |  | enum | Spherical, Aspherical, Cylindrical, Ellipsoidal, Hyeprboloidal, Paraboloidal, Other
| 4 | Mirror-APE | Reflectance | 4 |  | The amount of light the mirror reflects at a maximum. It is represented as a fraction, as a value from 0.0 to 1.0. |  | float with unit:none
| 5 | Mirror-APE | ReflectionAngle | 5 |  | The angle at which light is reflected by this mirror |  | float with unit:ReflectionAngleUnit
| 5 | Mirror-APE | RadiusOfCurvature | 5 |  | A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusOfCurvatureUnit
| 5 | Mirror-APE | MaterialName | 5 | MirrorMaterialName | This records the chemical composition or the nonproprietary trade designation of the glass. |  | enum | Calcium Fluoride, N-BK7 (Borosilicate Crown), UV-Grade Fused Silica (UVFS), Zerodur (Lithium Aluminum Ailicon Oxide), Zinc Selenide (ZnSe), Other
| 5 | Mirror-APE | GlassCode | 5 |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 3 | Mirror-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 2 | Dichroic |  | 2 |  | The dichromatic beamsplitter or dichroic mirror used for this filter combination.
| 3 | Dichroic | Type | 3 |  | The type of dichroic mirror used in this case |  | string | Longpass, Shortpass, Other
| 4 | Dichroic | FilterWheel | 4 |  | A filter 'wheel' in OME can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. [plain text string] |  | string
| 5 | Dichroic | Polarization | 5 |  | The polarization state of this dichroic |  | enum | Linear, Circular, Elliptical, Other
| 3 | Dichroic | TransmittanceRange | 3 |  | This records the range of wavelengths that are transmitted by the dichroic mirror. It also records the maximum amount of light transmitted | 1 … ∞ | Element
| 3 | Dichroic | ReflectanceRange | 3 |  | This records the range of wavelengths that are reflected by the mirror. | 1 … ∞ | Element
| 3 | TransmittanceRange | TransmittanceCutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 3 | TransmittanceRange | TransmittanceCutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutToleranceUnit
| 5 | TransmittanceRange | TransmittanceProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 3 | TransmittanceRange | Transmittance | 4 |  | The amount of light the filter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. |  | integer with unit:none
| 3 | ReflectanceRange | ReflectanceCutIn | 3 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 3 | ReflectanceRange | ReflectanceCutOut | 3 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutToleranceUnit
| 5 | ReflectanceRange | ReflectanceProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 2 | Beamsplitter |  | 2 |  | This is a mirror that reflects part of the light, regardless of wavelenght
| 4 | Beamsplitter | Transmittance | 4 |  | The amount of light the beamsplitter transmits at a maximum. A fraction, as a value from 0.0 to 1.0. |  | float with unit:none
| 5 | Beamsplitter | TransmissionAngle | 5 |  | The angle at which light is transmitted by this beamsplitter |  | float with unit:none
| 2 | Reflector |  | 2 |  | This is a total reflector mirror
| 2 | Detector-APE |  | 2 | Substitution : choose one child from the list below | This element describe the detector(s) available in this microscope. The Detector ID can be used as a reference within the Channel element in the Image element. The values stored in Detector represent the fixed values, variable values modified during the acquisition go in DetectorSettings | 1 … ∞ |  | Conventional CCD, EMCCD, Interline CCD, Conventional Intensified CCD, EBCCD, AnalogVideo, Conventional CMOS, Scientific CMOS, PMT, Photodiode, APD, Generic Detector, Other
| 2 | Detector-APE | Manufacturer | 2 | Manufacturer Spect | The manufacturer of the component |  | string
| 2 | Detector-APE | Model | 2 | Manufacturer Spect | The model of the component |  | string
| 5 | Detector-APE | SerialNumber | 5 | Manufacturer Spect | The serial number of the component |  | string
| 5 | Detector-APE | LotNumber | 5 | Manufacturer Spect | The lot number of the component |  | string
| 5 | Detector-APE | SpecsFile | 5 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 2 | Detector-APE | ID | 2
| 2 | Detector-APE | Name | 2 |  | A user defined name for this Detector. |  | string
| 2 | Detector-APE | RelayLensRef | 2 |  | This element refers to a Lens positioned at this Detector mount. |  | float with unit:RelayMagnificationUnit
| 3 | Detector-APE | Gain | 3 |  | The analog gain is the conversion unit for this Detector. It specifies how many electrons correspond to 1 signal unit. It is expressed in e-/Analog-to-Digit-Unit (ADU). |  | float with unit:none
| 3 | Detector-APE | Offset | 3 |  | The offset is an electronic baseline added to the output signal from the EMCCD sensor to ensure that the displayed signal level is always a positive number of counts. This the maximal nominal Offset of this detector. |  | float with unit:none
| 3 | Detector-APE | Zoom | 3 |  | The maximal nominal Zoom factor utilized for this detector. |  | float with unit:none
| 5 | Detector-APE | ReadNoise | 5 |  | Read noise includes thermal noise build-up and clock induced charge. This is generally expressed in e- rms.
| 5 | Detector-APE | DetectorNoiseModel | 5 | DetectorNoiseModel | The noise contribution that comes from the detector |  | enum | Poisson, Gaussian, Other
| 5 | Detector-APE | DarkCurrentRate | 5 |  | Median current detected in the absence of light. Generally expressed in e-/pixel/sec |  | float with unit:DarkCurrentRateUnit
| 5 | Detector-APE | PeakWavelength | 5 |  | The wavelength associated with maximum light intensity detected by this device. |  | float with unit:PeakWavelengthUnit
| 5 | Detector-APE | QuantumEfficiencyatPeakWavelength | 5 |  | The Quantum Efficiency observed at peak wavelength. Quantum Efficiency is generally defined as is the ratio of incident photon to converted electrons. Typically expressed as a number between 0 and 1. |  | float with unit:none
| 5 | Detector-APE | AmbientOperatingTemperature | 5 |  | The ambient temperature at which this detector is designed to operate. |  | float with unit:TemperatureUnit; default: Celsius[ºC]
| 5 | Detector-APE | AbientOperatingHumidity | 5 |  | The ambient humidity at which this detector is designed to operate. |  | float with unit:none
| 5 | Detector-APE | DetectorWavelengthRange | 5 |  | This records the range of wavelengths this detector is capable of sensing. Multiple of this elements can be used. | 1 … ∞ | Element
| 3 | Detector-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) |  | Extension of Reference
| 5 | DetectorWavelengthRange | WavelengthCutIn | 5 | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% |  | integer with unit: CutInUnit
| 5 | DetectorWavelengthRange | WavelengthCutOut | 5 | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% |  | integer with unit: CutOutToleranceUnit
| 5 | DetectorWavelengthRange | WavelenghtProfile | 5 | WavelengthBounds | The location of a file containing the wavelenght profile of this component |  | anyURI
| 2 | Camera-APE |  | 2 | Substitution : choose one child from the list below | This element describes the camera(s) devices available in this microscope. A camera is device that converts optical brightness into an electrical signal using a CCD(a semiconductor chip, which converts the intensity of light falling onto its surface at different points into a charge) and then reproduces an image of the specimen. |  |  | CCD, Intensified CCD, AnalogVideo, CMOS, Other
| 3 | Camera-APE | Illumination | 3 | CameraIllumination | The type of illumination the camera chip receives |  | enum | Front, Back, Other
| 3 | Camera-APE | ArrayWidth | 3 |  | The width of this camera chip expressed in pixel. |  | integer with unit:pixel
| 3 | Camera-APE | ArrayHeight | 3 |  | The height of this camera chip expressed in pixel. |  | integer with unit:pixel
| 3 | Camera-APE | PixelWidth | 3 |  | The physical width of each pixel |  | float with unit:PixelWidthUnit; default:µm
| 3 | Camera-APE | PixelHeight | 3 |  | The physical height of each pixel |  | float with unit:PixelHeightUnit; default:µm
| 3 | Camera-APE | ImageWidth | 3 |  | The width of the image generated by the camera |  | float with unit:ImageWidthUnit
| 3 | Camera-APE | ImageHeight | 3 |  | The width of the image generated by the camera |  | float with unit:ImageHeightUnit
| 5 | Camera-APE | PixelWellCapacity | 5 |  | The maximum amount of charge each pixel well can hold, measured in electron counts |  | integer with unit:none
| 5 | Camera-APE | RegisterWellCapacity | 5 |  | The maximum amount of charge each detector pixel can hold, measured in electron counts |  | integer with unit:none
| 5 | Camera-APE | VerticalClockSpeed | 5 |  | The maximal speed each vertical colunm of pixels in the camera chip, can be shifted down and read |  | float with unit:VerticalClockSpeedUnit
| 5 | Camera-APE | MaximumReadoutRate | 5 |  | The maximal pixel read out rate allowed by this camera |  | float with unit:MaximalReadoutRateUnit
| 2 | CCD |  | 2 |  | A charged coupled device camera
| 3 | CCD | Type | 3 | CCDType | The type of CCD camera used |  | enum | Conventional, EMCCD, Interline, Other
| 2 | IntensifiedCCD |  | 2 |  | An intensified charged coupled device camera
| 3 | IntensifiedCCD | Type | 3 | IntensifiedCCDType | The type of Intensified CCD camera used |  | enum | Conventional, EBCCD, Other
| 2 | AnalogVideo |  | 2 |  | Andevice used to capture analog videos
| 2 | CMOS |  | 2 |  | A complementary metal–oxide–semiconductor camera
| 3 | CMOS | Type | 3 | CMOSType | The type of CMOS camera used |  | enum | Conventional, Scientific, Other
| 2 | PhotoMultiplier |  | 2 |  | A photomultiplier tube, useful for light detection of very weak signals, is a photoemissive device in which the absorption of a photon results in the emission of an electron. These detectors work by amplifying the electrons generated by a photocathode exposed to a photon flux.
| 3 | PhotoMultiplier | Type | 3 | PhotoMultiplierType | The type of photomultiplier used. |  | enum | PMT, Photodiode, APD, Other
| 5 | PhotoMultiplier | RadiantSensitivity | 5 |  | The radiant sensitivity observed with this PhotoMultiplier at peak wavelength |  | float with unit:RadiantSensitivityAtPeakWavelengthUnit
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which 1): BOXES represent real-life entity-sets, about which information need to be collected. One way to think about it is that ENTITIES correspond to NOUNS. 2) LINES represent RELATIONSHIPS that connect entities with each other. Namely, RELATIONSHIPS correspond to VERBS (i.e. actions/states/occurrences). 3) 3) ATTRIBUTES corresponds to ADJECTIVES, which capture the actual information that needs to be collected about each entity.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on H2: How do we deal with OPTIONAL vs. MANDATORY fields vs. TIER system -Caterina Strambio De Castillia
Note on C18: We do not have any place in which the microscope acquisition software is described -Caterina Strambio De Castillia

## New OME proposal_Settings_v00-1
4DN-OME model
| Tier | Element | Attribute | Tier | Complex type | Description | Cardinality | Data type | Allowed values
IMAGE ACQUISITION Settings | 1 | Experiment |  | 1 |  | This element describes the type of experiment. The required Type attribute must contain one or more entries from the following list: FRET, Time-lapse, 4D+, HightThroughputScreen, FISH, Electrophysiology, Ion-Imaging, Co-localization, PGI/Documentation FRAP, Photo-ablation, Optical-Trapping, Photo-activation, Fluorescence-Lifetime, Spectral-Imaging, Other The required LabellingMethod attribute must contain one or more entries from the following list: FluorescentProtein, CASFISH, FISH, Single-Molecule-FISH, Indirect-Immunofluorescene, Immunocytochemistry, None, Other PGI/Documentation is not a 'data' image. The optional Description element may contain free text to further describe the experiment
| 1 | Experiment | ID | 1
| 1 | Experiment | Name | 1 |  | A user assigned name for this Experiment
| 2 | Experiment | Type | 2 | ExperimentType | A term to describe the type of experiment. |  | enum | FRET, TimeLapse, 4D+, HightThroughputScreen, Electrophysiology, IonImaging, Co-localization, PGI-Documentation, FRAP, PhotoAblation, OpticalTrapping, PhotoActivation, FLIM, SpectralImaging, SPIM, Other
| 1 | Experiment | LabellingMethod | 1 | ExperimentLabellingMethod | The method used for labelling the molecular component under study in this experiment |  | enum | FluorescentProtein, CASFISH, FISH, Single-Molecule-FISH, Immunofluorescene, Immunocytochemistry, None, Other
| 1 | Experiment | ExperimenterRef | 1 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME. | 1 … ∞ | Extension of Reference
| 1 | Experiment/Description | Experiment/Description | 1 |  | A description for the Experiment | 0 … ∞ | plain-text multi-line string
| 3 | Experiment | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 1 | Sample |  | 1 |  | A description of the sample that was imaged in this experiment | 1
| 1 | Sample | ID | 1
| 1 | Sample | Name | 1 |  | A user assigned name for this Sample |  | string
| 2 | Sample | MountingMedium | 2 |  | The mounting medium the sample is embedded in during acquisition |  | string
| 4 | Sample | MountingMediumRefractiveIndex | 4 |  | The measured refractive index of the medium the sample is embedded in during imaging |  | float with unit:MountingMediumRefractiveIndexUnit
| 1 | Sample | Sample/Description | 1 |  | A description for the Sample | 0 … ∞ | plain-text multi-line string
| 3 | Sample | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation |  | 3 |  | Defines a microbeam operation type and the region of the image it was applied to. The LightSourceSettingsRef element is a reference to a LightSourceSettings and in turn to the LightSource specified in the Instrument element which was used for this micromanipulation experiment. For example, a laser used for photobleaching. | 0 … ∞
| 3 | MicrobeamManipulation | ID | 3
| 3 | MicrobeamManipulation | Type | 3 | MicrobeamManipulationType | Defines the type of micromanipulation performed |  | enum | Fluorescence Recovery After Photobleaching, inverse FRAP, Fluorescence Loss In Photobleaching, Photo-Ablation, Photo-Activation, Uncaging, OpticalTrapping, Other
| 3 | MicrobeamManipulation | ExperimenterRef | 3 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME | 0 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | ROIRef | 3 |  | This empty element has a reference (the ROIID attribute) to a four dimensional ROI ( 'Region of Interest') that was used in this micromanipulation experiment | 1 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | LightSourceSettings | 3 |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 0 … ∞ | Element
| 3 | MicrobeamManipulation | MicrobeamManipulation/Description | 3 |  | A description for the Microbeam Manipulation experiment | 0 … ∞ | plain-text multi-line string
| 3 | MicrobeamManipulation | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 1 | Image |  | 1 |  | This element describes the actual image and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. The required attributes and elements are ID and Pixels. ExperimenterRef is required for all Images with well formed LSIDs. ImageType is a vendor-specific designation of the type of image this is. Examples of ImageType include 'STK', 'SoftWorx', etc. The Name attributes are in all cases the name of the element instance. In this case, the name of the image, not necessarily the filename. Physical size of pixels are microns[µm].
| 1 | Image |  | 1
| 1 | Image | Name | 1 |  | A user assigned name for this Image |  | string
| 1 | Image | AcquisitionDate | 1 |  | The acquisition date of the Image. |  | date
| 1 | Image (was Pixels) | SizeX | 1 |  | X size of pixel data array |  | integer with unit:none
| 1 | Image (was Pixels) | SizeY | 1 |  | Y size of pixel data array |  | integer with unit:none
| 1 | Image (was Pixels) | SizeZ | 1 |  | Number of Z-planes |  | integer with unit:none
| 1 | Image (was Pixels) | SizeC | 1 |  | Number of Channels used in this Image |  | integer with unit:none
| 4 | Image (was Pixels) | SizeT | 4 |  | Number of Timepoints used in this Image |  | integer with unit:none
| 1 | Image (was Pixels) | PhysicalSizeX | 1 |  | Physical width of a pixel |  | float with unitis:PhysicalSizeXUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeY | 1 |  | Physical height of a pixel |  | float with unitis:PhysicalSizeYUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeZ | 1 |  | Physical distance between optical planes |  | float with unitis:PhysicalSizeZUnit; default µm
| 4 | Image (was Pixels) | TimeIncrement | 4 |  | TimeIncrement is used for time series that have a global timing specification instead of per-timepoint timing info. For example in a video stream. Units are set by TimeIncrementUnit. |  | float with unit:TimeIncrementUnit
| 2 | Image (was Pixels) | DimensionOrder | 2 | PixelsDimensionOrder | The order in which the individual planes of data are interleaved in the Image |  | enum | XYZCT, XYZTC, XYCTZ, XYCZT, XYTCZ, XYTZC
| 2 | Image (was Pixels) | ChannelOrder | 2 |  | The order in which the channels were acquired and are stored |  | string
| 3 | Image (was Pixels) | PixelType | 3 | PixelType | The variable type used to represent the pixel intensity in the image |  | enum | int8, int16, int32, uint8, uint16, uint32, float, double, complex, double-complex, bit, Other
| 5 | Image (was Pixels) | SignificantBit | 5 |  | The number of bits within the type storing each pixel that are significant. e.g. you can store 12 bit data within a 16 bit type. This does not reduce the storage requirements but can be a useful indicator when processing or viewing the image data. |  | PositiveInteger
| 3 | Image (was Pixels) | Interleaved | 3 |  | How the channels are arranged within the data block |  | boolean | TRUE: if channels are stored RGBRGBRGB...; FALSE: if channels are stored RRR...GGG...BBB…
| 5 | Image (was Pixels) | BigEndian | 5 |  | This is true if the pixels data was written in BigEndian order. If this value is present it should match the value used in BinData or TiffData. If it does not a reader should honour the value used in the BinData or TiffData. This values is useful for MetadataOnly files and is to allow for future storage solutions. |  | boolean | TRUE: if pixel data was written in BigEndian mode
| 1 | Image | Image/Description | 1 |  | A description for the image | 0 … ∞ | plain-text multi-line string
| 3 | Image (was Pixels) | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 1 | Image | ExperimenterRef | 1 |  | This empty element has a required reference (the ExperimenterID attribute) to a Experimenter defined within OME. | 0 … ∞ | Extension of Reference
| 1 | Image | ExperimentGroupRef | 1 |  | This empty element has a reference (the ExperimenterGroup ID attribute) to a ExperimenterGroup defined within OME. | 0 … ∞ | Extension of Reference
| 1 | Image | ExperimentRef | 1 |  | This empty element has a reference (the Experiment ID attribute) to a Experiment defined within OME. | 1 | Extension of Reference
| 1 | Image | InstrumentRef | 1 |  | This empty element can be used (via the required Instrument ID attribute) to refer to an Instrument defined within OME | 1 | Extension of Reference
| 2 | Image | ROIRef | 2 |  | This empty element has a reference (the ROIID attribute) to a four dimensional ROI ( 'Region of Interest'). If they are not used, and the Image has more than one plane, the entire set of planes is assumed to be included in the ROI. Multiple ROIs may be specified. | 0 … ∞ | Extension of Reference
| 3 | Image | MicroBeamManipulationRef | 3 |  | This empty element has a reference (the MicroBeamManipulationID attribute) to a Microbeam Manipulation defined into OME. E.g. a microbeam operation, such as a laser used for photobleaching. | 0 … ∞ | Extension of Reference
| 4 | Image | ImagingEnvironment | 4 |  | This describes the environment that the biological sample was in, during the experiment. | 0 … ∞ | Extension of Reference
| 1 | Image | MicroscopeSettings | 1 |  | This holds the setting applied to a microscope as well as a reference to the microscope. The ID is the microscope used in this case. | 0 … ∞
| 3 | Image | TIRFSettings | 3 |  | Defines the TIRF settings used for this experiment | 0 … ∞
| 3 | Image | SamplePosition | 3 |  | This is used to specify a name and position for a sample position in the microscope's reference frame. | 0 … ∞
| 2 | Image | ObjectiveSettings | 2 |  | This holds the setting applied to an objective as well as a reference to the objective. The ID is the objective used in this case. | 0 … ∞
| 4 | ImagingEnvironment |  | 4 |  | This describes the environment that the biological sample was in during the experiment | 0 … ∞
| 4 | ImagingEnvironment | Temperature | 4 |  | The temperature the acquisition was performed at |  | float with unit:TemperatureUnit
| 4 | ImagingEnvironment | AirPressure | 4 |  | The airpressure the acquisition was performed at |  | float with unit:AirPressureUnit
| 4 | ImagingEnvironment | Humidity | 4 |  | The humidity the acqusition was performed at |  | float with unit:HumidityUnit
| 4 | ImagingEnvironment | CO2Percent | 4 |  | The CO2 concentration, expressed in percentage, the acquisition was performed at |  | float with unit:CO2PercentUnit
| 1 | MicroscopeSettings |  | 1 |  | This holds the setting applied to a Microscope as well as a reference to the Microscope (the ID attribute) in question. The ID referes to the microscope used in this case. | 0 … ∞
| 1 | MicroscopeSettings | ID (MicroscopeID) | 1
| 3 | MicroscopeSettings | FieldOfView | 3 |  | The Field of View defines the area of the specimen that is viewable via the eyepiece. This attribute caputers the measures size of the Field of View. |  | float with unit:FieldOfViewUnit[default mm]
| 1 | MicroscopeSettings | OptovarMagnification | 1 |  | This is the magnification associated with the Optovar used in this case |  | float with unit:none
| 1 | MicroscopeSettings | TotalEffectiveMagnification | 1 |  | This is the total effective magnification used in this case. This number can be obtained by multiplying the OptovarMagnification times the Objective/NominalMagnification or even better the Objective/CalibratedMagnification if available |  | float with unit:none
| 3 | TIRFSettings |  | 3 |  | Defines the TIRF settings used for this experiment | 0 … ∞
| 3 | TIRFSettings | ID | 3
| 3 | TIRFSettings | ThroughLens | 3 |  | The TIRF modality used in this case. True=Through the lens TIRF modality. False=TIRF modality in which the laser light is not travelling through the objective lens. |  | boolean
| 3 | TIRFSettings | TIRFAngle | 3 |  | The incidence angle used in this case |  | float with unit:TIRFAngleUnit
| 3 | TIRFSettings | Geometry | 3 |  | The TIRF geometry used in this case |  | enum | Spot, Ring, Other
| 3 | TIRFSettings/Description | TIRFSettings/Description | 3 |  | A description for the TIRF Settings | 0 … ∞ | plain-text multi-line string
| 3 | TIRFSettings | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | SamplePosition |  | 3 |  | SamplePosition is used to specify a name and location for a specific sample holder position in the microscope's reference frame. This element subsitutes a deprecated element called StageLabel | 0 … ∞
| 3 | SamplePosition | Name | 3 |  | A user assigned label for a specific sample holder position
| 3 | SamplePosition | SamplePositionX | 3 |  | The X position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionXUint
| 3 | SamplePosition | SamplePositionY | 3 |  | The Y position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionYUint
| 3 | SamplePosition | SamplePositionZ | 3 |  | The Z position of the SampleHolder within the microscope's reference frame |  | float with unit:PositionZUint
| 3 | SamplePosition | SampleHolderRef | 3 |  | This empty element has a reference (the SampleHolderID attribute) to a SampleHolder element defined within OME. | 1 … ∞ | Extension of Reference
| 2 | ObjectiveSettings |  | 2 |  | This holds the setting applied to an Objective e as well as a reference to the Objective (the ID attribute) in question. The ID referes to the Objective used in this case. | 0 … ∞
| 2 | ObjectiveSettings | ID (ObjectiveID) | 2
| 5 | ObjectiveSettings | CorrectionCollar | 5 |  | The CorrectionCollar is normally an adjustable ring on the objective. Each has an arbitrary scale on it so the values is unit-less. [units:none] |  | float with unit:none
| 2 | ObjectiveSettings | ImmersiumMedium | 2 | ObjectiveSettingsImmersionMedium | The actual immersion medium that was used in this case |  | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 2 | ObjectiveSettings | RefractiveIndex | 2 |  | The measured RefractiveIndex of the immersion medium utilized in this case. This is a ratio so it also unit-less. [units:none] |  | float with unit:none
| 4 | ObjectiveSettings | MeasuredRefractiveIndex | 4 |  | The temperature dependenent measured RefractiveIndex of ther immersion medium utilized in this case. This is a ratio so it also unit-less. [units:none] |  | float with unit:none
| 4 | ObjectiveSettings | Temperature | 4 |  | The temperature of the objective if different from ImagingEnvironment/Temperature |  | float with unit:ObjectiveTemperatureUnit
| 1 | Plane |  | 1 |  | The Plane object holds microscope stage and image timing data for a given channel/z-section/timepoint. | 1 … ∞
| 1 | Plane | ID | 1
| 1 | Plane | TheZ | 1 |  | The Z-section to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 1 | Plane | TheT | 1 |  | The timepoint to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 1 | Plane | TheC | 1 |  | The channel to which this plane corresponds. This is numbered from 0 |  | integer with unit:none
| 2 | Plane | Timestamp | 2 |  | This records the total time intercoursed since the beginning of the experiment |  | float with unit:TimestampUnit
| 2 | Plane | ExposureTime | 2 |  | The length of the exposure. |  | float with unit:ExposureTimeUnit.
| 5 | Plane | HashSHA1 | 5 |  | This optional element is a hash of the plane's image data. It is a choice between all the supported hash types. Currently the only method supported is SHA1 | 0 ... 1 | Hex40
| 3 | Plane | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | StageFocalPosition |  | 3 |  | This element is used to record the Z-axis position of a Plane with respect to the microscope's reference frame | 0 ... 1
| 3 | StageFocalPosition | ID | 3
| 3 | StageFocalPosition | Name | 3 |  | A user assigned name for this Stage Position
| 3 | StageFocalPosition | StagePositionZ | 3 |  | The Z-axis focal position
| 3 | StageFocalPosition | FocusingRef | 3 |  | This empty element has a reference (the FocusingID attribute) to a Focusing element defined within OME. | 0 ... 1
| 3 | StageFocalPosition | AutoFocusRef | 3 |  | This empty element has a reference (the AutofocusID attribute) to a Autofocus element defined within OME. | 0 ... 1
| 1 | Channel |  | 1 |  | There must be one per Channel in the Image, even for a single-plane image. Information about how each of the Channel was acquired is stored in the various optional *Ref elements. One useful feature of the model, is that this design allows to create LogicalChannels which are composed of one or more ChannelComponents. For example, an entire spectrum in an FTIR experiment may be stored in a single Logical Channel with each discrete wavenumber of the spectrum constituting a ChannelComponent of the FTIR Logical Channel. An RGB image where the Red, Green and Blue components do not reflect discrete probes but are instead the output of a color camera would be treated similarly - one Logical channel with three ChannelComponents in this case. The total number of ChannelComponents for a set of pixels must equal SizeC. The IlluminationType attribute is a string enumeration which may be set to 'Transmitted', 'Epifluorescence', 'Oblique', or 'NonLinear'. The user interface logic for labeling a given channel for the user should use the first existing attribute in the following sequence: Name -> Fluor -> EmissionWavelength -> ChannelComponent/Index. | 1 … ∞
| 1 | Channel | ID | 1
| 1 | Channel | Name | 1 |  | A name for the channel that is suitable for presentation to the user |  | string
| 5 | Channel | SamplesPerPixel | 5 |  | The number of samples the detector takes to form each pixel value. [units:none] Note: This is not the same as "Frame Averaging" - see Integration in DetectorSettings |  | PositiveInteger
| 2 | Channel | IlluminationType | 2 | ChannelIlluminationType | The method of illumination used to capture the channel |  | enum | Transmitted, Epifluorescence, Oblique, NonLinear, Other
| 3 | Channel | PinholeSize | 3 |  | The Pinhole Size utilized for this channel |  | float with unit:PinholeSizeUnit.
| 1 | Channel | AcquisitionMode | 1 | ChannelAcquistionMode | AcquisitionMode describes the type of microscopy performed for each channel |  | enum | WideField, LaserScanningConfocalMicroscopy, SpinningDiskConfocal, SlitScanConfocal, MultiPhotonMicroscopy, StructuredIllumination, SingleMoleculeImaging, TotalInternalReflection, FluorescenceLifetime, SpectralImaging, FluorescenceCorrelationSpectroscopy, NearFieldScanningOpticalMicroscopy, SecondHarmonicGenerationImaging, PALM, STORM, STED, TIRF, FSM, LCM, Other, BrightField, SweptFieldConfocal, SPIM,
| 1 | Channel | ContrastMethod | 1 | ChannelContrastMethod | ContrastMethod describes the technique used to achieve contrast for each channel |  | enum | Brightfield, Phase, DIC, HoffmanModulation, ObliqueIllumination, PolarizedLight, Darkfield, Fluorescence, Other
| 3 | Channel | ExcitationWavelength | 3 |  | The excitation wavelength used for this channel |  | float with unit:ExcitationWavelengthUnit
| 3 | Channel | EmissionWavelength | 3 |  | The emission wavelength used for this channel |  | float with unit:EmissionWavelengthUnit
| 2 | Channel | Fluorophore | 2 |  | This is the name of the fluorophore used to produce this channel [plain text string] |  | string
| 5 | Channel | NDFilter | 5 |  | The NDfilter attribute is used to specify the combined effect of any neutral density filters used. The amount of light the filter transmits at a maximum expressed as a fraction with value from 0.0 to 1.0. NOTE: This was formerly described as "units optical density expressed as a PercentFraction". This was how the field had been described in the schema from the beginning but all the use of it has been in the opposite direction, i.e. as a amount transmitted, not the amount blocked. This change has been made to make the model reflect this usage. |  | float with unit:none
| 5 | Channel | PockelCellSetting | 5 |  | The PockelCellSetting used for this channel. This is the amount the polarization of the beam is rotated by |  | integer with unit:none
| 1 | Channel | Color | 1 |  | A color used to render this channel - endocded as RGBA. The default value is "-1" corresponding #FFFFFFFF, so solid white |  | string
| 3 | Channel | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | LightSourceSettings |  | 3 |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 1
| 3 | LightSourceSettings | ID (LightSourceID) | 3
| 3 | LightSourceSettings | Attenuation | 3 |  | The Attenuation of the light source. A fraction, as a value from 0.0 to 1.0 |  | float with unit:none
| 3 | LightSourceSettings | Wavelength | 3 |  | The Wavelength of the light source |  | float with unit:WavelengthUnit
| 3 | LightSourceSettings | OutputPower | 3 |  | Light power measured in the back of the objective |  | float with unit:OutputPowerUnit
| 3 | LightSourceSettings | PowerAtSample | 3 |  | Light power measured at the sample interface |  | float with unit:PowerAtSampleUnit
| 2 | LightPath |  | 2 |  | A description of the light path | 1
| 2 | LightPath | ID | 2 |  |  |  | LSID
| 2 | LightPath | Name | 2 |  | Any user defined name/description for this element |  | string
| 3 | LightPath | LightSourceCouplingRef | 3 |  | This empty element can be used (via the required LightSourceCouplingID attribute) to refer to a LightSourceCoupling element defined within OME | 0 … ∞ | Extension of Reference
| 2 | LightPath | FilterSetRef | 2 | Reference | This empty element has a reference (the FilterSet ID attribute) to the FilterSet placed in the light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | ExcitationFilterRef | 2 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Excitation light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | DichroicRef | 2 | Reference | This empty element has a reference (the Mirror ID attribute) to the Mirror/Dichroic placed in the light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | EmissionFilterRef | 2 | Reference | This empty element has a reference (the Filter ID attribute) to the Filters placed in the Emission light path | 0 … ∞ | Extension of Reference
| 2 | LightPath | AdditionalOptics | 2 |  | A description of additional optics placed in the light path | 0 … ∞ | Element
| 3 | LightPath | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics |  | 2 |  | A description of additional optics placed in the light path | 0 … ∞
| 2 | AdditionalOptics | ID | 2
| 2 | AdditionalOptics | Name | 2 |  | Any user defined name/description for this element |  | string
| 3 | AdditionalOptics | AdditionalLensRef | 3 | Reference | This empty element has a reference (the LensID attribute) to the an additional Lens placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | PrimsRef | 3 | Reference | This empty element has a reference (the PrismID attribute) to the a Prism placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | PolarizationOpticsRef | 3 | Reference | This empty element has a reference (the PorarizationOpticsID attribute) to the a PolarizationOptics placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AdditionalFilterRef | 3 | Reference | This empty element has a reference (the Filter ID attribute) to the an additional Filter placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AdditionalMirrorRef | 3 | Reference | This empty element has a reference (the MirrorID attribute) to the an additional Mirror placed in the light path | 0 … ∞ | Extension of Reference
| 3 | AdditionalOptics | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞
| 3 | DetectorSettings-APE |  | 3 |  | This holds the setting applied to a Detector e as well as a reference to the Detector (the ID attribute) in question. The ID referes to the Detector used in this case | 1
| 3 | DetectorSettings-APE | ID (DetetectorID) | 3 |  |  |  | DetectorID
| 3 | DetectorSettings-APE | EffectiveOffset | 3 |  | This is the actual offset used for this image. The offset is an electronic baseline added to the output signal from the EMCCD sensor to ensure that the displayed signal level is always a positive number of counts. |  | float with unit:none
| 3 | DetectorSettings-APE | Integration | 3 |  | This is the number of sequential frames that have been averaged for this image, to improve the signal-to-noise ratio |  | integer with unit:none
| 3 | DetectorSettings-APE | CameraFieldOfView | 3 |  | The Camera Field of View is the diameter of the field visible by the camera as measured in the specimen plane. It is expressed in mm and it is defined as the Cameara Field Number divided by the Total Effective Magnification of the system. In turn the Field Number is defined as the diameter of the view field in mm measured at the intermediate image plane. |  | float with unit:CameraFieldOfViewUnit[default mm]
| 3 | CameraSettings |  | 3 |  | This holds the setting applied to a Camera e as well as a reference to the Camera (the ID attribute) in question. The ID referes to the Camera used in this case.
| 3 | CameraSettings | Binning | 3 | Binning | Represents the number of pixels that are combined to improve readout rates, signal to noise ratios albeit at the expense of reduced spatial resolution |  | enum | 1x1, 2x2, 4x4, 8x8, Other
| 3 | CameraSettings | AnalogToDigitalConverter | 3 |  | The bit depth used to report the light signal in digital units. |  | integer with unit:bit
| 3 | CameraSettings | EMGain | 3 |  | Gain applied to the detector signal. This is the electronic gain (as apposed to the inherent gain) that is set for the detector for this image |  | integer with unit:none
| 5 | CameraSettings | ReadoutRate | 5 |  | Effective pixel read out rate used for this image |  | float with unit:ReadOutRateUnit
| 3 | CameraSettings | FrameRate | 3 |  | Effective frame rate used for this image |  | float with unit:FrameRateUnit
| 3 | CameraSettings | OperatingTemperature | 3 |  | Effective cooling temperature used for this image |  | float with unit:TemperatureUnit; default:Celsius[°C]
| 3 | PhotoMultiplierSettings |  | 3 |  | This holds the setting applied to a PhotoMultiplier e as well as a reference to the PhotoMultiplier (the ID attribute) in question. The ID referes to the PhotoMultiplier used in this case.
| 3 | PhotoMultiplierSettings | PMTGain | 3 |  | In general, the gain is a measure of the ability a two-port circuit to add energy to an input signal and therefore increase its power or amplitude. Specifically, PMTGain refers to the ratio of output to input current set for the acquisition of this image |  | float with unit:none
| 3 | PhotoMultiplierSettings | PMTVoltage | 3 |  | This refers to the voltage applied to the photomultiplier for the acquisition of this image |  | float with unit:PMTVoltageUnit
| 3 | PhotoMultiplierSettings | EffectiveZoom | 3 |  | The Zoom or "Confocal Zoom" or "Scan Zoom" utilized for this image |  | foat with unit:none
| 3 | OpticalCalibration |  | 3 |  | This describes the procedure that was used for optical calibration and the resulting calibration results |  |  |  | Alternative name: PSFCalibration
| 3 | OpticalCalibration | ID | 3
| 3 | OpticalCalibration | Date | 3 |  | The date in which this calibration was performed.
| 3 | OpticalCalibration | OpticalCalibrationSoftware | 3 |  | The name of the software tool utilized to generate the OpticalCalibration reoprt. Such a tool generally analyzes microscopy images of fluorescent beads in order to determine quantitatively the resolution and chromatic aberrations of a fluorescence microsocope over the full field of view. An example of such a tool is PSFj (http://www.knoplab.de/psfj/sample-page/)
| 3 | OpticalCalibration | FWHMmedianBeadSize | 3 |  | Observed median bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile | Optional | float with unit:
| 3 | OpticalCalibration | FWHMmedianTheoretical | 3 |  | Theoretical (expected) median bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile, on the basis of the given optical conditions. | Optional | float with unit:
| 3 | OpticalCalibration | FWHMminBeadSize | 3 |  | Observed minimum bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile | Optional | float with unit:
| 3 | OpticalCalibration | FWHMminTheoretical | 3 |  | Theoretical (expected) minimum bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile, on the basis of the given optical conditions. | Optional | float with unit:
| 3 | OpticalCalibration | FWHMmaxBeadSize | 3 |  | Observed maximum bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile | Optional | float with unit:
| 3 | OpticalCalibration | FWHMmaxTheoretical | 3 |  | Theoretical (expected) maximum bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile, on the basis of the given optical conditions. | Optional | float with unit:
| 3 | OpticalCalibration | FWHMzBeadSize | 3 |  | Observed axial bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile | Optional | float with unit:
| 3 | OpticalCalibration | FWHMzTheoretical | 3 |  | Theoretical (expected) axial bead size expressed as the Full Width at Half Maximum (FWHM) of the fitted Bead Spread Function (BSF) intensity profile, on the basis of the given optical conditions. | Optional | float with unit:
| 3 | OpticalCalibration | Planarity | 3 |  |  | Optional | float with unit:
| 3 | OpticalCalibration | Asymmetry | 3 |  | The average over the Field of View (FOV) of the observed Bead Spread Function (BSF) asymmetry, calculated as the ratio of the smallest over the largest FWHM of the fitted intensity profile. | Optional | float with unit:none
| 3 | OpticalCalibration | AsymmetryWeightedTheta | 3 |  | The average over the Field of View (FOV) of the observed Bead Spread Function (BSF) asymmetry, calculated as the smallest over the largest FWHM of the fitted PSF intensity profile. | Optional | float with unit:none
| 3 | OpticalCalibration | LateralChromaticShift | 3 |  | Lateral offset between two color channels | Optional | float with unit:
| 3 | OpticalCalibration | AxialChromaticShift | 3 |  | Axial offset between two color channels | Optional | float with unit:
| 3 | OpticalCalibration | Description | 3 |  | A pertinent description of the calibration procedure | 0 … ∞ | string
| 3 | OpticalCalibration | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | ColoredBeads |  | 3 |  | This element describes the beads that were used to acquire the image | 1 | Element
| 3 | ColoredBeads | Manufacturer | 3 | Manufacturer Spect | The manufacturer of the component |  | string
| 3 | ColoredBeads | Model | 3 | Manufacturer Spect | The model of the component |  | string
| 3 | ColoredBeads | SerialNumber | 3 | Manufacturer Spect | The serial number of the component |  | string
| 3 | ColoredBeads | LotNumber | 3 | Manufacturer Spect | The lot number of the component |  | string
| 3 | ColoredBeads | SpecsFile | 3 |  | Attach a PDF file with the specifications for the component as provided by the manufacturer |  | file
| 3 | ColoredBeads | ID | 3
| 3 | ColoredBeads | Name | 3 |  | A user assigned name for this beads preparation |  | Denomination
| 3 | ColoredBeads | BeadSize | 3 |  | The size of the colred beads utilized | float with unit:BeadSizeUnit
| 3 | ColoredBeads | BeadConcentration | 3 |  | The concentration of the colored beads utilized | float with unit:BeadConcentration
| 3 | ColoredBeads | BeadColor (array?) | 3 |  | The color(s) emitted by these beads | array of strings
| 3 | ColoredBeads | Type | 3 |  | The type of colred beads used in this case |  | string
| 3 | ColoredBeads | Description | 3 |  | A pertinent description of the calibration procedure | 0 … ∞ | string
| 3 | ColoredBeads | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | ColoredBeadsImageRef |  | 3 |  | An image stack of colored beads imaged using the same instrument and settings employed to acquire the image this calibration refers to. | 0 … ∞ | Image
| 3 | ColoredBeadsImageRef | ID (ImageID) | 3
| 3 | OpticalCalibrationReportRef |  | 3 |  | Colored beads imaging report generated by the OpticalCalibrationSoftware referred to in OpticalCalibration. | 0 … ∞ | FileURI
| 3 | OpticalCalibrationReportRef | ID (OpticalCalibrationReportID) | 3 | [5 empty] | Make this name more general PSFCalibrationReport
| 3 | IntensityCalibration-APE |  | 3 |  | This describes the procedures that were used for field, excitation and detector calibration and the resulting calibration results |  |  |  | This is for Field, Excitation and Detector calibration
| 3 | IntenstiyCalibration-APE | ID | 3
| 3 | IntenstiyCalibration-APE | Date | 3 |  | The date in which this calibration was performed.
| 3 | IntenstiyCalibration-APE | Description | 3 |  | A pertinent description of the calibration procedure | 0 … ∞ | string
| 3 | IntenstiyCalibration-APE | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | IntensityCalibrationToolSettings |  | 3 | [5 empty] | This is the abstracted/generic name for MetaMax
| 3 | IntensityCalibrationToolSettings | ID (CalibrationToolID) | 3
| 3 | IntensityCalibrationToolSettings | Settings_1 | 3
| 3 | IntensityCalibrationToolSettings | Settings_2 | 3 | [5 empty] | Do we need these?
| 3 | IntensityCalibrationToolSettings | Settings_3 | 3
| 3 | IntensityCalibrationToolSettings | AnnotationRef | 3 | CommentAnnotation | The AnnotationRef element is a reference to an element derived from the CommonAnnotation element (I.e. A simple comment annotation) | 0 … ∞ | Extension of Reference
| 3 | FieldCalibration | FlatFieldImageRef | 3
| 3 | FlatFiedlImageRef |  | 3 |  | An image describing field dependent differences in response | 0 … ∞ | Image
| 3 | FlatFiedlImageRef | ID (ImageID) | 3
| 3 | ExcitationCalibration | ObservedExcitationWavelength | 3 |  | Center of the observed excitation wavelength
| 3 | ExcitationCalibration | ObservedExcitationPower | 3 |  | Observed excitation light irradiance as measured at the sample plane
| 3 | DetectorCalibration | DetectorReadMap | 3
| 3 | DetectorCalibration | DetectorConversionMap | 3
| 3 | DetectorReadMap |  | 3 |  | A multi-dimensional map describing the pixel by pixel variations in detected intensity and and intensity variance obtained under increasing light exposure intensities starting from the absence of light and using a suitable intensity calibration apparatus. |  |  |  | a 4 dimensional map/matrix reporting intensity µ and variance per pixel obtained under dark + different illumination intensities
| 3 | DetectorReadMap | ID (DetectorReadMapI) | 3 |  |  |  | FileURI
| 3 | DetectorConversionMap |  | 3 |  | A two dimensional map describing the observed pixel by pixel photometric conversion factor used to convert incoming counts to photons |  |  |  | this is the output of processing of the IlluminationMap. Reflects the noise model and reflects what the producer of the image did to produce a conversion factor
| 3 | DetectorConversionMap | ID (DetectorConversionMap) | 3 |  |  |  | FileURI
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on E2: What is the meaning of Tier5 as far as Settings? It seems to me that Tier5 as it is defined now only applies to HARDWARE but not as much to SETTINGS… -Caterina Strambio De Castillia
Note on C72: This seems too specific. Should be generalized to include other similar experiments
Note on D155: HOW DOES THIS GET MEASURED? -Caterina Strambio De Castillia

## Element categories_v00-01
Category | Sub-category | Element
MicroscopicBasic |  | Instrument, MicroscopyBody, MicroscopyTable
SampleHolderAndFocusing |  | SampleHolder, AutoFocus
SampleHolderAndFocusing | SampleHolderAndFocusing.Focusing | MechanicalStage, PiezoElectricStage, PiezoElectricGroup
Magnification |  | Objective, Optovar
LightSource |  | Laser, Arc, Filament, LightEmittingDiode, GenericExcitationSource
LightSourceCoupling |  | LightGuide, Direct
LightPath |  | LightPath, FilterSet, Filter
Mirror |  | Reflector, Beamsplitter, Dichroic
AdditionalOptics |  | Lens, Prism, PolarizationOptics, AdditionalFilter, AdditionalMirror
Detector |  | PhotoMultiplier, GenricDetector
Detector | Detector.Camera | CCD, IntensifiedCCD, AnalogVideo, CMOS