## Tier system_v01-00
-- required metadata --
Category | ID | Name | Description | Example | Example Labelling | Optical calibration | Intensity calibration | Mechanical calibration | Experiment/ Sample | Microscope specifications | Image acquisition settings
Descriptive | 1 | Documentation | Reporting qualitative effects without quantification | Transfection control, viability assay | FISH, Immuno Fluorescence | recommended annually (not required) | not recommended | not recommended | - experiment description and date - sample description - experimenter name | - microscope manufacturer, name and type - light source type - objective magnification and NA | - acquisition date - dimension size - physical pixel size x, y, and z - optovar magnification - channel name, color, contrast method and acquisition mode
| 2 | Simple Quantification | Identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required) | not recommended | not recommended | - mounting medium | - light source manufacturer, wavelenght - objective manufacturer and correction - filter/dichroic transmittance range - detector type, manufacturer | - exposure time - immersion medium name and refractive index - image dimension order, channel order - channel illumination type, fluorophore
Analytical | 3 | Advanced Quantification | Identification and localization of refraction-limited particles, super-resolution microscopy | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing | SM FISH, CasFISH, Proximity Ligation Assay (PLA), SM-FP | required quarterly | recommended | recommended |  | detailed light source, magnification, focusing and autofocus, light path, filter, dichroic, additional optics and detector specification & settings
| 4 | Life Cells Imaging | Tracking of intracellular dynamics | Cell tracking, single particle tracking, dynamic expression level quantification | dCas9-based labelling, fluoresent protein labelling, SM-FP | required monthly | highly-recommended | highly-recommended | - imaging temperature, air pressure, humidity and CO2 percentage - mounting medium measured refractive index | - microscope table and light source coupling specifications | - acquisition time settings - sample positioning settings - objective temperature - measured immersion medium refractive index
| 5 | Pioneer | Full reproducibility of microscopic set up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition |  |  | all the metadata specified by the data model - including any novel technology-specific metrics

## Tier system_v01-00_FORMATTED
-- required metadata --
Category | ID | Name | Description | Example | Example Labelling | Optical calibration | Intensity calibration | Mechanical calibration | Experiment/ Sample | Microscope specifications | Image acquisition settings
Descriptive | 1 | Documentation | Reporting qualitative effects without quantification | Transfection control, viability assay | FISH, Immuno Fluorescence | recommended annually (not required) | not recommended | not recommended | - experiment description and date - sample description - experimenter name | - microscope manufacturer, name and type - light source type - objective magnification and NA | - acquisition date - dimension size - physical pixel size x, y, and z - optovar magnification - channel name, color, contrast method and acquisition mode
| 2 | Simple Quantification | Identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required) | not recommended | not recommended | - mounting medium | - light source manufacturer, wavelenght - objective manufacturer and correction - filter/dichroic transmittance range - detector type, manufacturer | - exposure time - immersion medium name and refractive index - image dimension order, channel order - channel illumination type, fluorophore
Analytical | 3 | Advanced Quantification | Identification and localization of refraction-limited particles, super-resolution microscopy | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing | SM FISH, CasFISH, Proximity Ligation Assay (PLA), SM-FP | required quarterly | recommended | recommended |  | detailed light source, magnification, focusing and autofocus, light path, filter, dichroic, additional optics and detector specification & settings
| 4 | Life Cells Imaging | Tracking of intracellular dynamics | Cell tracking, single particle tracking, dynamic expression level quantification | dCas9-based labelling, fluoresent protein labelling, SM-FP | required monthly | highly-recommended | highly-recommended | - imaging temperature, air pressure, humidity and CO2 percentage - mounting medium measured refractive index | - microscope table and light source coupling specifications | - acquisition time settings - sample positioning settings - objective temperature - measured immersion medium refractive index
| 5 | Pioneer | Full reproducibility of microscopic set up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition |  |  | all the metadata specified by the data model - including any novel technology-specific metrics

## Tier system_v01-00_SUMMARY
Category | ID | Name | Description | Example | Example Labelling | Optical calibration | Intensity calibration | Mechanical calibration
Descriptive | 1 | Documentation | Reporting qualitative effects without quantification | Transfection control, viability assay | FISH, Immuno Fluorescence | recommended annually (not required) | not recommended | not recommended
| 2 | Simple Quantification | Identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required) | not recommended | not recommended
Analytical | 3 | Advanced Quantification | Identification and localization of refraction-limited particles, super-resolution microscopy | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing | SM FISH, CasFISH, Proximity Ligation Assay (PLA), SM-FP | required quarterly | recommended | recommended
| 4 | Life Cells Imaging | Tracking of intracellular dynamics | Cell tracking, single particle tracking, dynamic expression level quantification | dCas9-based labelling, fluoresent protein labelling, SM-FP | required monthly | highly-recommended | highly-recommended
| 5 | Pioneer | Full reproducibility of microscopic set up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition

## Tier system_v01-00_MINIMAL
Category | ID | Name | Description | Example | Labelling | Quality control
Descriptive | 1 | Documentation | Reporting qualitative effects without quantification | Transfection control, viability assay | FISH, Immuno Fluorescence | not required
| 2 | Simple Quantification | Identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly
Analytical | 3 | Advanced Quantification | Identification and localization of refraction-limited particles, super-resolution microscopy | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing | SM FISH, CasFISH, Proximity Ligation Assay (PLA), SM-FP | required quarterly
| 4 | Life Cells Imaging | Tracking of intracellular dynamics | Cell tracking, single particle tracking, dynamic expression level quantification | dCas9-based labelling, fluoresent protein labelling, SM-FP | required monthly
| 5 | Pioneer | Full reproducibility of microscopic set up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition