## Element categories_v01-00
Nr. | Extension | Domain | Order in Menu | Category | Sub-category | Element
1 | Core | MicrosocpeHardwareSpecifications | 1 | MicroscopyEssentials |  | Instrument, MicroscopeStand
2 | Core | MicrosocpeHardwareSpecifications | 2 | LightSource |  | Arc, Filament, GenericExcitationSource, Laser, LightEmittingDiode
3 | Core | MicrosocpeHardwareSpecifications | 4 | Magnification |  | Objective
4 | Core | MicrosocpeHardwareSpecifications | 5 | FluorescenceLightPath |  | FilterCube, ExcitationFilter, EmissionFilter, StandardDichroic
5 | Core | MicrosocpeHardwareSpecifications | 14 | Detector |  | GenericDetector
6 | Core | MicrosocpeHardwareSpecifications | - | ChildElement |  | TransmittanceRange
7 | Core | Experimental | - | SamplePreparation |  | Experiment, Experimenter, ExperimenterGroup, Sample, SampleHolder, CoverGlass, ImmersionLiquid, Fluorophore
8 | Core | ImageAcquisitionSettings | - | Channel |  | Channel
9 | Core | ImageAcquisitionSettings | - | LightPath |  | LightPath, FilterSet
10 | Core | ImageAcquisitionSettings | - | Image |  | Image, Pixels, Plane
11 | Core | ImageAcquisitionSettings | - | SampleConditions |  | ImagingEnvironment
12 | Core | ImageAcquisitionSettings | - | Settings |  | DetectorSettings, LightSourceSettings, ObjectiveSettings
13 | Basic | MicrosocpeHardwareSpecifications | 1 | MicroscopyEssentials |  | AcquisitionSoftware, MicroscopeTable, EnvironmentalControlDevice
14 | Basic | MicrosocpeHardwareSpecifications | 2 | LightSource |  | MultiLasrEngine
15 | Basic | MicrosocpeHardwareSpecifications | 3 | LightSourceCoupling |  | LightGuide, FreeBeam
16 | Basic | MicrosocpeHardwareSpecifications | 4 | Magnification |  | MagnificationChanger
17 | Basic | MicrosocpeHardwareSpecifications | 5 | SamplePositioning | SamplePositioning.Stage | MechanicalStage, PiezoElectricStage, StageInsert
18 | Basic | MicrosocpeHardwareSpecifications | 6 | SamplePositioning | SamplePositioning.Focusing | FocusStabilizationDevice, IndividualObjectiveFocusing, Z-Drive, ObjectiveTurretFocusing
19 | Basic | MicrosocpeHardwareSpecifications | 7 | OpticalAssemblies |  | BeamExpander, Collimator, Condenser, MicroscopeTube
20 | Basic | MicrosocpeHardwareSpecifications | 8 | Apertures |  | DarkFieldStop, HoffmanPlate, IrisDiaphragm, PhaseRing, VarelRing, Shutter
21 | Basic | MicrosocpeHardwareSpecifications | 9 | Holders |  | FilterSlider, FilterTurret, FilterWheel, OpticsTurret
22 | Basic | MicrosocpeHardwareSpecifications | 10 | Filters |  | GenericFilter, NeutralDensityFilter
23 | Basic | MicrosocpeHardwareSpecifications | 11 | Mirrors |  | GenericDichroic, ReflectingMirror, Beamsplitter,
24 | Basic | MicrosocpeHardwareSpecifications | 12 | Lenses |  | GenericLens, BeamExpanderLens, BeamCollimatorLens, CouplingLens, RelayLens, TubeLens
25 | Basic | MicrosocpeHardwareSpecifications | 13 | AdditionalOptics |  | PolarizationOptics, Prism
26 | Basic | MicrosocpeHardwareSpecifications | 15 | Detector | Detector.PointDetector | HybridPhotoDetector, PhotDiode, PhotoMultiplierTube
27 | Basic | MicrosocpeHardwareSpecifications | 16 | Detector | Detector.Camera | AnalogVideo, CCD, CMOS, IntensifiedCCD
28 | Basic | MicrosocpeHardwareSpecifications | - | VendorInformation |  | AdditionalVendorInformation, VendorBinData
29 | Basic | MicrosocpeHardwareSpecifications | - | ChildElement |  | LEDModule, Map, Pump, ReflectanceRance, TransmittanceRange, WavelengthRange, IlluminationWavelengthRange
30 | Basic | Experimental | - | SamplePreparation |  | Sample, SampleHolder, CoverGlass, ImmersionLiquid, Fluorophore
31 | Basic | ImageAcquisitionSettings | - | Settings |  | AdditionalMirrorSettings, CameraSettings, MicroscopeSettings, MicroscopeTableSettings, PhotoMultiplierSettings, PrismSettings, ShutterSettings,
32 | Basic | ImageAcquisitionSettings | - | ChildElement |  | LightPathMap, OpticalComponentPosition, ReflectionWavelengthRangeSettings, TransmissionWavelengthRangeSettings, PlaneTranformationMatrix
33 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 8 | Apertures |  | DiffractionGrating, MaskingPlate, Pinhole
34 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 12 | Lenses |  | MicrolensArray
35 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 16 | Detector | Detector.Camera | WavefrontSensor
36 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 17 | Confocal |  | RasterScanner, SpinningDisk
37 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 18 | LightModulation |  | AcoustoOpticalBeamSplitter, AcoustoOpticalDeflector, AcoustoOpticalModulator, AcoustoOpticalTuneableFilter, DeformableMirror, ElectroOpticalModulator, LiquidCrystalTunableFilter, PockelsCell, SpatialLightModulator
38 | Advanced+Confocal | MicrosocpeHardwareSpecifications | 19 | ModalitySpecific |  | TIRFHardwareModule
39 | Advanced+Confocal | Experimental | - | BeamManipulation |  | MicrobeamManipulation
40 | Advanced+Confocal | ImageAcquisitionSettings | - | Settings |  | TIRFSettings, AcoustoOpticalTuneableFilterSettings, AcoustoOpticalBeamSplitterSettings, AcoustoOpticalDeflectorSettings, AcoustoOpticalModulatorSettings, ConfocalScannerSettings, DiffractiveGratingSettings, MaskingPlateSettings, PinHoleSettings, DeformableMirrorSettings, ElectroOpticalModulatorSettings, LiquidCrystalTuneableFilterSettings, SpatialLightModulatorSettings
41 | Calibration and Performance | MicrosocpeHardwareSpecifications | 20 | CalibrationTools |  | IntensityCalibrationTool
42 | Calibration and Performance | ImageAcquisitionSettings | - | CalibrationSamples |  | CalibrationSlide, ColoredBeadsSlide, DNAOrigami, FluorescentReferenceSlide, MultiColorMicrospheres
43 | Calibration and Performance | ImageAcquisitionSettings | - | CalibrationProcedures |  | CalibrationMap, CalibrationSoftware, ChromaticRegistrationEvaluation, DetectorCalibration, FieldCalibration, MechanicalCalibration, OpticalCalibration,
44 | Calibration and Performance | ImageAcquisitionSettings | - | ChildElement |  | PSFEquationFile, OpticalCalibrationReport, Cell, Row, IntensityCalibrationReport

## 4DN-BINA OME-Basic Ext_Spc_v02
4DN-OME Model |  |  | APE = Abstract Parent Element
[15 empty] | BINA WORKING GROUP | [7 empty] | Methods checklist |  |  | QUAREP comments
| Tier | M&M (Montero, Brown, Cameron, Itano and Strambio-De-Castillia) | Element | Attribute or Subelement | Complex type | Description | Data type | Allowed values | Cardinality/Required? | OME 6-2016 | OME Cardinality/Required? | OME Data type | SPACER | DECISION | Damir Comments | Pablo Comments | Arturo comments | Caterina comments | Jim comments | Nathalie comments | Ulrike comments | SPACER2 | Michelle Itano | Lisa Cameron | Paola Motero Llopin | WG7 Member 1 [Enter your name and Choose a color] | WG7 Member 2 [Enter your name and Choose a color] | WG7 Member n [Enter your name and Choose a color]
Microscope Speecifications | 1 |  | Instrument |  |  | Y | - |  | 1 | same | - | - |  | MADE Instrument EXTEND MicroSpecs | Does the Instrument have a ManufacturerSpec as well?
| 1 |  | Instrument | ID |  | A Unique Identifier for this component. | LSID |  | R | same |  | LSID | [6 empty] | I love this! Need to find it, but another org was trying to do this for core facilities....
| 1 |  | Instrument | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞ | same | 0... ∞ | Extension of Reference |  | CHANGE CARDINALITY FOR ALL AnnotationRef | Allow 0 ... infinity cardinality?
| 1 |  | Instrument | Name |  | A User-defined Name for this component. | Denomination |  | R
| 1 | Y | ManufacturerSpecs (extension base) |  | ExtensionBase | This describes the Manufacturer specifications for this component (i.e., Manufacturer Name, Model designation and Catalog Number). | ExtensionBase |  | - | same | - | -
| 1 | Y | ManufacturerSpecs (extension base) | Manufacturer | Manufacturer Specs | This field refers to the company, research group or individual who manufactured this hardware component. | string |  | R | same |  | string
| 1 | Y | ManufacturerSpecs (extension base) | Model | Manufacturer Specs | This field refers to the Model name or any other designation the Manufacturer uses for this component. | string |  | R | same |  | string
| 1 | Y | ManufacturerSpecs (extension base) | CatalogNumber | Manufacturer Specs | This field records the Catalog, Part or Serial Number used by the Manufacturer to uniquely identify this component. | string |  | R | was Serial Number |  | string |  | Changed to Not required (applies to all Manufacturer Specs)
| 3 |  | ManufacturerSpecs (extension base) | LotNumber | Manufacturer Specs | The Lot Number used by the Manufacturer to refer to a specific production batch or lot for this component. | string |  |  | same |  | string |  | Changed to Not required (applies to all Manufacturer Specs)
| 3 |  | ManufacturerSpecs (extension base) | SpecsFile | FileAnnotation | This element refers to a PDF file that contains the specifications for this hardware component as provided by the Manufacturer. | Element |  | 0 ...∞ |  |  |  |  | Changed to Not required (applies to all Manufacturer Specs) |  |  |  |  | Will this just describe Microscope body or entire system as installed?
| 3 |  | ManufacturerSpecs (extension base) | MapAnnotation |  | A Map Annotation. The contents of this is a list of Key/Value pairs. | Element |  | 0 …∞ |  |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | ManufacturerSpecs (extension base) | AdditionalVendorInformation |  | This element is intendended to be used as a last-resort solution to capture vendor-specific Additional Information about this hardware component, which cannot be described using either available model elements (Classes or Attributes) or the Map Annotation feature of the schema. The use of this feature should not be intended as a way to hide information that the Imaging Community has designated as essential to ensure rigor and reproducibility. In case this feature is used, the Vendor is fully responsible for this information, including but not limited to, making sure this information is appropriately recorded and kept up to date. | Element |  | 0 …∞ |  |  |  |  | ADDED. This AdditionalVendorInformation element is going to be a subelement of ManufacturerSpec which is the Extension base for all Microscope components described in this Model.CHANGED DESCRIPTION TO MAKE SURE IT IS CLEAR THAT THE USE OF THIS SHOULD BE LAST RESORT | Will that also extend to the Settings entries? Vendors will also want to have a place for this type of stuff there. | [5 empty] | I second Damir
| 3 |  | MapAnnotation |  | Annotation | A Map Annotation. The contents of this is a list of Key/Value pairs. |  |  | 0 ... ∞ | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | ID | AnnotationID | A Unique Identifier for this component. | LSID |  | R | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | Namespace |  | We recommend the inclusion of a Namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type. | anyURI |  | R | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | Annotator (ExperimenterID) | ExperimenterID | The Annotator is the person who attached this annotation.e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator. | ExperimenterID |  |  | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | Description |  | This element contains a description for the annotation.[plain-text multi-line string] | Element |  | 0 … 1 | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | Value |  | This is a Mapping of Key/Value pairs. | Element |  | 1 | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | MapAnnotation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞ | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | Value |  |  | This is a Mapping of Key/Value pairs. |  |  |  | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | Value | M |  | This is a Key/Value pair used to build up a Mapping. The Element and Attribute Name are kept to single letters to minimize the length at the expense of readability as they are likely to occur many times. | Element |  | 0 … ∞ | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | M |  |  | This is a Key/Value pair used to build up a Mapping. The Element and Attribute Name are kept to single letters to minimize the length at the expense of readability as they are likely to occur many times. |  |  |  | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | M | K |  | This is the key of this Key/Value pair. | string |  | R | same (minor description changes) |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 |  | AdditionalVendorInformation |  |  | This element is intendended to be used as a last-resort solution to capture vendor-specific Additional Information about this hardware component, which cannot be described using either available model elements (Classes or Attributes) or the Map Annotation feature of the schema. The use of this feature should not be intended as a way to hide information that the Imaging Community has designated as essential to ensure rigor and reproducibility. In case this feature is used, the Vendor is fully responsible for this information, including but not limited to, making sure this information is appropriately recorded and kept up to date. | [7 empty] | ADDED | Add to description that Manufacturer/vendor is encouraged to use the MapAnnotation option instead since that can be easily interpreted by the community and provide incentive to adopt into the standard. | [5 empty] | But core components have to be specified; i.e. a vendor standard has to be defined
| 3 |  | AdditionalVendorInformation | ID |  | A Unique Identifier for this model element. | LSID |  | R |  |  |  |  | ADDED
| 3 |  | AdditionalVendorInformation | VendorName | AdditionalVendorInformation | The Name of the vendor responsible for this Information | string | [6 empty] | ADDED
| 3 |  | AdditionalVendorInformation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED | not sure how this can be codified in ER diagram
| 3 |  | AdditionalVendorInformation | VendorBinData | AdditionalVendorInformation | The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream. | Element |  | 1 ... ∞ |  |  |  |  | ADDED |  |  |  | Propose to ADD
| 3 |  | VendorBinData |  | base64Binary | The contents of this element are base64-encoded. These are not CDATA sections, just base64 streams. |  |  | 1 ... ∞ |  |  |  |  | ADDED |  |  |  | Propose to ADD
| 3 |  | VendorBinData | Compression |  | Specifies the compression scheme used to encode this data. | string | [6 empty] | ADDED |  |  |  | Propose to ADD
| 3 |  | VendorBinData | BigEndian |  | This is true if the binary data was written in BigEndian order. This is dependent on the system architecture of the machine that wrote the pixels. True for essentially all modern CPUs other than Intel and Alpha. All Binary data must be written in the same endian order. | boolean |  | R
| 3 |  | VendorBinData | Length |  | Character count attribute for the BinData field. This is the length of the base-64 encoded block. It allows easy skipping of the block when parsing the file. [unit:bytes] | long with unit:bytes |  | R |  |  |  |  | ADDED |  |  |  | Propose to ADD
| 1 |  | CommentAnnotation |  |  | An annotation that consists of a simple multi-line comment. | - |  | - | same | - | -
| 1 |  | CommentAnnotation | ID |  | A Unique Identifier for this component. | AnnotationID |  | R | same | R | AnnotationID
| 3 |  | CommentAnnotation | Namespace |  | We recommend the inclusion of a Namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type. | anyURI |  |  | same |  | anyURI
| 1 |  | CommentAnnotation | Annotator |  | The Annotator is the person who attached this annotation. e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator. | ExperimenterID |  |  | same |  | ExperimenterID
| 1 |  | CommentAnnotation/AnnotationRef | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0... ∞ | same | 0... ∞ | Extension of Reference
| 1 |  | CommentAnnotation/Description | Description |  | This element contains a description for the annotation.[plain-text multi-line string] | Element |  | 0 ... 1 | same | 0 ... 1 | Element
| 1 |  | CommentAnnotation/Value | Value |  | This element contains the actual content of this annotation.; [multi-line string] | Element |  | 1 | same | 1 | Element
THIRD | 1 | Y | MicroscopeStand-APE |  | Manufacturer Specs | This element refers to the core structure of this Microscope, which might also be referred to as Body, Frame, or Platform. | Substitution : choose one of the children elements from the list below | InvertedMicroscope, UprightMicroscope | 1 ... ∞ | same | 0 ..1 |  |  | CHANGED NAME TO Stand, Decision >>DESCRIPTION >> Include a better description that include the alternativer Names very clearly. Changed to an APE to faciliate the implementation of different Canvast drawings. |  |  |  | ISO vocabulary dostinguishes between Microscope Stand and Microscope Base. I do not think we should distinguish but we should incorporate the term Microscpe Base in the description.
| 1 | Y | MicroscopeStand-APE | Type | MicroscopeType | This field describes the general category to which this Microscope belongs. A Compound Microscope is a Microscope in which the primary image is generated by an objective or an objective and a tube-lens, and in which the Sample is observed through an eyepiece. A Custom Microscope is specially designed for specialized purposes. An Electrophysiology Microscope is a Microscope dedicated to electrophysiological observation and manipulation. A Lightsheet Microscope is a Microscope in which in contrast to epifluorescence microscopy only a thin slice (usually a few hundred nanometers to a few micrometers) of the sample is illuminated perpendicularly to the direction of observation. A Stereo Microscope (also known as Dissecting Microscope) is a low-power, binocular Microscope with long free working distance used for dissecting. Generally in this type of Microscope the object is observed by each eye from a slightly different angle, thus causing stereoscopic perception. | enum | Compound, Custom, Electrophysiology, Lightsheet, StereoMicroscope, Other |  | same |  | enum |  | Edited enum; ADDED Light sheet; REMOVED Dissection because it is the same as Stereo. Edited desription | suggest to add LightSheet to enum | [5 empty] | Multiple types should be possible in case the scope combines multiple scope types and /or the user has to indicate with which subtype he is working with
| 1 |  | MicroscopeStand-APE | ID |  | A Unique Identifier for this component. | LSID |  | R
| 1 |  | MicroscopeStand-APE | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | MicroscopeStand-APE | Origin |  | This field describes the origin of this Microscope Stand. Specifically it specifies whether this Microscope Stand was purchased from a Commercial vendor and is being used as-is (i.e., Commercial- as is), purchased from a Commercial vendor and subsequently Modified for a specialized application (i.e., Commercial- custom-modified) or entirely Custom-built from scratch (i.e., Custom-built). | enum | Commercial-as is, Commercial-custom modified, Custom-built | [5 empty] | ADDED TO ALIGN WITH SUGGESTION FROM Claire Brown Michelle Itano, Paula Montero-Lopin and Lisa Cameron |  |  |  | PROPOSE TO ADD
| 3 |  | MicroscopeStand-APE | EyepieceFieldNumber |  | The Field Number represents the diaphragm size of the Microscope eyepiece expressed in mm. | float with unit:none | [13 empty] | If available; ASI RAMM platforms only come with a frame for the stage
| 1 |  | MicroscopeStand-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 …∞ | same |  | Extension of Reference
| 1 | Y | InvertedMicroscope |  |  | This is a Microscope in which the object is observed from beneath the stage. | [7 empty] | ADDED
| 1 | Y | UprightMicroscope |  |  | This is a Microscope in which the object is observed from above the stage. | [7 empty] | ADDED
THIRD | 1 | Y | AcquisitionSoftware |  | Manufacturer Specs | This is a commercial, or open-source Software, script, macro, plugin or module utilized for any aspets of the Image Acquisition process. | - |  | 1 … ∞ |  |  |  |  | Add more field to describe a custom made software >> URL, GIT repository, Language, PATCH OR HOTFIX to indicate an addition to the software >> ADD MACRO/SCRIPT Sub-Element | I do not agree about moving this section to Settings. I consider this to be a "complete" inventory of the components of the microscopy system and the software is an integral part of that. |  |  |  |  | not sure it should be moved to setting. | Will this also include software written in LabVIEW, Python, Matlab etc to control the scope due to the inflexibilty of the vendors software?
| 1 | Y | AcquisitionSoftware | ID |  | A Unique Identifier for this component. | LSID |  | R
| 1 | Y | AcquisitionSoftware | Name |  | This field records the Name or any other designation the Author uses for this Software. | string |  | R
| 1 | Y | AcquisitionSoftware | Version |  | This field records the Version of this Software. | string |  | R |  |  |  |  | MADE Require | [5 empty] | need an entry for the last update/hotfix version
| 1 | Y | AcquisitionSoftware | Author |  | This field records the developer of this Software. This could be the Manufacturer if this is a commercial proprietary sofware or the Author(s) if this is an open-source software. | string |  | R
| 1 | Y | AcquisitionSoftware | WebsiteURL |  | This field records the URL address of any Website where the Software can be obtained. In case of open-source Software this field should record the URL for the Git repository from which to obtain the source-code. | anyURI |  | R |  |  |  |  | ADDED
| 2 |  | AcquisitionSoftware | PrimaryPublicationID |  | This field captures the a Unique Identifier for the Primary Publication describing this Software. Examples include, Digital Object Identifier (DOI), PubMed Central Identifier (PMCID), ArXiv.org ID etc. | string
| 2 |  | AcquisitionSoftware | OpenSource |  | This records whether this Software is OpenSource (True) or proprietary (False). | boolean | [6 empty] | ADDED
| 2 |  | AcquisitionSoftware | Macro |  | This field records whether this Software is a Macro or a Scprit (True) or is it a full free-standing program. | boolean | [6 empty] | ADDED
| 2 |  | AcquisitionSoftware | Build |  | The specific build of this Software. | string
| 2 |  | AcquisitionSoftware | Patch |  | This recodes the version for an optional Patch, HotFix or Update used for this Sofware to run this Instrument. A Patch (aka HotFix, or Update) is a set of changes to a computer program or its supporting data designed to update, fix, or improve it. This includes fixing security vulnerabilities, and other bugs, with such patches usually being called bugfixes or bug fixes, and improving the functionality, usability or performance. | string | [6 empty] | ADDED
| 3 |  | AcquisitionSoftware | ReleaseDate |  | The date in which this Software was released. | Date
| 3 |  | AcquisitionSoftware | Launguage |  | This records the computer Language this Software is developed in. | string | [6 empty] | ADDED
| 1 |  | AcquisitionSoftware | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 …∞ | [5 empty] | Allow 1 ... infinity?
| 1 |  | AcquisitionSoftware | AcquisitionSoftware |  | This is a commercial, or open-source Software, script or macro utilized to run this Instrument during Image Acquisition. | Element |  | 0 … ∞ |  |  |  |  | ADDED | Why this line?
| 1 | Y | AcquisitionSoftware |  |  | This is a commercial, or open-source Software, script or macro utilized to run this Instrument during Image Acquisition. |  |  | 0 … ∞
THIRD | 2 |  | MicroscopeTable |  | ManufacturerSpecs | This element describes the Table used to support this Instrument during Image Acquisition. | - |  | 0 … ∞ |  |  |  |  | ADDED MicroscopeTableSettings element and insert the attribute Floating (boolean): This recods whether the MicroscopeTable was floating (True) or not (False) during the acquistion of this Image. | [5 empty] | how many Instrument per table?
| 2 |  | MicroscopeTable | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 |  | MicroscopeTable | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | MicroscopeTable | MultipleMicroscopes |  | This field records whether (True) or not (False) multiple Microscopes share this Microscope Table. This is relevant because multiple Microscopes using the same table might affect stability and vibrations during Image Acquisition. | boolean | [6 empty] | ADDED MultipleMicroscopes (boolean)
| 2 |  | MicroscopeTable | Type | MicroscopeTableTYpe | This field describes the type of this Microscope Table. | enum | Air_Active, Air_Passive, Non-isolating Frame, Simple Bench, Stone, Other | R |  |  |  |  | CHANGED ENUMERATION
| FileAnnotation |  | MicroscopeTable | TableWidth |  | The Width of this Microscope Table. | float with unit:TableDepthUnit
| 3 |  | MicroscopeTable | TableDepth |  | The Depth of this Microscope Table | float with unit:TableWidthUnit
| 2 |  | MicroscopeTable | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞
| 1 | Y | EnvironmentalControlDevice |  | ManufacturerSpecs | This element describes an Environmental Control Device available for this Instrument. | - |  | 0 … ∞ |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 1 |  | EnvironmentalControlDevice | ID |  | A Unique Identifier for this component. | EnvironmentalControlDeviceID |  | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 1 |  | EnvironmentalControlDevice | Name |  | A User-defined Name for this component. | Denomination | [6 empty] | ADDED MultipleMicroscopes (boolean)
| 1 | Y | EnvironmentalControlDevice | Type | EnvironmentalControlType | This field captures the general category of this Environmental Control Device. | enum | Cage-incubator enclosure, Objective heather, Stage-top chamber, Other | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 1 | Y | EnvironmentalControlDevice | HumidityControl |  | This field captures whether this Environmental Control Device provides Humidity control and what type of control it provides. If no Humidity control is provided please enter "None". | enum | None, Active, Passive, Mineral oil, Water-bath, Other | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 1 | Y | EnvironmentalControlDevice | CO2Control |  | This field captures whether this Environmental Control Device provides CO2 control and what type of control it provides. If no CO2 control is provided please enter "None". | enum | None, Gas mixer, Pre-mixed, Other | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 2 |  | EnvironmentalControlDevice | O2Control |  | This field captures whether (True) or not (False) this Environmental Control Device provides vibration control. | boolean |  | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 2 |  | EnvironmentalControlDevice | VibrationControl |  | This field captures whether (True) or not (False) this Environmental Control Device provides O2 control. | boolean | [6 empty] | ADDED MultipleMicroscopes (boolean)
| 2 |  | EnvironmentalControlDevice | StandCompatibility | EnvironmentalStandCompatibility | This field captures what Microscope Stand type this Environmental Control Chamber is compatible with. | enum | Inverted, Upright, Other | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 2 |  | EnvironmentalControlDevice | Perfusion |  | This field caputers whether (True) or not (False) this Environmental Control Device supports sample Perfusion. | boolean |  | R
| 2 |  | EnvironmentalControlDevice | LidMaterial |  | In case this Environmental Control Device consists of a Sample Holder Chamber, this field captures the Material that constitutes the Lid of the Chamber. | string
| 3 |  | EnvironmentalControlDevice | MaxTemberature |  | This field captures the Maximum Temperature that can be stably maintained by this Environmental Control Device. | float with unit:MinTemperatureUnit |  | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 3 |  | EnvironmentalControlDevice | MinTemperature |  | This field captures the Minimum Temperature that can be stably maintained by this Environmental Control Device. | float with unit:MaxTemperatureUnit |  | R |  |  |  |  | ADDED MultipleMicroscopes (boolean)
| 3 |  | EnvironmentalControlDevice | TemperatureAccuracy |  | This field captures the Temperature control Accuracy provided by this Environmental Control Device. | float with unit:TemperatureAccuracyUnit | [6 empty] | ADDED MultipleMicroscopes (boolean)
| 3 |  | EnvironmentalControlDevice | TTL |  | This field captures whether (True) or not (False) this Environmental Control Device can be controlled by a TTL signal. | boolean | [6 empty] | ADDED MultipleMicroscopes (boolean)
| 2 |  | EnvironmentalControlDevice | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ |  |  |  |  | ADDED MultipleMicroscopes (boolean)
THIRD | 1 | Y | LightSource-APE |  | Manufacturer Specs | This element describes one or more Light Source(s) available for this Instrument. An Instrument may have several Light Sources of different types (Laser, Light Emitting Diode, Filament, Arc etc.). A typical fluorescence Microscope has two Light Sources, one for Transmitted light and one for Fluorescence light. | Substitution : choose one of the children elements from the list below | Arc, Filament, GenericExcitationSource, Laser, LightEmittingDiode, MultiLaserEngine | 1 ... ∞ | same | 0 ... ∞ | SubstitutionGroup |  | DO NOT DIFFERENTIATE BETWEEN TRANSMITTED AND FLUORESCENCE; | My answer to Caterina's Q: In principle the same light sources can be used in either mode. The definition in the Settings that describes how the light source is used should describe in which mode it is used in that particular configuration. |  |  |  | I agree with Damir - also reflection mode on confocal, for instance.
| 1 |  | LightSource-APE | ID (LightSourceID) |  | A Unique Identifier for this component. | LightSourceID |  | R | same | R | LSID
| 1 |  | LightSource-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same | 0 ... ∞ | Extension of Reference
[10 empty] | Power (moved to WavelengthRange) |  |  |  | MOVED TO WAVELENGTH RANGE
| 1 |  | LightSource-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | LightSource-APE | PowerMode | PowerMode | The statistic used by the Manufacturer to report the nominal Power output. | enum | Average, Maximum, Median, Other | [5 empty] | ADDED MODE (MAX Averate etc.)
| 1 | Y | Arc |  |  | An Arc lamp is a diskharge lamp that produces light by an electric arc (also called a voltaic arc), which is typically generated between metal electrodes and travels through an inert gas in a glass bulb. The inert gas can be Mercury (He), Xenon (Xe), or a mixture of the two. The choice of the gas is important because its nature determines the spectral behavior of the lamp. Refer to Type for more informations about this aspect. | - |  | - | same | - | - |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Arc | Type | ArcType | This field records the inert gas vapor contained in the Arc lamp bulb. The nature of the gas determines the spectral behavior of the lamp. At low pressure, Hg lamps emit a characteristic line spectrum, which becomes more uniform as the lamp heats up. On the other hand, Xe lamps emit light of high luminance, high color temperature, and an almost continuous spectrum distributed from the ultraviolet to the infrared. | enum | Hg, Xe, HgXe, Other | R | same |  | enum |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 1 |  | Arc | IlluminationWavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞ |  |  |  |  | ADDED
| 1 | Y | Filament |  |  | A Filament lamp is an incandescent Light Source with a wire filament (usually tungsten), which is heated by the passage of an electric current to such a high temperature that it glows with visible light (incandescence). The emitted spectrum is continuous, and it approximates that of a black-body radiator. | - |  | - | same | - | - |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Filament | Type | FilamentType | This field records the type of Filament lamp Light Source. A Halogen lamp is a Filament lamp whose envelope contains halogen vapor, which reduces the loss of tungsten from the filament. This, in turn, allows to reach a high filament temperature and consequent higher luminance, higher color temperature, and longer operating life than a Standard Filament lamp of the same input power. | enum | Standard, Halogen, Other | R | same |  | enum |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 1 |  | Filament | IlluminationWavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞ |  |  |  |  | ADDED
| 1 |  | GenericExcitationSource |  |  | The GenericExcitationSource element is used to represent a Light Source as a collection of Key/Value pairs, stored in a Map. One of the other Light Source objects should always be used in preference to this if possible. | - |  | - | same | - | -
| 2 |  | GenericExcitationSource | Map |  | This is a Mapping of Key/Value pairs used to store customized metadata that describes this element. | Element |  | 0 ... 1 | same | 0 ... 1 | Element
| 1 | Y | Laser |  |  | This element describes a Light Source, which emits coherent radiation at high spectral concentrations of radiance and minimal solid angles. This includes proper Light Amplification by Stimulated Emission of Radiation (LASER) sources that emit light through a process of optical amplification based on the stimulated emission of electromagnetic radiation. It also includes Optical Parametric Oscillator (OPO) sources that use a kind of LASER resonator but based on optical gain from parametric amplification in a nonlinear crystal rather than from stimulated emission. | - |  | - | - | - | - |  | Name changed back to Laser (was LaserLine briefly; Laser Modulator was added (for now it is in the Advanced Extension) | I prefer the Name "Laser" here. For me, laser line refers to one of the wavelengths from a multi-line laser. |  |  | Reference to Laser Modulation of various types will be put in LightSourceSettings |  | type of attenuator? laser rack and/or module if multplie lasers
| 1 | Y | Laser | Type | LaserType | This field captures the general category of this Laser. Please note that Optical Parametric Oscillators (OPO) are listed here even though they are technically not a LASER. They produce optical gain not through stimulated emission but via parametric amplification in a non-linear crystal. | enum | Chemical, Dye, Gas, MetalVapor, OpticalParametricOscillator, SolidState, SemiconductorLaserDiode, Other |  | same |  | enum |  | Reduced number of Types; reinstated OpticalParametricOscillator (OPO); removed WhiteLaser and Fiber laser from ENUM; changed description. | Make the OPO an ElectroOpticalDevice, it's not a laser, Okay to remove Chemical and MetalVapor since we're not shooting down satellites in orbit, isn't WhiteLightLaser and FiberLaser the same thing and isn't that just another SolidState laser?
| 1 | Y | Laser | Tuneable |  | This field records whether (True) or not (False) this Laser is Tuneable to different Wavelengths | boolean |  |  | same |  | boolean
| 1 | Y | Laser | Pulse |  | Whether this Laser emits light in the form of optical Pulses (i.e., light flashes; True) or as a Continuous Wave (CW; False). | boolean |  |  | same |  | enum |  | Pulse was deprecated and substituted with ContinuousWave; Made boolean
| 2 |  | Laser | LaserMedium | LaserMedium | This field captures the material out of which this Laser is constructed. | enum | Ag, Alexandrite, Ar, ArCl, ArFl, ArKr, CO, CO2, CoumarinC30, Cu, EMinus, ErGlass, ErYAG, GaAlAs, GaAs, H2O, HeCd, HeNe, HFl, HoYAG, HoYLF, Kr, KrCl, KrFl, N, NdGlass, NdYAG, Rhodamine6G, Ruby, TiSapphire, Xe, XeBr, XeCl, XeFl, Other |  | same |  | enum |  | MODIFIED ENUMERATION CONTENT but LEFT FIELD
| 2 |  | Laser | IsPumped |  | Whether (True) or not (False) this is a pumped Laser. | boolean |  | R | was Pump |  | boolean |  | ADDED; changed to IsPump to avoid confusion with dependent element
| 3 |  | Laser | RepetitionRate |  | This field (also known as Pulse Frequency) describes the frequency of laser pulses if Pulse is True. The default unit is Hz. | float with unit:RepetitionRateUnit; [default Hz] |  | R | same |  | float |  |  | I added to the description
| 3 |  | Laser | FrequencyMultiplication |  | This field captures the optical frequency multiplication factor that may be specified for this Laser. | positiveInt with unit:none |  | R | same |  | positiveInt |  | WAS ADDED BACK TO ALIGN WITH OME 2016
| 3 |  | Laser | Pump |  | The Pump element is a reference to a Light Source. It is used as a sub-Element of a Laser element to specify the Light Source for this Laser's pump (if any). | Element |  | 0 ... 1 | same | 0 ... 1 | Element |  | CHANGED DESCRIPTION
[10 empty] | Wavelength |  |  |  | REMOVED because incorporated in WavelengthRange
[10 empty] | PockelCell |  |  |  | REMOVED because a PockelCell element was added to the Advanced Extension | Why added back?
| 1 | Y | Laser | ModulationMechanism | LaserModulation | This field captures the Modulation Mechanism used for this Laser. | enum | Acousto-Optical Beam Splitter, Acousto-Optical Tunable Filter, Acousto-Optical Modulator, Direct, None, Other | [5 empty] | ADDED TO ALIGN WITH SUGGESTION FROM Claire Brown Michelle Itano, Paula Montero-Lopin and Lisa Cameron |  |  |  | Is MODULATION the same as TUNING? Can we substitute "Tuneable" with "ModulationMechanism" (enum including "non")
| 2 |  | Laser | LaserClass | LaserClass | This describes the safety Class this Laser belongs to. Laser Classes are defined based on their potential for causing injury to humans’ eyes and skin. | enum | 1, 1M, 2, 2M, 3R, 3B, 4 | [5 empty] | ADDED; changed to IsPumped to make it consistent with IsPump
| 2 |  | Laser | IsPump |  | Wheter (True) or not (False) this laser serves as a Pump for a downstream laser. | boolean |  | R |  |  |  |  | ADDED; changed to IsPump to avoid confusion with dependent element
| 3 |  | Laser | Position |  | This field captures the position of this Laser on a Multi Laser Engine. | positiveInt with unit:none | [6 empty] | ADDED |  |  |  | Propose to add in response to Nathalie comment (line 96)
| 3 |  | Laser | PulseDuration |  | Describes the duration of each Laser pulse if ContinuousWave is False. | float with unit:PulseDurationUnit [default ms] |  | R | [11 empty] | These durations can change and make the laser less efficient / Therefore it is neccessary to indicate the date when the duration was measured
| 2 |  | Laser | Acousto-OpticalDeviceRef | Reference | This empty element refers to an Acousto-Optical Device component used to modulate the Power of this Laser. | Extension of Reference |  | 0 ... ∞ |  |  |  |  | ADDED TO ALIGN WITH SUGGESTION FROM Claire Brown Michelle Itano, Paula Montero-Lopin and Lisa Cameron |  |  |  | PROPOSE TO ADD
| 1 | Y | Laser | IlluminationWavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 1 ... ∞ |  |  |  |  | ADDED
| 1 | Y | LightEmittingDiode |  |  | A Light Emitting Diode (LED) is a solid-state semiconductor Light Source that emits light when current flows through it. Electrons in the semiconductor recombine with electron holes, releasing energy in the form of photons. The color of the light (i.e., the energy of the photons) is determined by the energy required for electrons to cross the bandgap of the semiconductor. White light may be obtained by using multiple semiconductors or a layer of light-emitting phosphor on the semiconductor device. | - |  | - | same | - | - |  | Checked ISO 10934-1-2002 and LED is absent. Wikipedia uses the term freely (https://en.wikipedia.org/wiki/Light-emitting_diode); for now left the current Name; changed description to include Solid State; stil not clear whether the TERM Light Emitting Diode is encumbered by the Zeiss licence or it can be used freely. | Suggestion to change the Name to SolidState: direct LED illumination is encumbered by Zeiss patent so Manufacturers such as Lumencor use a LED-driven scintillator. SolidState would be more generic and would include novel laser-driven solidstate devices. Update: pretty much all other vendors have sub-licensed the Zeiss patent and this is no longer an issue.
| 2 |  | LightEmittingDiode | Type | LightEmittingDiodeType | This field captures the general category of this Light Emitting Diode. | enum | Collimated, Mounted, Unmounted, Other | R |  |  |  |  | Changed description.
| 1 | Y | LightEmittingDiode | LEDModule |  | A Light Emitting Diode (LED) Light Source is often composed of multiple modules each with its own Wavelength characteristics. This element describes an individual LED module. Multiple of this elements may be used. | Element |  | 1 … ∞
| 2 |  | MultiLaserEngine |  |  | This describes a Light Source device that contains multiple Lasers built into a single box. This element contains Reference(s) to individual Laser components built separately. | [7 empty] | ADDED | for multi-laser boxes frequently used in confocal microscopy, similar to LED below
| 2 |  | MultiLaserEngine | SingleMode |  | Whether the output mode of this Multi Laser Engine is single-mode (True) or multi-mode (False). | boolean | [6 empty] | ADDED | suggest changing Name to SingleMode
| 2 |  | MultiLaserEngine | NumberOfLasers |  | This field records the number of Laser lines associated with this Multi Laser Engine. | integer with unit: none | [6 empty] | ADDED | question: is this # of lasers or # of laser lines?
| 3 |  | MultiLaserEngine | OutputPortNumber |  | This field records the number of Light Guide ports used as output by this Multi Laser Engine. | integer with unit: none |  | R |  |  |  |  | ADDED
| 3 |  | MultiLaserEngine | MultiportSwitchTime |  | This field records the time it takes for this Multi Laser Engine to switch between Laser lines. | float with unit:MultiPortSwitchTime |  | R |  |  |  |  | ADDED
| 2 |  | MultiLaserEngine | Laser |  | This element describes a Light Source, which emits coherent radiation at high spectral concentrations of radiance and minimal solid angles. This includes proper Light Amplification by Stimulated Emission of Radiation (LASER) sources that emit light through a process of optical amplification based on the stimulated emission of electromagnetic radiation. It also includes Optical Parametric Oscillator (OPO) sources that use a kind of LASER resonator but based on optical gain from parametric amplification in a nonlinear crystal rather than from stimulated emission. | Element |  | 1 … ∞ |  |  |  |  | ADDED
| 3 |  | Pump |  |  | The Pump element is a reference to a Light Source. It is used as a sub-Element of a Laser element to specify the Light Source for this Laser's pump (if any). | - |  | 0 ... 1 | same | 0 ... 1 | -
| 3 |  | Pump | ID (LightSourceID) |  | This is the Unique Identifier of the Light Source used in this case. | LightSourceID |  | R | same | R | LightSourceID
| 1 | Y | LEDModule |  |  | A Light Emitting Diode (LED) Light Source is often composed of multiple modules each with its own Wavelength characteristics. This element describes an individual LED module. Multiple of this elements may be used. |  |  | 1 ... ∞ |  |  |  |  | CHANGED to utilize as Wavelength Profile as Sub-Element
| 3 |  | LEDModule | ProjectionAngle |  | This field records the angle at which this LED Module projects the light beam. | float with unit:ProjectionAngleUnit
| 1 | Y | LEDModule | IlluminationWavelengthRange |  | The nominal wavelength at which the amount of light power output is at a maximum. | Element |  | 0 ... 1
| 1 | Y | IlluminationWavelengthRange |  |  | This element records a range of Wavelengths at which this Light Source emits, according to the Manufacturer. Multiple of this element may be used. |  |  | - |  |  |  |  | CHANGED CARDINALITY
| 1 | Y | IlluminationWavelengthRange | IlluminationPower |  | This field records the nominal Power output emitted at Peak Wavelength by this Light Source, as advertised by the Manufacturer. | float with unit:IlluminationPowerUnit [default W] |  | R |  |  |  |  | MOVED HERE FROM Light Source and LEDModule
| 2 |  | IlluminationWavelengthRange | PeakWavelength | WavelengthRange | This field records the Wavelength associated with the peak response rate of this device. | float with unit:PeakWavelengthUnit |  | R
| 2 |  | IlluminationWavelengthRange | CutOn | WavelengthRange | This field records the Wavelength(s) below which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED CARDINALITY
| 2 |  | IlluminationWavelengthRange | CutOff | WavelengthRange | This field records the Wavelength(s) above which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED CARDINALITY
| 3 |  | IlluminationWavelengthRange | WavelengthProfile | FileAnnotation | This element consists of a file containing the Wavelength profile of this hardware component. | Element |  | 0... ∞ |  |  |  |  | CHANGED CARDINALITY
THIRD | 2 |  | LightSourceCoupling-APE |  | ManufacturerSpecs | This element describes the device(s)/method(s) used to couple an available Light Source(s) with this Microscope Stand to transfer light to the specimen. A typical light Microscope has a Light Coupling device for each available Light Source. | Substitution : choose one child from the list below | FreeBeam, LightGuide | 0 … ∞ | [11 empty] | The definitions of LightSourceCoupling and LightGuide sound very similar; But first: fiber or free beam?
| 2 |  | LightSourceCoupling-APE | ID |  | A Unique Identifier for this component. | LightSourceCouplingID |  | R
| 2 |  | LightSourceCoupling-APE | Name |  | A User-defined Name for this component. | Denomination
| 3 |  | LightSourceCoupling-APE | Transmittance |  | The maximum amount of light this Light Source Coupling device is capable of transmitting at the Peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R
| 2 |  | LightSourceCoupling-APE | CouplingLensRef |  | This is a reference to a Lens positioned at the connection between this Light Source coupling element and the Microscope Stand. | Extension of Reference |  | 0 … 1
| 2 |  | LightSourceCoupling-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … 1
| 2 |  | LightGuide |  |  | A Light Guide coupling between the Light Source and the Microscope Stand generally consists of a casing enclosing a medium (i.e., typically fiber optic) capable of transferring light from the illuminator to the back of the Microscope Stand. | [7 empty] | STET to LightGuide |  |  |  | Change Name to WaveGuide?
| 2 |  | LightGuide | Type | LightGuideType | This field records the general category of Light Guide used in this case. | enum | Fiber optics, Liquid, Other | R
| 3 |  | LightGuide | RefractiveIndex |  | This field records the Refractive Index of the material used to construct this Light Guide. | float with unit:none |  | R
| 3 |  | LightGuide | NumericalAperture |  | The Numerical Aperture (NA) of this device expressed as a floating-point (real) number. NA can be calculated from the Angular Aperture (alpha) using the following formula, NA = n(sin alpha), and is a measure of how much of the incident light actually passes through this device. The larger the angular aperture, the more information this device can transmit. Expected range 0.02 - 1.5. | float with unit:none |  | R |  |  |  |  | ACCEPT CHANGED DEFINITION |  |  |  | Changed definition of Numerical Aperture
| 3 |  | LightGuide | AcceptanceAngle |  | The angle within which the light beam produced by a Light Source may be coupled with the bound modes of this Light Guide. | float with unit:AcceptanceAngleUnit
| 3 |  | LightGuide | Geometry | LightGuideGeometry | This field records the cross-section shape of this Light Guide. | enum | Diffuse Axial, Backlight, Line Light, Ring Light, Spot Light, Other
| 3 |  | LightGuide | MaterialName | LightGuideMaterialName | This field records the chemical composition or the non-proprietary trade designation of the material used to construct this Light Guide. | enum | Borosilicate, Quartz, Silica, Other
| 3 |  | LightGuide | SingleMode |  | This field records whether (True) this Light Guide is single-mode. If not (False), then the guide is assumend to be multi-modal. | boolean
| 3 |  | LightGuide | WaveguideMode | LightGuideWaveguideMode | This field records the transverse mode (or superposition of such modes) that are being imposed by this light guide onto the light wave traversing it. These modes generally follow different propagation constants. | enum | Transverse electric, Transverse electromagnetic, Transverse magnetic, Hybrid, Other
| 3 |  | LightGuide | Diameter |  | This fields records the outer diameter of this Light Guide. | float with unit:DiameterUnit [default mm]
| 3 |  | LightGuide | WavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 1 ... ∞ |  |  |  |  | ACCEPT CHANGED CARDINALITY |  |  |  | Changed cardinality
| 2 |  | FreeBeam |  |  | A Free Beam coupling between the Light Source and the Microscope Stand consists of either a direct mount or an optical path that directs light across free space to the back of the Microscope Stand. | [7 empty] | CALLED FreeBeam; Changed description |  |  |  | Call this FREE BEAM?
| 2 |  | FreeBeam | DirectMount |  | Whether light traveling via this Light Source Coupling device from the Light Source to the Microscope Stand travels via a direct mount attached to the Microscope Stand (True), or across an optical path through the air (False). | boolean |  | R |  |  |  |  | CHANGED TO BOOLEAN; changed description
| 3 |  | FreeBeam | MountType |  | In case the DirectMount field is True, this field records the type of Direct Mount utilized to couple the Light Source to this Microscope Stand. | string |  | R |  |  |  |  | CHANGED DESCRIPTION | Better keep as string, these lamp mounts are very Microscope vendor specific, not at all the same as the camera mounts you list. |  |  | Should we use an Enum? Would this be the same types of mount as the ones listed for Camera Mount? The ones listed there are: C-mount, T-mount, SLR adapter, D-SLR adapter, Other
| 2 |  | WavelengthRange |  |  | This element records a range of Wavelengths this component is designed for, according to the Manufacturer. Multiple of this element may be used. |  |  | 1 ... ∞ |  |  |  |  | CHANGED CARDINALITY
| 2 |  | WavelengthRange | PeakWavelength |  | This field records the Wavelength associated with the peak response rate of this device. | float with unit:PeakWavelengthUnit |  | R
| 2 |  | WavelengthRange | CutOn | WavelengthBounds | This field records the Wavelength(s) below which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED CARDINALITY
| 2 |  | WavelengthRange | CutOff | WavelengthBounds | This field records the Wavelength(s) above which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED CARDINALITY
| 3 |  | WavelengthRange | WavelengthProfile | FileAnnotation | This element consists of a file containing the Wavelength profile of this hardware component. | Element |  | 0... ∞ |  |  |  |  | CHANGED CARDINALITY
THIRD | 1 | Y | Objective |  | Manufacturer Specs | The Microscope's Objective lens consists of a lens, its mount, and any associated parts. It is the part of the imaging system, which forms a primary image of the object, either alone or in conjunction with a tube lens. The Objective typically consists of a compound lens consisting of several simple lenses (elements), usually arranged along a common axis. | - |  | 1 ... ∞ | same | 0 ... ∞ | - |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 1 |  | Objective | ID |  | A Unique Identifier for this component. | LSID |  | R | same | R | LSID
| 1 |  | Objective | ImmersionType | ImmersionTypeList | According to ISO 10934-1:2002, an Immersion Liquid is a liquid (commonly immersion oil, water or glycerol) specified as suitable for use in the space between the front of an immersion lens and the object NOTE Because the immersion liquid is considered in the computing of corrections to be part of the lens, its refractive index and dispersion (or Abbe number) are critical. Please refer to ISO 8036-1 for more details on Immersion Oil. This field describes the general category of the Immersion Liquid this Objective is designed for, according to the Manufacturer. If the general category is not "Air", the actual Immersion Liquid that was used during the Acquisition of a specific Image must be specified in Objective Settings. | enum | Air, Dipping, Glycerol, Multi, Mineral Oil, Silicone Oil, Water, Other |  | was Immersion |  | enum |  | WAS IN OME; added back
| 1 | Y | Objective | Correction | ObjectiveCorrection | This field captures the type of optical correction (i.e. chromatic correction, or flatness correction, etc.) associated with this Objective. | enum | Achro, Achromat, Achroplan, Apo, Fl, Fluar, Fluor, Fluotar, Lambda, Neofluar, Pl, Plan, Plano, PlanApo, PlanFluor, PlanNeoFluar, SuperFluor, UPlanApo, UPlanFl, UV, VioletCorrected, Other | R | same |  | enum
| 1 | Y | Objective | LensNA |  | This field captures the nominal Numerical Aperture (NA) of this Objective as specified by the Manufactuer, expressed as a floating-point (real) number (e.g. 1.27 NA). NA is a number that was defined initially by Abbe to characterize Objectives and Condensers. NA can be calculated by the expression n sin alpha, where n is the Refractive Index of the medium between the Lens and the Sample, and alpha is half the Angular Aperture of the Lens. NA is a measure of how much of the illumination that leaves the specimen actually passes through the Lens. The larger the NA, the more information the lens can transmit. Expected range 0.02 - 1.5. | float with unit:none |  | R | same |  | float |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | Changed LensNA definition
| 1 | Y | Objective | Magnification |  | This field captures the Magnification of the Objective as specified by the Manufacturer - i.e., 60 represents a 60X lens. | float with unit:none |  | R | same |  | float |  |  | How about CalibratedMagnification here rather than in Settings? |  |  | CalibratedMagnification sould definitely be a Spec
| 2 |  | Objective | CalibratedMagnification |  | This field captures the Magnification of the Objective as measured via a calibration process- i.e., '59.987' for a 60x lens. | float with unit:none |  |  | same |  | float |  | CALIBRATED MAGNIFICATION TO BE MOVED BACK TO SPECS
| 2 |  | Objective | Iris |  | This field records whether (True) or not (False) the Objective was fitted with an Iris. Manufacturers offer high NA Objectives, which are equipped with an internal Iris Diaphragm that increases image contrast during digital imaging to compensate for image flare artifacts due to light emitted from areas outside the focal plane. Opening or closing the Iris Diaphragm determines the size of the Objective rear aperture yielding a variable NA range between 0.5 and the Objective's upper limit. | boolean |  | R | same |  | boolean |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Objective | WorkingDistance |  | This field reports the Working Distance of this Objective under normal operating conditions. The Working Distance is the distance (in air, or in the specified immersion liquid) between the front of an Objective and the surface of the cover glass or of the specimen if uncovered. | float with unit:OpticalDistanceUnit; [default mm] |  | R | same |  | float |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. Changed unit |  |  |  | Changed unit
| 1 |  | Objective | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same | 0 ... ∞ | Extension of Reference
| 1 |  | Objective | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | Objective | LightType | ObjectiveLightType | This field specifies whether this Objective is designed for a special Light application (e.g., IR or UV). | enum | Infrared, Ultraviolet, Visible | R
| 1 | Y | Objective | DIC |  | This field specifies whether (True) the Objective is designed for Differential Interference Contrast (DIC) or not (False). DIC is a contrast modality due to double-beam interference in which two waves that fall on the object plane or image plane are separated laterally by a distance similar to the minimum resolvable distance. | boolean |  | R |  |  |  |  | DECISION ADDED DIC; Changed description to incorporate ISO 10934-1-2002 standard definitions. | Added to address the objective's DIC capability
| 1 | Y | Objective | ContrastModulation | ContrastModulationPlate | This field specifies the presence in this Objective of a contrast modulation plate, such as a Phase Ring (Ph1, Ph2, Ph3), a Hoffman modulator, a VAREL modulator, or similar. | enum | None, Ph1, Ph2, Ph3, Hoffman, VAREL, Other | R |  |  |  |  | DECISION ADDED ContrastModulationPlate | Not sure if DIC belongs here since the objective itself doesn't modulate anything. For DIC it just needs to be made from strain-free glass. Maybe better to add another attribute for that as shown above? |  |  |  |  | how about DIC?
| 1 |  | Objective | InfinityCorrected |  | This field records whether (True) or not (False) this Objective is infinity corrected. An Infinity-corrected Objective is corrected for an infinite sample to primary image distance and needs to be used with a Tube Lens of appropriate focal length to obtains its nominal magnification. | boolean |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Objective | PhaseContrastDesignation | PhaseContrastDesignationType | If the Objective is designed for Phase, this field specifies the Phase Contrast designation of the Phase Plate found in this Objective (see: https://www.microscopyu.com/techniques/phase-contrast/phase-contrast-Microscope-configuration). This designation has to match the Phase Number of the Phase Ring associated with the front focal plane of the Condenser. | enum | Phase, PHACO, PC, PhL, Ph1, Ph2, Ph3, Ph4, DL, DLL, DM, ADL, PL, PLL, PM, PH, NL, NM, BM, NH | [5 empty] | DECISION ADDED PhaseContrastType
| 2 |  | Objective | DippingMedium | DippingType | In case the Immersion Type for this Objective is designated as "Dipping", this field captures the type of Dipping Medium this Objective is compatible with. Examples include, water-based (biological buffers, culture media, other water-based solutions) organic-based [FocusClear (CLARITY); CUBIC-1, CUBIC-2, CUBIC-R1, CUBIC-R2; 2,2-thiodiethanol (TDE); ethyl cinnamate; Benzyl Benzoate and BB-PEG (PEGASOS); Benzyl Alcohol, Benzyl Benzoate (BABB); dibenzyl ether (DBE); 100% Ethanol]. | enum | Orgnic-based, Water-based, Other
| 2 |  | Objective | CorrectionCollar |  | This field records whether (True) or not (False) the Objective was fitted with a Correction Collar. When using high NA dry Objectives (NA 0.8 or greater), cover glass thickness variations of only a few micrometers, immersion medium, and temperature changes result in dramatic image degradation due to aberration. To compensate for these errors, the more highly corrected Objectives are equipped with a Correction Collar mechanism, which is provided to adjust the central lens group position to compensate for deviations from correct optical path length due to changes in Cover Glass Thickness, the nature of the Immersion Liquid, different Numerical Apertures, and Temperature. | boolean |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Objective | CorrectionCollarType | CorrectionCollarTypeList | This field records the type of Correction Collar with which this Objective is fitted. | enum | Coverglass Thickness, Immersion Liquid, Numerical Aperture, Temperature | [5 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 3 |  | Objective | ImageDistance |  | This field records the distance in air between the rear principal plane of this Objective and the primary Image plane. | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED
| 3 |  | Objective | ObjectDistance |  | This field records the distance in air between the front principal plane of this Objective and the field plane in which the object (i.e., specimen) is situated. | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED
| 3 |  | Objective | FrontFocalLength |  | This field records the distance (f) measured along the optical axis from the front principal plane of this Lens to the front focal plane (F). | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED
| 3 |  | Objective | BackFocalLenght |  | This field records the distance (f') measured along the optical axis from the rear principal plane of this Lens to the rear focal plane (F'). | float with unit:OpticalDistanceUnit[default mm] | [6 empty] | ADDED
| 3 |  | Objective | ParfocalizingDistance |  | This field records the distance between the mounting plane of this Objective and the specimen plane. | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | Changed Name and description to incorporate ISO 10934-1-2002 standard definitions.
| 3 |  | Objective | ObjectiveViewField |  | This field records the maximal diameter of the view field allowed by this Objective, expressed in mm, and measured at the intermediate image plane. NB: The field of view is influenced by the magnification and diameter of the ocular (eyepiece) field diaphragm, but also by the limits imposed by the design of the Objective. | float with unit:none
THIRD | 1 | Y | OpticalAssembly-APE |  | Manufacturer Specs | This is a device that might include one or more Lenses, other Optics and Optics Holders. Optical Devices might be custom built or they might also be available as pre-assembled units. Such devices might be embedded within the Microscope Stand or exist as independent units. Examples include the Condenser and the Microscope Tube. | Substitution: | BeamCollimator, BeamExpander, Condenser, MicroscopeTube | 0 … ∞ |  |  |  |  | CREATED; ADDED DESCRIPTION |  |  |  | propose to call it OpticalDevice
| 1 |  | OpticalAssembly-APE | ID |  | A Unique Identifier for this component. | LSID |  | R
| 1 |  | OpticalAssembly-APE | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | OpticalAssembly-APE | LensRef |  | This refers to an available Lens, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 1 |  | OpticalAssembly-APE | OpticalApertureRef |  | This refers to an available Optical Aperture, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 1 | Y | OpticalAssembly-APE | PolarizationOpticsRef |  | This refers to an available Polarization Optics, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 1 | Y | OpticalAssembly-APE | PrismRef |  | This refers to an available Prism, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 2 |  | OpticalAssembly-APE | FilterRef |  | This refers to an available Filter, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 2 |  | OpticalAssembly-APE | MirrorringDeviceRef |  | This refers to an available Mirroring Device, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 2 | Y | OpticalAssembly-APE | OpticsHolderRef |  | This refers to an available Optics Holder, which is part of this Optical Device. | Extension of Reference |  | 0 … ∞
| 1 |  | OpticalAssembly-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 2 |  | BeamCollimator |  |  | Beam Collimators are Lens Systems that transform a strongly diverging light beam into a collimated beam, i.e., a beam where light propagates essentially only in one direction, and the beam divergence is weak. Beam Collimators typically contain some kind of Focusing Lens, or sometimes a curved Mirror, and usually some opto-mechanics for stable mounting and possibly adjustment. Beam Expanders can be custom built or they are also available as pre-assembled optical devices. | [7 empty] | ADDED |  |  |  | added definitions
| 2 |  | BeamExpander |  |  | Beam Expanders are Lens Systems whose role is to increase the diameter of a collimated input beam to a larger collimated output beam. Beam Expanders typically are afocal systems in which object rays enter and exit parallel to the optical axis. As such a Beam Expander does not have a focal length. In most cases, Beam Expanders are realized as optical telescopes consisting of two lenses. Beam Expanders can be custom built or they are also available as pre-assembled optical devices. | [7 empty] | ADDED |  |  |  | added definitions
| 3 |  | BeamExpander | Configuration | BeamExpanderConfiguration | This field captures the Configuration of this Beam Expander. | enum | Galileian telescope, Keplerian telescope, Other
| 1 | Y | Condenser |  |  | A Condensers is a part of the illuminating system of the microscope which consists of one or more lenses (or mirrors), their mounts, and typically a Diaphragm. They typically also include a Turret system to facilitate switching between different Optics. The are located above the Light Source and under the sample in an Upright Microscope, and above the stage and below the Light Source in an Inverted Microscope. They are designed to collect, control and concentrate light from the Light Source into a cone of light that illuminates the specimen. The aperture and angle of the light cone must be adjusted (via the size of the Diaphragm) for each different Objective. | [7 empty] | ADDED |  |  |  | added definitions
| 1 | Y | Condenser | Correction | CondenserCorrection | This field captures any type of Aberration (i.e., failure of an optical system to produce a perfect image) this Condenser is design to correct. | enum | Astigmatism, Chromatic, Chromatic and Curvature, Coma, Curvature, Distorsion, Spherical, Other
| 2 |  | Condenser | Type | CondenserCategory | This field captures the general category this Condenser belongs to according to ISO 10934-1:2002. Specifically, this category refers to the type of Correction or Contrast Enhancement Method this Condenser is designed to achieve. According to ISO 10934-1:2002, options include: 1) Abbe, this is a condenser of simple design introduced by Abbe, in which there is only limited correction for spherical aberration and none for chromatic aberration. 2) Achromatic-aplanatic, this is a in which chromatic aberrations and spherical aberrations have been reduced. 3) Cardioid, this is a dark field condenser for transmitted light illumination that has spherical aberration and coma (i.e., aberration in which the image of an off-axis point object is deformed so that the image is shaped like a comet) correction. 4) Dark-field, this is a dark-ground condenser designed for dark-field microscopy, which for reflected-light microscopy might be located within the mount of the objective. 5) Pancratic, this is a condenser containing a variable “zoom” (pancratic) lens which allows the size of the illuminated field at the object to be varied while the illuminated field diaphragm remains of constant size. 6) Phase-contrast, this is condenser designed for phasecontrast microscopy. 6) Universal, this is a condenser designed for multiple contrast methods such as bright field, dark field, phase contrast, differential interference contrast, polarized light and modulation contrast. | enum | Abbe, Achromatic-Aplanatic, Cardioid, Dark-field, Pancreatic, Phase-contrast, Universal, Other | R |  |  |  |  | ADDED |  |  |  | Added. OK?
| 2 |  | MicroscopeTube |  |  | Illumination Tubes are Lens systems, which might be embedded within the Microscope Stand and therefore not immediately accessible. They are essential for infinity-correction and should be regarded as integral parts of the Objective Lens System when magnification and correction are considered. | [7 empty] | ADDED |  |  |  | Added. OK?
| 2 |  | MicroscopeTube | Type | MicroscopeTubeType | This field captures the category of this Tube. | enum | Body, Intermediate, Viewing, Other | R
| 3 |  | MicroscopeTube | MechanicalLength |  | For Objectives corrected for a finite primary image distance, this is the length of the Tube in its simplest form, and it corresponds to the distance in air between the Objective and the Camera port. It commonly has a value of 160 mm (ISO 9345-1). For infinity-corrected objectives, the mechanical tube length is hypothetically considered to be infinite. | float with unit:LengthUnit ; [default mm]
| 3 |  | MicroscopeTube | OpticalLenght |  | This is the distance between the back focal plane of the Objective and the primary Image plane. This distance is relevant only when using finite primary Image distance Objectives. | float with unit:LengthUnit ; [default mm]
| 3 |  | MicroscopeTube | Factor |  | Factor (q) by which the lateral Magnification at the primary image is changed by an intermediate Lens or Lens System inserted between the Objective and the primary Image. | float with unit:none
| 1 | Y | OpticalAperture-APE |  | Manufacturer Specs | This element represents any mechanical, non optical device that is placed in the Light Path and spatially modifies the light beam. Examples include, Pin Holes, Phase Rings, Dark Field Stops, Hoffman Modulators, Diffraction Gratings etc. This element can be associated with an optional Optics Turret. | Substitutions: | DarkFieldStop, DiffractionGrating, FourierFilter, PhaseRing, PinHole, HoffmanPlate, IrisDiaphragm, VarelRing | [5 empty] | Changed to Optical Aperture; make sure Ulrike agrees. | Didn't see anywhere else to put things such as these. Develop this further? Who? |  |  | HAD PROPOSED OptoMechanicalDevice BUT THEY are not really MECHANICAL IN NATURE BECAUSE THEY DO NOT MOVE; PROPOSE TO CALL IT OPTICAL APERTURE; OPTICAL MASK; SPATIAL FILTER??
| 1 |  | OpticalAperture-APE | ID |  | A Unique Identifier for this component. | FilterID |  | R
| 1 |  | OpticalAperture-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | OpticalAperture-APE | OpticsHolderPosition |  | This field records the position of this Optical Aperture in a Optics Turret. | positiveInteger with unit:none |  | R | [5 empty] | Ok, many of these devices will be mounted in a condenser turret but some (apertures) will be singly somewhere else. If in a multi-holder, do we also want to record the position of this Device in that holder? |  |  | Answer: this is an optional field so if it is not needed it will be left out.
| 3 |  | OpticalAperture-APE | OuterDiameter | OptoMechanicalElementSize | This field captures the physical outer diameter of the component | float with unit:DiameterUnit [default mm]
| 1 |  | OpticalAperture-APE | AnnotationRef | CommentAnnotation | This field is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 2 |  | DarkFieldStop |  |  | This element refers to a central circular stop (i.e., opaque disk) usually used in dark-field microscopy and placed in the front focal plane of a Condenser to occlude all the direct light which would fall within the aperture of the objective. |  |  | 0 … ∞ |  |  |  |  | MODIFIED DESCRIPTION TO ALIGN WITH ISO 109341-1
| 3 |  | DarkFieldStop | CenterStopDiameter | DarkFieldStopSize | This field captures the physical diameter of the central disk of this Dark Field Stop. | float with unit:DiameterUnit [default mm] | [11 empty] | Not sure if I missed diskussion - but how about oblique lighting?
| 1 | Y | PhaseRing |  |  | This element (also called Phase Annulus) refers to an annular (ring-like) slit used to produce contrast modulation following Zernike, in which the image contrast of a phase object is enhanced by altering the phase and amplitude of the direct light with respect to those of the diffracted light. This illuminating aperture Diaphragm is placed in the front focal plane of the Condenser and acts in conjunction with an appropriate phase plate, usually in the form of an annulus, placed in the back focal plane of the objective (or in a succeeding plane conjugate with this). |  |  | 0 … ∞ |  |  |  |  | MODIFIED DESCRIPTION to align with ISO 10934-1:2002 |  |  |  | I have seen this referred to as Phase Stop; ISO has no Name that I can find. DECISION?
| 1 | Y | PhaseRing | PhaseNumber | PhaseRingNumber | This field captures the Phase Number used to identify the size of the annular slit of this Phase Ring. The Phase Number must match the Phase Plate found within a Phase-contrast capable Objective lens. | enum | PhL, Ph1, Ph2, Ph3, Ph4 | R |  |  |  |  | LEAVE HERE AND ALLOW SELECTION IN SETTINGS; modified description | I would think it should be here. You would have a selection of PhaseRings in your inventory and in Settings you pick one of the available ones. |  |  | Changed Name and Description. QUESTION >>> SHOULD THIS BE MOVED TO SETTINGS? THIS IS CHANGED DURING ACQUISTION SO I THINK SO
| 3 |  | PhaseRing | CenterStopDiameter |  | This filed records the physical diameter of the center-stop (i.e. opaque disk) of this Phase Ring. | float with unit:DiameterUnit [default mm]
| 3 |  | PhaseRing | SlitWidth |  | This field records the physical width of the annular slit of this Phase Ring. | positivefloat with unit:WidthUnit | [6 empty] | CALLED SLITWIDTH |  |  |  | Propose to call it SlitWidth
| 2 |  | HoffmanSlit |  |  | This element represents an off-axis Slit aperture positioned in the front focal plane of the Condenser. It is used alongside a Hoffman contrast modulator (i.e., optical amplitude spatial filter) present in the back focal plane of the objective or in a succeeding conjugate plane to produce contrast modulation as described by Hoffman (i.e., Hoffman Contrast Modulation, HMC). Hoffman modulation contrast is not hampered by the use of birefringent materials (such as plastic Petri dishes) in the optical pathway, so the technique is more useful for examining specimens in containers constructed with polymeric materials. |  |  | 0 … ∞ |  |  |  |  | Changed Name to Hoffman Slit; Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | I have seen the term Plate often used to refer to a component placed within the Objective. Propose to call this Hoffman Slit or Hoffman Slit Plate instead.
| 3 |  | HoffmanSlit | Design | HoffmanSlitDesign | This field captures the design of the Slit of this Hoffman Slit. | enum | Center, Offset, Other | [5 empty] | CHANGED |  |  |  | Proposed to change from Type to Design
| 2 |  | IrisDiaphragm |  |  | This element represents a Diaphragm, positioned trasversally to the optical axis bounded by multiple leaves, usually metal, arranged so as to provide an opening of variable size, which is adjustable by means of a mechanical control. |  |  | 0 … ∞ |  |  |  |  | CHANGED NAME; Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | IrisDiaphragm | MaxApertureDiameter |  | This field records the Maximum physical Diameter of the Aperture of this Iris Diaphragm. | float with unit:ApertureDiameterUnit [default mm]
| 3 |  | IrisDiaphragm | Shape | IrisShape | This field captures the Shape of the aperture of this Iris Diaphragm. | enum | Diamond, Round, Slit, Square, Other
| 2 |  | VarelRing |  |  | This element represents a stop (i.e., opaque structure) shaped like a ring sector, which is used on the illumination side, to permit unilateral inclined illumination. It was developed to work in combination with Phase Objectives to yield contrast even when light passes through birefringent materials, such as plastic Petri dishes. |  |  | 0 … ∞ |  |  |  |  | Changed description |  |  |  | Is this a compound device?
| 1 | Y | Shutter |  |  | This element describes a Shutter available for this Instrument. Shutters are devices that open and close, and they may be mechanically or electro-mechanically controlled. Shutters are typically used to control the Light that raches the sample (i.e., to limit phototoxicity and photobleaching especially in case of live specimens), or the Detecrtor at any one time. They can also be used for Excitation or Emission Wavelength selection. | [7 empty] | ADDED in response to Motero, Itano, Cameron
| 1 | Y | Shutter | ResponseTime |  | This field (also known as Speed) records the time it takes for this Shutter to completely open and close, according to the Manufacturer. | float with unit:ResponseTimeUnit | [6 empty] | ADDED in response to Motero, Itano, Cameron
| 2 |  | Shutter | Function | ShutterFunction | This field records whether this Shutter is used for the control of illumination (especially relevant for live imaging), the selection of Excitation or Emission wavelenghts or the control of the Light that reaches the Detector. | enum | Detection control, Illumination control, Wavelength selection, Other | [5 empty] | ADDED in response to Motero, Itano, Cameron
| 2 |  | Shutter | OperationFrequency |  | This field records the Frequency of Operation of this Shutter, according to the Manufactuers. | float with unit:OperationFrequencyUnit
| 2 |  | Shutter | Motorized |  | This field describes whether this device is Motor (True) or Manually (False) operated. | boolean | [6 empty] | ADDED in response to Motero, Itano, Cameron
| 2 |  | Shutter | MaxApertureDiameter |  | This field records the Maximum physical Diameter of the Aperture of this Shutter. | float with unit:ApertureDiameterUnit | [6 empty] | ADDED in response to Motero, Itano, Cameron
| 2 |  | Shutter | Multimode |  | This field records whether (True) multiple operation modes are available for this Shutter or not (False). Typical operation modes are are: Fast, Soft and Neutral Density. The Fast mode provides the quickest open and close action sequence for which the shutter is capable. The Soft mode gradually opens and closes the shutter to reduce vibration. Control of light intensity without affecting wavelength is accommodated by a Neutral Density shutter mode (for more details see https://www.microscopyu.com/applications/live-cell-imaging/the-automatic-microscope). | boolean | [6 empty] | ADDED in response to Motero, Itano, Cameron
| 1 | Y | StageInsert |  | ManufacturerSpecs | This element describes an optional Stage Insert mounted on the Stage. |  |  | 0 … ∞ |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. CHANGED TO STAGE INSERT; MADE INDEPENDENT OF SAMPLE POSITIONING | Made into separate class since it isn't a specialization of a SamplePositiong device |  |  | Propose to change Name to SpecimenHolder
| 1 |  | StageInsert | ID |  | A Unique Identifier for this component. | StageInserID |  | R
| 1 |  | StageInsert | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | StageInsert | SampleHolderType |  | This field describes the type of Specimen Holder this Stage Insert is designed to hold. | enum | Chambered Coverslip, Culture Plate, Microscope Slide, Multiwell Plate, Petri Dish, Other | R |  |  |  |  | ADDED RES + POS LINEARITY etc. for XY axes |  |  |  | Propose to add
| 2 |  | StageInsert | TemperatureControlled |  | This field records whether this Stage Insert is Temperature Controlled (True), or not (False). | boolean |  | R |  |  |  |  | ADDED RES + POS LINEARITY etc. for XY axes; CHANGED NAME and DESCRIPTION | Heating is a property of the insert, not of the stage itself. |  |  |  | Cooling?
| 2 |  | StageInsert | NumberOfSlots |  | This field describes the number of Specimen Holder slots on this Stage Insert. | integer with unit:none | [6 empty] | ADDED | Added
| 2 |  | StageInsert | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
THIRD | 1 | Y | SamplePositioning-APE |  | ManufacturerSpecs | This element describes a Sample-Optics Positioning device available for this Instrument. An Instrument may have more than one of these devices of different types. For example, a typical setup might include a Mechanical Stage, a Stage Insert, a Turret Objective Focusing, and a Piezoelectric Objective Focusing device. | Substitution: | StageInsert, MechanicalStage, PiezoElectricStage, TurretObjectiveFocusing, PiezoElectricObjectiveFocusing | 1 … ∞ |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. Modified description; added the capacity to control XY vs Z movement, resolution etc. | Changed name back to SamplePositioning-APE |  |  | Motorized vs, Manual
| 1 |  | SamplePositioning-APE | ID |  | A Unique Identifier for this component. | SamplePositioningID |  | R
| 1 |  | SamplePositioning-APE | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | SamplePositioning-APE | Encoded |  | This field describes whether this Sample Positioning Device is equipped with a system (i.e., encoder) to monitor movement and provide automatic positional feedback for high-precision operations, and it specifies the type of encoders it has. | enum | None, Linear, Rotary, Other | R |  |  |  |  | ADDED | Propose to add. Update: No, high-precision has no quantitative meaning. Actually would be good to specify whether Rotary or Linear encoded so I made it an enum. See: http://www.asiimaging.com/products/stages-old/linear-encoder-options/ |  |  | propose to call it HighPrecision instead of Endocded
| 1 | Y | SamplePositioning-APE | Motorized |  | This field describes whether (True) or not (False) this Sample Positioning Device is fitted with a Motorized component (i.e., ball-screw drives or rack mechanisms, sometimes equipped with precision calibration) to primarily assist in the translational movement of the sample in the X and Y or Z directions. A Motorized Stage is also often referred to as a Scanning Stage. | boolean |  | R |  |  |  |  | ADDED
| 1 | Y | SamplePositioning-APE | ZMaxVelocity |  | This field describes the maximum Z-axis Velocity at which this Sample Positioning Device is capable of moving. | float with unit:ZMaxVelocityUnit | [6 empty] | CHANGED TO Z specific
| 1 | Y | SamplePositioning-APE | ZRepetability |  | This field captures the expected Accuracy with which this Sample Positioning Device is capable of returning to a previously visited Z-axis Position, as specified by the Manufacturer. This metric is important for example for Tiling. | float with unit:ZRepetabilityUnit | [6 empty] | CHANGED TO Z specific
| 1 | Y | SamplePositioning-APE | ZTravelRange |  | This field describes the total range in the Z-axis that this Sample Positioning Device can travel. | float with unit:ZTravelRangeUnit | [6 empty] | REMOVED REQUIRED;
| 2 |  | SamplePositioning-APE | MovementAxes |  | This field describes the axes or combination of axes along which this Sample Positioning Device is capable of moving. | enum | X, Y, Z, XY, XYZ, Θ, Other | R
| 2 |  | SamplePositioning-APE | BuiltIn |  | This field indicates whether this Sample Positioning Device is built into the Microscope Stand (True) or it is removable and interchangeable (False). | boolean |  | R |  |  |  |  | ADDED
| 2 |  | SamplePositioning-APE | ZReproducibility |  | This field captures the expected Precision with which this Sample Positioning Device is capable of maintaining a given Z-axis Position over time during as specified by the Manufacturer. This metric is generally referred to as Reproducibility and it might also be referred to as Focal Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:ZReprducibilityUnit |  | R
| 2 |  | SamplePositioning-APE | ZSettlingTime |  | This field describes the expected time it takes for this Sample Positioning Device to stabilize at the target Z-axis position, as specified by the Manufacturer. | float with unit:ZSettlingTimeUnit |  | R |  |  |  |  | CHANGED TO Z specific
| 2 |  | SamplePositioning-APE | ZResolution |  | This field describes the Z-axis motion Resolution of this Sample Positioning Device as specified by the Manufacturer. | float with unit:ZResolutionUnit |  | R |  |  |  |  | CHANGED TO Z specific
| 2 |  | SamplePositioning-APE | ZPositionLinearityError |  | This field describes the Z-axis positioning error (typically expressed in % error) associated with this Sample Positioning Device as specified by the Manufacturer. | float with unit:none | [6 empty] | CHANGED TO Z specific
| 3 |  | SamplePositioning-APE | Type |  | This field describes the general category of this Sample Positioning Device. | enum | Centering, DC Servo Motor, Gliding, Levelling, Linear, Rigid Stand, Stepper, Translational, Other | [5 empty] | ADDED; MOVE THIS TO GRANDMOTHER level | No, important to keep DC in the enum. Positioning accuracy and speed of DC motors is quite different from stepper motors. Another Name for DC motor is servo motor; is that more familiar? |  |  | Propose to move Type to the generic Stage-APE [propose to add Levelling as a type. Should we add Heating as a type?]
| 3 |  | SamplePositioning-APE | RotationAngle |  | The angle by which this Sample Positioning Device is capable of rotating. | float with unit:RotationAngleUnit [default deg]
| 2 |  | SamplePositioning-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 | Y | FocusingDevice-APE |  |  | A Focusing mechanism is used to obtain maximum sharpness in an Image by changing the distance between a Sample and the optical system forming the Image. In most cases, a Z-axis Focusing mechanism moves the Objective along the optical axis.
| 1 | Y | TurretObjectiveFocusing |  |  | A Focusing mechanism is used to obtain maximum sharpness in an Image by changing the distance between a Sample and the optical system forming the Image. In most cases, a Z-axis Focusing mechanism moves the Objective along the optical axis. This element describes a Focusing Device associated with a rotating Objective Turret associated with this Instrument. This type of device is often operated using a rack and pinion mechanism. | [7 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions. Changed Name | Is this different from the item below? |  |  | Propose to call it MechanicalObjectiveFocusing
| 1 | Y | IndividualObjectiveFocusing |  |  | A Focusing mechanism is used to obtain maximum sharpness in an Image by changing the distance between a Sample and the optical system forming the Image. In most cases, a Z-axis Focusing mechanism moves the Objective along the optical axis. This element describes an optional objective nano-Focusing Device associated with this Instrument. These devices might be driven using piezoelectric mechanisms or linear motor actuators associated with crossed roller guideways. They typically achieve higher precision, accuracy, and speed than traditional Focusing Devices. | [7 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions. Changed Name
| 2 |  | IndividualObjectiveFocusing | PiezoElectric |  | This field describes whether this device is Piezoelectric driven (True) or not (False). | boolean |  | R
| 2 | Y | ZDrive |  |  | This element describes a Focusing Device designed to electronically control a moving a mechanical part of the Microscope such as the Stage (see Stage-APE), or the Microscope Stand itself (e.g. most dissecting Microscopes). An example of such devices is the MFC-2000 Z-Axis Drive and Controller (http://www.asiimaging.com/controllers/mfc-2000-z-axis-drive-and-controller/#tab-id-4). | [7 empty] | MAKE THIS A SPECIALIZATION OF MECHANNICAL STAGE
| 1 | Y | Stage-APE |  |  | This element describes a platform positioned at right angle to the optical axis of this Microscope, which is designed to either directly hold a Specimen Holder containing a Sample to be observed, or a Stage Insert, which itself holds the Sample Holder. In many cases, this device is capable of mechanical movements (as in a Mechanical Stage) to allow easy positioning and rotation of the Sample in the X-, Y-, and Z-axes. | [7 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions. Changed Name
| 1 | Y | Stage-APE | XYMaxVelocity |  | This field describes the maximum XY-axis Velocity at which this Stage is capable of moving. | float with unit:XYMaximumVelocity | [6 empty] | ADDED RES + POS LINEARITY etc. for XY axes
| 1 | Y | Stage-APE | XYRepetability |  | This field captures the expected Accuracy with which this Stage is capable of returning to a previously visited XY-plane Position, as specified by the Manufacturer. This metric is important for example for Tiling. | float with unit:XYRepetabilityUnit | [6 empty] | ADDED RES + POS LINEARITY etc. for XY axes
| 2 |  | Stage-APE | XTravelRange |  | This field describes the Total Range in the X-axis that this Stage can Travel. | float with unit:XTravelRangeUnit |  | R
| 2 |  | Stage-APE | YTravelRange |  | This field describes the Total Range in the Y-axis that this Stage can Travel. | float with unit:YTravelRangeUnit |  | R
| 2 |  | Stage-APE | XYReproducibility |  | This field captures the expected Precision with which this Stage is capable of maintaining a given XY-plane Position over time during as specified by the Manufacturer. This metric is generally referred to as Reproducibility and it might also be referred to as Lateral Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:XYReprducibilityUnit |  | R
| 2 |  | Stage-APE | XYSettlingTime |  | This field describes the expected time it takes for this Stage to stabilize at the target XY-plane position, as specified by the Manufacturer. | float with unit:XYSettlingTimeUnit |  | R |  |  |  |  | ADDED RES + POS LINEARITY etc. for XY axes
| 2 |  | Stage-APE | XYResolution |  | This field describes the XY-plane motion Resolution of this Stage as specified by the Manufacturer. | float with unit:XYResolutionUnit |  | R |  |  |  |  | ADDED RES + POS LINEARITY etc. for XY axes
| 2 |  | Stage-APE | XYPositionLinearityError |  | This field describes the XY-axis positioning error (typically expressed in % error) associated with this Stage as specified by the Manufacturer. | float with unit:none | [6 empty] | ADDED RES + POS LINEARITY etc. for XY axes
| 1 | Y | MechanicalStage |  |  | This element describes a Stage fitted with ball-screw drives or rack mechanisms (sometimes equipped with precision calibration) to primarily assist in the translational movement of the Sample in the X and Y directions. A Mechanical Stage may be built into the Microscope Stand or it may consist of a removable insert. It may be manually operated, or it might be driven by linear electrodynamic motors, such as for example, in computer-controlled digital Microscopes. In some rare cases, a Mechanical Stage might also be used for Focusing (Z-axis movement), i.e., in simple surveying Microscopes. | [7 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions. | Add MotorizedStage (or see above) |  |  |  |  | yes for motorized stage
| 1 | Y | PiezoElectricStage |  |  | This describes an optional piezoelectric-driven sample nano-positioning device associated with this Instrument. These devices typically achieve higher precision, accuracy and speed with respect to traditional Mechanical Stages. | [7 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions.
THIRD | 1 | Y | FocusStabilizationDevice |  | Manufacturer Specs | This element describes an optional focus stabilization device (also known as, Auto Focus) that may be used for bringing an object automatically into focus for this instrument. Such devices can be controlled either via an imaging software and/or by a hardware device that detects and adjusts the objective or sample positions. |  |  | 0 … ∞ |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. Changed Named and Description |  |  |  | Change to FocusStabilizationDevice? |  | Definite Focus ? (for time lapse)
| 1 |  | FocusStabilizationDevice | ID |  | A Unique Identifier for this component. | AutoFocusID |  | R
| 1 |  | FocusStabilizationDevice | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | FocusStabilizationDevice | Mechanism | FocusStabilizationType | This field refers to the mechanism employed by this Focus Stabilization Device. Software (also referred to as Image) based methods rely on the Acquisition of a Z-series of images followed by image analysis to determine the best focal position based primarily on contrast comparisons. Hardware based methods rely on an external light beam (typically infrared) that is reflected off the coverslip and whose reflection is measured to determine and correct the focal offset of the sample. | enum | Hardware based, Software (aka Image) based, Other | R |  |  |  |  | CHANGED FROM OPTICAL TO HARDWAREW BASED; IMPROVE DESCRIPTION; Nathalie will look at how ther Image based version is implemented on her scopes;
| 1 | Y | FocusStabilizationDevice | Algorithm |  | In case this Focus Stabilization Device is Software Based, this field refers to the name of the Algorithm used to maintain Focus. | string
| 1 | Y | FocusStabilizationDevice | TwoStep |  | This field records whether this Focus Stabilization Device utilizes a 2-step progressive procedure (True) or not (False). | boolean
| 1 | Y | FocusStabilizationDevice | Wavelength |  | This field refers to the Wavelength of the beam that is relevant for this Focus Stabilization Device. For example this might represent the Wavelength of the beam that is reflected against the Coverglass to obtain a measurement of the focal offset of the sample at a given time during Acquisition. | float with unit:WavelengthOfReflectedBeam | [6 empty] | ADDED NEED BETTER DESCRIPTION
| 1 |  | FocusStabilizationDevice | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
THIRD | 1 | Y | OpticsHolder-APE |  | Manufacturer Specs | This element represents a Holder for Filters, Dichroic and other Optical Element that is associated with this Instrument. These Holders come in different arrangements such as Cubes, Sliders, Turrets, and Wheels. While some of these Holders, such as Filter Cubes, might be pre-assebled by the Manufacturer and pruchased as a whole, in the majority of cases they are assembled on site depending on the needs of the user. | Substitutions: | FilterCube, FilterTurret, FilterWheel, OpticsTurret | 0 … ∞ | was FilterSet |  | Element |  | CHANGED NAME |  |  |  | Generalized to OpticsHolder with Children FilterCube, FilterWheel, FilterSlider and OpticsHolder
| 1 |  | OpticsHolder-APE | ID |  | A Unique Identifier for this component. | LSID |  | R | same | R | LSID
| 1 |  | OpticsHolder-APE | Name |  | A User-defined Name for this component.
| 3 |  | OpticsHolder-APE | Motorized |  | This field describes whether this device is Motor (True) or Manually (False) operated. | boolean | [6 empty] | ADDED NEED BETTER DESCRIPTION
| 3 |  | OpticsHolder-APE | ResponseTime |  | This field records the time it takes for this Optics Holder to switch from one position to another, according to the Manufacturer. | float with unit:ResponseTimeUnit | [6 empty] | ADDED NEED BETTER DESCRIPTION
| 3 |  | OpticsHolder-APE | NrOfSlots |  | This field records how many Slots this Optics Holder has. | integer with unit:none |  | R
| 3 |  | OpticsHolder-APE | FilterCube |  | This field records whether or not this holder holds Filter Cubes (True) or individual Filters (False). | boolean |  | R
| 1 |  | OpticsHolder-APE | AnnotationRef | CommentAnnotation | This field is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 | Y | FilterCube |  |  | This element represents a Filter Cube containing the filters and mirror used in epi-fluorescence microscopy to separate fluorescence excitation and emission light. A Filter Cube is typically composed of one Excitation, one Emission, and one Dichroic Filter. It might be pre-assebled by the Manufacturer and pruchased as a whole or it might be custom assembled on site depending on the needs of the user. | - |  | - | FilterSet | -
| 1 |  | FilterCube | ExcitationFilterRef | Reference | This field refers to an available Excitation Filter, which is part of this Filter Cube. | Extension of Reference |  | 0 ... 1 | ExcitationFilterRef | 0 ... ∞ | Extension of Ref
| 1 |  | FilterCube | StandardDichroicRef | Reference | This field refers to an available Dichroic Mirror, which is part of this Filter Cube. | Extension of Reference |  | 1 | DichroicRef | 0 ... 1 | Extension of Ref |  | CHANGED TO REQUIRED
| 1 |  | FilterCube | EmissionFilterRef | Reference | This field refers to an available Emission Filter, which is part of this Filter Cube. | Extension of Reference |  | 0 ... 1 | EmissionFilterRef | 0 ... ∞ | Extension of Ref
| 2 |  | FilterCube | OpticsHolderPosition |  | When relevant, this field records the position of this Filter Cube in a Optics Holder. | positiveInteger with unit:none
| 1 |  | FilterCubeTurret |  |  | This element represents a circular holder for multiple Filter Cubes. | [7 empty] | CHANGED TO FILTERCUBETURRET
| 2 |  | FilterCubeTurret | FilterCubeRef |  | This field refers to an available FilterCube, which is housed in this Optics Holder. | Extension of Reference |  | 1 … ∞
| 1 | Y | FilterWheel |  |  | This element represents a circular holder for multiple individual Filters or Dichroics.
| 2 |  | FilterWheel | ExcitationFilterRef | Reference | This field refers to an available Excitation Filter, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | FilterWheel | StandardDichroicRef | Reference | This field refers to an available Dichroic Mirror, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞
| 2 |  | FilterWheel | EmissionFilterRef | Reference | This field refers to an available Emission Filter, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 1 | Y | FilterSlider |  |  | This element represents a linear holder for either Filter Cubes or individual Filers.
| 2 |  | FilterSlider | ExcitationFilterRef | Reference | This field refers to an available Excitation Filter, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | FilterSlider | StandardDichroicRef | Reference | This field refers to an available Dichroic Mirror, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞
| 2 |  | FilterSlider | EmissionFilterRef | Reference | This field refers to an available Emission Filter, which is part of this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | FilterSlider | FilterCubeRef |  | This field refers to an available FilterCube, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞
| 2 |  | OpticsTurret |  |  | This element represents a circular holder for multiple optical elements.
| 2 |  | OpticsTurret | LensRef |  | This field refers to an available Lens, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | OpticsTurret | MirroringDeviceRef |  | This field refers to an available Mirrorig Device, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | OpticsTurret | OpticalApertureRef |  | This field refers to an available Optical Aperture, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED |  |  |  | DO WE NEED TO REFERENCE OTHER ITEMS?
| 2 |  | OpticsTurret | PrismRef |  | This field refers to an available Prism, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
| 2 |  | OpticsTurret | PolarizationOpticsRef |  | This field refers to an available Polarization Optics, which is housed in this Optics Holder. | Extension of Reference |  | 0 … ∞ |  |  |  |  | ADDED
THIRD | 1 | Y | Filter-APE |  | Manufacturer Specs | This element describes an optical Filter positioned in the Light Path for one or more Channels available for this Instrument. An optical Filter is a device designed to selectively control the Wavelengths, color temperature, vibration direction, and/or intensity of the radiation which it transmits or reflects. An Instrument may have several Filters of different types, such as an Excitation, Emission, Neutral Density or an additional Generic Filter positioned in the Light Path for different purposes. For each fluorescence Image there should be at least one Filter element specified per Channel. This element can be associated with an optional Optics Holder, such as a Filter Cube, Filter Turret, Filtr Wheel, or Filter Slider. Note: Filter Holder is not the same as a Filter Set. | - |  | 1 ... ∞ | same | 1 ... ∞ | - |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 1 | Y | Filter-APE | ID |  | A Unique Identifier for this component. | FIlterID |  | R | same | R | LSID
| 1 | Y | Filter-APE | Type | FilterType | This field captures the Type of Filter used in this case. | enum | Bandpass, BandStop, Multipass, Notch, Longpass, ShortPass, Other | R | same |  | enum |  | Removed ND , Heat and Tuneable filter |  |  |  | Removed ND , Heat and Tuneable filter; Added Edge
| 2 |  | Filter-APE | FilterHolderPosition |  | This field records the position of this element in a Filter Holder of any kind, such as a Filter Slider, Turret or Wheel. | integer with unit:none |  |  | FilterWheel |  | string |  | Changed to FilterHolderPostion; has to change description accordingly; FilterHolderRef was left out as there is a reference in the FilterHolder to the Filter | Add a Ref to a FilterHolder-APE? |  |  | PROPOSE TO REMOVE THIS AND HAVE A GENERIC FILTER HOLDER APE THAT WOULD SUPERSEED FILTER SET
| 1 |  | Filter-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same | 0 ... ∞ | Extension of Ref
| 1 |  | Filter-APE | TransmittanceRange |  | This element records the Range (also known as Band) of Wavelengths that are transmitted by this optical component. It also records the maximum amount of light transmitted. Multiple of these elements can be used. | Element |  | 1 ... ∞ | same | 0 ... 1 | Element
| 1 |  | Filter-APE | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | Filter-APE | CoatingMethod | CoatingMethod | This field records the type of Coating used in this case. An optical coating is one or more thin layers of material deposited on an optical component such as a Lens, Filter or Mirror, which alters the way in which the optic reflects and transmits light. | enum | HardCoated, SoftCoated, Other | R |  |  |  |  | Added; Changed to enum | How about naming this FabricationMethod and allowing an enum for Dielectric, HardCoated, EBeam, Other. Makes it easy to add new FabricationMethods when they are developed. |  |  | Added
| 1 | Y | Filter-APE | LightPathLocation |  | This field records the location of this component in the Microscope Light Path. The most common positions are: Infinity Space, Camera Port or Illumination Port. | enum | Infinity space, Camera Port, Illumination port, Other | R
| 2 |  | Filter-APE | Technology | FilterTechnologyType | This field records the filtering Technology utilized for this optical element. The main types of technologies used in fluorescence analysis are Absorption (aka colored) glass, Polymer or Thin-film coating. | enum | Absorption Colored Glass, Metallic-coating, Polymer-Based, Thin-film interference-coating, Other | [5 empty] | Added |  |  |  | Added
| 2 |  | Filter-APE | AttenuationCoefficient |  | This field records the Attenuation Coefficient of this Filter. The Attenuation Coefficient (aka blocking coefficient) is a material property that characterizes the degree of transparency of a Lens. The Attenuation Coefficient is often defined in units of optical density (OD), whereas OD= -log10(T), and T=Transmittance (0≤T≤1). | float with unit:none |  | R |  |  |  |  | Added |  |  |  | Added
| 3 |  | Filter-APE | AttenuationMethod | AttenuationMethodType | This field records the mean by which this optical element achieves out-of-band transmission attenuation. | enum | Absorption, Reflection, Refraction, Other | R |  |  |  |  | Added |  |  |  | Added
| 3 |  | Filter-APE | Diameter |  | This field describes the Diameter of this optical element. | float with unit:DiameterUnit [default mm] | [6 empty] | Added; Changed from Clear Aperture |  |  |  | Added
| 3 |  | Filter-APE | Thickness |  | This field describes the Thickness of this optical element. | float with unit:ThicknessUnit | [6 empty] | Added
| 3 |  | Filter-APE | AngleOfIncidence |  | This field records the optimal angle of incidence (AOI) for this Filter. The AOI is the angle between the optical axis of the incident light and the axis normal to the surface of the Mirror. In specular reflection, AOI is by definition equal to the Angle Of Reflection (AOR). | float with unit:AngleOfIncidenceUnit | [6 empty] | Added |  |  |  | Added
| 3 |  | Filter-APE | Polarization | FilterPolarizationType | This field records the polarization effect of this optical element. Any thin-film interference Coating that is used at non-normal angles-of-incidence will cause some amount of polarization, the precise effect varying greatly with wavelength and with the particular Coating design. | enum | Circular, Elliptical, Planar, Other | [5 empty] | Added |  |  |  | Added
| 3 |  | Filter-APE | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0... ∞
| 1 | Y | ExcitationFilter |  |  | An optical Filter is a device designed to selectively control the Wavelengths, color temperature, vibration direction, and/or intensity of the radiation which it transmits or reflects. This element refers to an Excitation Filter whose purpose is to selectively attenuate the Light that is utilized to Excite the Sample depending on Wavelength. This element can be associated with an optional Filter Holder.
| 1 | Y | EmissionFilter |  |  | An optical Filter is a device designed to selectively control the Wavelengths, color temperature, vibration direction, and/or intensity of the radiation which it transmits or reflects. This element refers to an Emission Filter whose purpose is to selectively attenuate the Fluorescent Light that is emitted by the Sample depending on Wavelength. This element can be associated with an optional Filter Holder.
| 1 | Y | NeutralDensity |  |  | An optical Filter is a device designed to selectively control the Wavelengths, color temperature, vibration direction, and/or intensity of the radiation which it transmits or reflects. This element refers to an optical Filter whose purpose is to attenuate the intensity of the transmitted Light independent of the Wavelength. |  |  | 0 ... ∞ |  |  |  |  | Added ND filter
| 2 |  | NeutralDensity | NDType | NDType | This field refers to the general category of this Neutral Density Filter. Neutral Density Filters come in two general types. Those made of darkened glass that absorb light and those coated with a metal film that reflect light. | enum | Absorbing, Reflecting, Other | R | [8 empty] | PROPOSE TO MAKE IT TIER 1
| 3 |  | NeutralDensity | OpticalDensity |  | This field measures the absolute value of the logarithm of base 10 of the power transmission factor. For example, an optical density of 3 means an attenuation of optical powers by the factor 10^3 = 1000. | float with unit:none |  | R |  |  |  |  | Changed to OpticalDensity | Call it OpticalDensity instead?
| 1 | Y | GenericFilter |  |  | An optical Filter is a device designed to selectively control the Wavelengths, color temperature, vibration direction, and/or intensity of the radiation which it transmits or reflects. This element refers to a Generic Filter positioned along the Light Path in addition to the canonical fluorescence Excitation and the Emission Filters. | [7 empty] | Changed Name to Generic Filter; Changed description | How about calling this simply "Filter"? |  |  | Issue with calling it filter is that people might get confused and think this is the canonical Excitation or Emission filter.
THIRD | 1 | Y | MirroringDevice-APE |  | Substitution : choose one child from the list below | This element describes a Mirror positioned in the Ligth Path for one or more Channels available for this Instrument. An Instrument may have several Mirrors of different types, such as Dichroic, Beamsplitter or Reflector. This element can be associated with an optional Optics Holder, such as a Filter Cube, Filter Turret, Filtr Wheel, or Filter Slider. | Substitution : choose one child from the list below | Dichroic, Beamsplitter, Reflector | 1 ... ∞ | was Dichroic element |  | Element
| 1 |  | MirroringDevice-APE | ID |  | A Unique Identifier for this component. | MirrorID |  | R | same | R | LDIS
| 1 |  | MirroringDevice-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same | 0 ... ∞ | Extension of Ref
| 1 |  | MirroringDevice-APE | Name |  | A User-defined Name for this component. | Denomination
| 3 |  | MirroringDevice-APE | MirrorType | MirrorType | This field records the general category of this Mirror. | enum | Convex, Concave, Planar, Other | R
| 3 |  | MirroringDevice-APE | Diameter |  | This field describes the Diameter of this optical element. | float with unit:DiameterUnit | [6 empty] | Moved from Dichroic because it is a general Attribute. |  |  |  | Added
| 3 |  | MirroringDevice-APE | AngleOfIncidence |  | This field records the optimal angle of incidence (AOI) for this Mirror. The AOI is the angle between the optical axis of the incident light and the axis normal to the surface of the Mirror. In specular reflection AOI is by definition equal to the Angle Of Reflection (AOR). | float with unit:AngleOfIncidenceUnit
| 3 |  | MirroringDevice-APE | Geometry | MirrorGeometry | This field records the geometrical shape of the surface(s) of this Mirror. | enum | Aspherical, Cylindrical, Ellipsoidal, Hyeprboloidal, Paraboloidal, Spherical, Other
| 3 |  | MirroringDevice-APE | RadiusOfCurvature |  | The radius of curvature of this Mirror. A spherical lens or Mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. | float with unit:RadiusOfCurvatureUnit
| 3 |  | MirroringDevice-APE | SubstrateType |  | This field records the type of Substrate that is used for this Mirror. It is typically a ground and polished piece of optical glass. | string
| 3 |  | MirroringDevice-APE | SubstrateMaterial | SubstrateMaterialName | This field records the polarization effect of this optical element. Any thin-film interference Coating that is used at non-normal angles-of-incidence will cause some amount of polarization, the precise effect varying greatly with wavelength and with the particular Coating design. | enum | Calcium Fluoride, N-BK7 (Borosilicate Crown), UV-Grade Fused Silica (UVFS), Zerodur (Lithium Aluminum Ailicon Oxide), Zinc Selenide (ZnSe), Other
| 3 |  | MirroringDevice-APE | GlassCode |  | If available, this field records the international Glass Code of the material that constitutes this Mirror. The Glass Code designation is based on U.S. military standard MIL-G-174, and it consists of a six-digit number specifying the glass according to its Refractive Index nd at the Fraunhofer d- (or D3-) line, and its Abbe Number Vd also taken at that line. The resulting Glass Code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit Glass Code of 517642. | string
| 2 |  | MirroringDevice-APE | OpticsHolderPosition |  | This field records the position of this Mirroring Device in a Optics Holder of any kind, such as a Slider, Turret or Wheel. | postitiveInteger with unit:none | [6 empty] | Changed to FilterHolderPostion; changed Description accordingly; FilterHolderRef was left out as there is a reference in the FilterHolder to the Filter; MOVED TO MIRRORING DEVICE (was in Dichroic)
| 1 | Y | Dichroic-APE |  |  | A Dichroic Mirror (also often referred to as Dichroic Beamsplitter, Dichromatic Mirror or Dichroic) is a Mirror that spectrally separate light by transmitting and reflecting light as a function of wavelength. In the context of fluorescence microscopy, they are typically used to reflect the excitation light, but transmit the emitted fluorescence. Dichroic Mirrors can be categorized in different manners. However, the most important criteria for choosing a Dichoric Mirror are the Transmittance and the Reflectance Wavelength ranges.
| 1 | Y | Dichroic-APE | Type | FilterType | This field captures the Type of Dichroic used in this case. | enum | Bandpass, BandStop, Multipass, Notch, Longpass, ShortPass, Other | R |  |  |  |  | Removed ND , Heat and Tuneable filter
| 1 | Y | Dichroic-APE | CoatingMethod | CoatingMethod | This field records the type of Coating used in this case. An optical coating is one or more thin layers of material deposited on an optical component such as a Lens, Filter or Mirror, which alters the way in which the optic reflects and transmits light. | enum | HardCoated, SoftCoated, Other | R |  |  |  |  | Added; Changed to enum
| 1 | Y | Dichroic-APE | LightPathLocation | LightPathLocationOptions | This field records the location of this component in the Microscope Light Path. The most common positions are: Infinity Space, Camera Port or Illumination Port. | enum | Infinity space, Camera Port, Illumination port, Other | R
| 2 |  | Dichroic-APE | FilterHolderPosition |  | This field records the position of this element in a Filter Holder of any kind, such as a Filter Slider, Turret or Wheel. | integer with unit:none | [6 empty] | Added
| 2 |  | Dichroic-APE | Technology | FilterTechnologyType | This field records the filtering Technology utilized for this optical element. The main types of technologies used in fluorescence analysis are Absorption (aka colored) glass, Polymer or Thin-film coating. | enum | Absorption Colored Glass, Metallic-coating, Polymer-Based, Thin-film interference-coating, Other | [5 empty] | Added
| 2 |  | Dichroic-APE | AttenuationCoefficient |  | This field records the Attenuation Coefficient of this Dichroic. The Attenuation Coefficient (aka blocking coefficient) is a material property that characterizes the degree of transparency of a Lens. The Attenuation Coefficient is often defined in units of optical density (OD), whereas OD= -log10(T), and T=Transmittance (0≤T≤1). | float with unit:none |  | R |  |  |  |  | Added
| 3 |  | Dichroic-APE | AttenuationMethod | AttenuationMethodType | This field records the mean by which this optical element achieves out-of-band transmission attenuation. | enum | Absorption, Reflection, Refraction, Other | R |  |  |  |  | Added
| 3 |  | Dichroic-APE | Thickness |  | This field describes the Thickness of this optical element. | float with unit:ThicknessUnit | [6 empty] | Added
| 3 |  | Dichroic-APE | Polarization | FilterPolarizationType | This field records the polarization effect of this optical element. Any thin-film interference coating that is used at non-normal angles-of-incidence will cause some amount of polarization, the precise effect varying greatly with wavelength and with the particular coating design. | enum | Circular, Elliptical, Planar, Other | [5 empty] | Added
| 3 |  | Dichroic-APE | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0... ∞
| 3 |  | Dichroic-APE | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0... ∞
| 1 |  | Dichroic-APE | TransmittanceRange |  | This element records the Range (also known as Band) of Wavelengths that are transmitted by this optical component. It also records the maximum amount of light transmitted. Multiple of these elements can be used. | Element |  | 1 … ∞
| 3 |  | Dichroic-APE | ReflectanceRange |  | This element records the Range (also known as Band) of Wavelengths that are reflected by this optical componewnt. It also records the maximum amount of light reflected. Multiple of these elements can be used. | Element |  | 0 … ∞
| 1 | Y | StandardDichroic |  |  | A Dichroic Mirror (also often referred to as Dichroic Beamsplitter or simply as Dichroic) is a Mirror that spectrally separate light by transmitting and reflecting light as a function of wavelength. In the context of fluorescence microscopy, they are typically used to reflect the excitation light, but transmit the emitted fluorescence. Dichroic Mirrors can be categorized in different manners. However, the most important criteria for choosing a Dichoric Mirror are the Transmittance and the Reflectance Wavelength ranges. This element describes a Dichroic positioned in the standard Ligth Path for one or more Channels available for this Instrument. | [7 empty] | STET | Is this not superfluous? Why not create one or more Dichroics as needed? |  |  |  |  | used as beam splitters for 2-cam Acquisition?
| 1 | Y | AdditionalDichroic |  |  | A Dichroic Mirror (also often referred to as Dichroic Beamsplitter or simply as Dichroic) is a Mirror that spectrally separate light by transmitting and reflecting light as a function of wavelength. In the context of fluorescence microscopy, they are typically used to reflect the excitation light, but transmit the emitted fluorescence. Dichroic Mirrors can be categorized in different manners. However, the most important criteria for choosing a Dichoric Mirror are the Transmittance and the Reflectance Wavelength ranges. This element describes a generic Additional Dichroic positioned in the Ligth Path for one or more Channel available for this Instrument. | [7 empty] | STET | Superfluous?
| 2 |  | Beamsplitter |  |  | This element describes a Mirror that reflects part of the light, regardless of Wavelength. | [14 empty] | It can be a half-silvered mirror, but usually it consists of two glass prisms
| 2 |  | Beamsplitter | Transmittance |  | The field records the maximum amount of light this Beamsplitteris capable of transmitting at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R | [11 empty] | Add Manufacturer, Serial Number and Model as well
| 3 |  | Beamsplitter | TransmissionAngle |  | The angle at which light is transmitted by this Beamsplitter. | float with unit:TransmittanceAngleUnit
| 2 |  | Beamsplitter | Reflectance |  | The field records the maximum amount of light this Beamsplitteris capable of reflecting at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R
| 3 |  | Beamsplitter | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0... ∞
| 3 |  | Beamsplitter | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0... ∞
| 2 |  | ReflectingMirror |  |  | This element describes a total reflection Mirror.
| 2 |  | ReflectingMirror | Reflectance |  | The field records the maximum amount of light this Mirror is capable of reflecting at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R
| 3 |  | ReflectingMirror | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0... ∞
| 1 | Y | TransmittanceRange |  |  | This element records the range of Wavelengths that are transmitted by this optical component. It also records the maximum amount of light transmitted. Multiple of these elements can be used. |  |  | 1 … ∞ | CutIn |  | positiveFloat
| 2 |  | TransmittanceRange | Transmittance |  | This field records the maximum amount of light this optical component is capable of transmitting in this Wavelength Range, at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | percentFraction with unit:none |  |  | same |  | percentFraction
[10 empty] | CutOut |  | positiveFloat | [9 empty] | Wavelength
[10 empty] | CutInTolerance |  | nonNegativeFloat | [9 empty] | FWHMBandwidth
[10 empty] | CutOutTolerance |  | nonNegativeFloat | [9 empty] | Wavelength
| 1 | Y | TransmittanceRange | Wavelength |  | This field records the key Wavelength] that characterizes this Transmittance Range (also known as Transmission band). In the case of Bandpass, Bandstop, Multipass or Notch Filters, Wavelength often refers to the Central Wavelength of each of the transmission band(s); alternatively, in case of a Longpass or a Shortpass Filters, Wavelength refers to the CutOn Wavelength (i.e., below which there is less than 50% transmittance) or the CutOff Wavelengths (i.e., above which there is less than 50% transmittance) respectively. In the case of Dichroics, Wavelength most commonly refers to the Edge Wavelength which might be either the CutOn or the CutOff Wavelength depending on individual cases. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED | Good question: I looked at how the various Manufacturers (Chroma, Semrock, Omega, Edmund, Newport, etc.) describe their filters and most of them actually use the Center Wavelength and Full-Width Half Max. Maybe better to switch to that? |  |  | Should we change Change to CutOn CutOff? | agreed with center with FWHM
| 1 | Y | TransmittanceRange | FWHMBandwidth |  | This field records the width of this Transmittance Range (also known as Transmission band). Specifically, the Full Width at Half Maximum (FWHM) represents the bandwidth of the Wavelength spectrum curve measured between those points whose transmission amplitude is 50% of the maximum. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED
| 3 |  | ReflectanceRange |  |  | This element records the range of Wavelengths that are reflected by this optical componewnt. It also records the maximum amount of light reflected. Multiple of these elements can be used. |  |  | 0 … ∞
| 3 |  | ReflectanceRange | Wavelength |  | This field records the key Wavelength that characterizes this Reflectance Range (also known as Reflection band). In most cases, Wavelength refers to the CutOn Wavelength (i.e., below which there is less than 50% transmittance) but it might also refer to the CutOff wavelength (i.e., above which there is less than 50% transmittance) or the Center Wavelength of the Reflectance Band. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED
| 3 |  | ReflectanceRange | FWHMBandwidth |  | This field records the width of this Reflectance Range (also known as Reflection band). Specifically, the Full Width at Half Maximum (FWHM) represents the bandwidth of the Wavelength spectrum curve measured between those points whose transmission amplitude is 50% of the maximum. | positiveFloat with unit:WavelengthUnit |  | R |  |  |  |  | CHANGED
| 3 |  | ReflectanceRange | Reflectance |  | This field records the maximum amount of light this optical component is capable of reflecting in this Wavelength Range, at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | percentFraction with unit:none
THIRD | 1 | Y | Lens-APE |  | Manufacturer Specs | This element describes additional lenses positioned in the Light Path for one or more Channels available for this Instrument. A Lens is a component of glass or transparent plastic material, usually circular, with two primary surfaces that are ground and polished in a specific manner designed to produce either a convergence or divergence of light passing through the material. An Instrument may have several Lenses with different functions, such as a Magnification Changer (also known as Optovar), a Coupling Lens connecting the Light Source with the Microscope Stand, a Relay Lens positioned at the Detector mount, or additional Lenses placed in the Light Path for various purposes. This element can be associated with an optional Optics Turret. | Substitution: | MagnificationChanger, GenericLens, BeamExpanderLens, BeamCollimatorLens, CouplingLens, RelayLens, TubeLens | 0 … ∞ |  |  |  |  | Changed Defition; Added Function to Settings
| 1 |  | Lens-APE | ID |  | A Unique Identifier for this component. | LensID |  | R
| 1 |  | Lens-APE | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | Lens-APE | Magnification |  | This field records the nominal Magnification provided by this Lens as specified by the Manufacturer. | positiveInteger with unit:none |  | R
| 1 | Y | Lens-APE | LensNA |  | This field captures the Numerical Aperture (NA) of this Lens expressed as a floating-point (real) number. NA is a number that was defined initially by Abbe to characterize Objectives and Condensers. NA can be calculated by the expression n sin alpha, where n is the Refractive Index of the medium between the Lens and the Sample, and alpha is half the Angular Aperture of the Lens. NA is a measure of how much of the illumination that leaves the specimen actually passes through the Lens. The larger the NA, the more information the lens can transmit. Expected range 0.02 - 1.5. | float with unit:none |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | Added missing Aperture
| 2 |  | Lens-APE | Infinity |  | This field records whether (True) or not (False) this Lens is infinity corrected. An Infinity-corrected Lens is fixed for an infinite sample to primary image distance. It needs to be used with a Tube Lens of appropriate focal length to obtain the specified magnification. | boolean |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | Added missing
| 2 |  | Lens-APE | RefractiveIndex |  | This field records the Refractive Index of the material this Lens is made of, as specified by the Manufacturer. The Refractive Index is the ratio of the speed of light (more precisely, the phase velocity) in a vacuum versus that in a given medium. | float with unit:none |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 2 |  | Lens-APE | WorkingDistance |  | This field reports the Working Distance of this Lens under normal operating conditions. The Working Distance is the distance (in air, or in the specified immersion liquid) between the front of an Objective and the surface of the cover glass or of the specimen if uncovered. | float with unit:WorkingDistanceUnit |  | R |  |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | Changed key term and Description
| 3 |  | Lens-APE | Type | LensType | This field records the general category to which this Lens belongs. | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other | R
| 3 |  | Lens-APE | Composite |  | This field records whether this Lens is Composite (True) or Simple (False). | boolean |  | R
| 3 |  | Lens-APE | Geometry | LensGeometry | This field records the geometrical shape of this Lens' surface(s). | enum | Spherical, Aspherical, Cylindrical, Ball, Hemispheric-ball, Cone, Drum, Rod, Toric | R
| 3 |  | Lens-APE | AttenuationCoefficient |  | This field records the Attenuation Coefficient of this Lens. The Attenuation Coefficient (aka blocking coefficient) is a material property that characterizes the degree of transparency of a Lens. The Attenuation Coefficient is often defined in units of optical density (OD), whereas OD= -log10(T), and T=Transmittance (0≤T≤1). | float with unit:none |  | R
| 3 |  | Lens-APE | ImageDistance |  | This field records the distance in air between the rear principal plane of this Lens and the primary Image plane. | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED |  |  |  | ADD?
| 3 |  | Lens-APE | ObjectDistance |  | This field records the distance in air between the principal front plane of this Lens and the field plane in which the object (i.e., specimen) is situated. | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED |  |  |  | ADD?
| 3 |  | Lens-APE | FrontFocalLength |  | This field records the distance (f) measured along the optical axis from the front principal plane of this Lens to the front focal plane (F). | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED |  |  |  | Change description to incorporate ISO 10934-1-2002 standard definitions?
| 3 |  | Lens-APE | BackFocalLenght |  | This field records the distance (f') measured along the optical axis from the rear principal plane of this Lens to the rear focal plane (F'). | float with unit:OpticalDistanceUnit; [default mm] | [6 empty] | ADDED |  |  |  | ADD?
| 3 |  | Lens-APE | Diameter |  | This field records the physical diameter of this Lens. | float with unit:DiameterUnit [default mm]
| 3 |  | Lens-APE | AbbeNumber |  | This field records the Abbe number of this Lens. The Abbe Number is a measure of the material's dispersion (variation of Refractive Index versus Wavelength), with high values indicating low dispersion. | float with unit:none
| 3 |  | Lens-APE | RadiusOfCurvature |  | This field records the Radius of Curvature of this Lens. A spherical Lens or Mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the Radius of Curvature of the surface. | float with unit:RadiusUnit
| 3 |  | Lens-APE | MaterialName | LensMeterialName | This field records the chemical composition or the nonproprietary trade designation of the Material this Lens is made of. | enum | alpha-Barium Borate, N-BK7 (Borosilicate Crown), Calcium Fluoride , UV-Grade Fused Silica (UVFS), Germanium, Magnesium Fluoride, N-SF11 (Dense Flint), Polytetrafluoroethylene (PFTE), Silicon (Si), Zinc Selenide (ZnSe), Other
| 3 |  | Lens-APE | Coating |  | An optical Coating consists of one or more thin layers of material deposited on an optical component such as a Lens or Mirror, which alters the way in which the optic reflects and transmits light. This field records the optical Coating used for this Lens. | string | [6 empty] | ADDED
| 3 |  | Lens-APE | GlassCode |  | If available, this field records the international Glass Code of the material that constitutes this Lens. The Glass Code designation is based on U.S. military standard MIL-G-174, and it consists of a six-digit number specifying the glass according to its Refractive Index nd at the Fraunhofer d- (or D3-) line, and its Abbe Number Vd also taken at that line. The resulting Glass Code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit Glass Code of 517642. | sring
| 3 |  | Lens-APE | Density |  | This field records the Density of the material this Lens is made of. | float with unit:DensityUnit
| 1 |  | Lens-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 | Y | MagnificationChanger |  |  | A Magnification Changer (also known as Optovar because one of the first magnification changing devices was developed by Zeiss and given this proprietary Name) provides additional magnification that can be beneficial when a desired field of view is required for the imaging sensor. Magnification Changers are placed before the Tube Lens within the infinity space of the optical system. Typical magnification increments are 1.25x, 1.6x, and 2x. | [7 empty] | THIS WAS CALLED OPTOVAR AND WAS A SEPARATE ENTITY;
| 2 |  | AdditionalLens |  |  | This element refers to any generic Additional Lens positioned in the Light Path for a given Channel. | [7 empty] | Changed to Name to GenericLens; ADDED FUNCTION in Settings | [5 empty] | perhaps switching lens and adaptor lens from spinning disk head can be included here.
| 2 |  | BeamExpanderLens |  |  | A Beam Expander Lens is a Lens designed to work in combination with other optical components to form a Beam Expander device. As such, its role is to increase the diameter of a collimated input beam to a larger collimated output beam. | [7 empty] | ADDED |  |  |  | Changed definition; please double check
| 2 |  | BeamCollimatorLens |  |  | A Beam Collimator Lens is a Lens designed to work in combination with other optical components to form a Beam Collimator device. As such, its role is to transform a strongly diverging light beam into a collimated beam, i.e., a beam where light propagates mainly only in one direction, and the beam divergence is weak. | [7 empty] | ADDED |  | QUESTION: DO WE NEED TO ADD A REFERENCE TO AN OPTO-MECHANICAL DEVICE? |  | Changed definition; please double check
| 1 | Y | CondenserLens |  |  | A Condenser Lens renders a divergent beam from a point source into a parallel or converging beam to illuminate an object. It acts to gather light from the Light Source and concentrate it into a cone of light that illuminates the specimen. The aperture and angle of the light cone must be adjusted (via the size of the diaphragm) for each different objective Lens with different numerical apertures. Conserser Lenses are typically housed within a specific component of the Microscope called the Condenser. | [7 empty] | ADDED |  | QUESTION: DO WE NEED A CONDENSER UNIT? |  | Changed definition; please double check
| 2 |  | CouplingLens |  |  | This element refers to a Lens positioned at the mount connecting the Light Source Coupling device and the Microscope Stand.
| 1 | Y | RelayLens |  |  | A Relay Lens is often used in combination with other Lenses for transferring an Image into a different optical plane. A typical example of Relay Lens is the Lens positioned at the mount connecting a Detector with the Microscope Stand. | [7 empty] | CHANGED DEFINITION
| 2 |  | TubeLens |  |  | A Tube Lens is an intermediate Lens, located within the Microscope's Illumination Tube, which is designed to operate as an essential component of infinity-corrected Objectives, and which should be regarded as an integral part of the Objective lens system when magnification and correction are considered. In some Microscope Tubes, Lenses might be located within the Microscope Stand and might not be immediately accessible. | [7 empty] | ADDED; Changed definition |  |  |  | Changed definition; please double check |  |  | Very important / The tube lens should get its own section like the optovar
THIRD | 1 | Y | PolarizationOptics |  |  | This element describes a Polarization device positioned in the Ligth Path for this Instrument. A Polarization device is an optical filter rotatable through 360° that lets light waves of a specific polarization pass through while blocking light waves of other polarizations. This element can be associated with an optional Optics Turret. |  |  | 0 … ∞ |  |  |  |  | Changed definition.
| 1 |  | PolarizationOptics | ID |  | A Unique Identifier for this component. | PolarizationOpticsID |  | R
| 1 |  | PolarizationOptics | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | PolarizationOptics | Function | PolarizationOpticsFunction | This field records whether this Polarization Optics is used to generate (Polarizer) or to analyze (Analyzer) polarized light. | enum | Polarizer, Analyzer, Other | [5 empty] | Removed Cross Polarizer and created Cross Polarizer field |  |  |  | Removed Cross Polarizer and created Cross Polarizer field
| 2 |  | PolarizationOptics | CrossPolarizer |  | This field records whether (True) or not (False) this Polarization Optics is a combination of an Analyzer and a Polarizer positioned orthogonal to each other. | boolean | [6 empty] | ADDED |  |  |  | PROPOSED
| 2 |  | PolarizationOptics | BeamSplitter |  | This field records whether (True) or not (False) this is PolarizationOptics has a beam splitter function. | boolean | [6 empty] | ADDED
| 2 |  | PolarizationOptics | FaradayIsolator |  | This field records whether (True) or not (False) this PolarizationOptics is a Fraday Isolator. A Faraday Isolator is an optical component which allows light to pass through in only ONE direction and therefore protects laser systems against back reflection. It uses the Faraday magneto-optical effect to turn the polarization of light between various polarizers. | boolean | [6 empty] | ADDED FaradayIsolator
| 2 |  | PolarizationOptics | DIC |  | This field records whether (True) or not (False) this PolarizationOptics is used for Differential Interference Contrast (DIC) microscopy. | boolean | [6 empty] | ADDED FaradayIsolator |  |  |  | Added
| 2 |  | PolarizationOptics | Construction | PolarizationOpticsConstruction | This field records the modality with which this PolarizationOptics is constructed. | enum | Achromatic, Zero Order, Multiple Order, Other
| 2 |  | PolarizationOptics | Retardation | PolarizationOpticsRetardation | This field records the Phase shift that exists between the polarized light component projected along the fast axis and the part projected along the slow axis. | enum | Full Wave, Half Wave, Quarter Wave, Other
| 3 |  | PolarizationOptics | Fabrication | PolarizationOpticsFabrication | This field records the way in with which this PolarizationOptics is fabricated. | enum | Birefringent, Film, Polarizing Beamsplitter, Polarizing BandPass, Polymer, Quarz, Variable, Wire Grid, Other
| 3 |  | PolarizationOptics | MaterialName | PolarizationOpticsMaterialName | This field records the chemical composition or the nonproprietary trade designation of the Material this Polarization Optics is made of. | enum | alpha-Barium Borate, Calcite, Calcium Fluoride, N-BK7 (Borosilicate Crown), Rutile (TiO2), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 3 |  | PolarizationOptics | GlassCode |  | If available, this field records the international Glass Code of the material that constitutes this Lens. The Glass Code designation is based on U.S. military standard MIL-G-174, and it consists of a six-digit number specifying the glass according to its Refractive Index nd at the Fraunhofer d- (or D3-) line, and its Abbe Number Vd also taken at that line. The resulting Glass Code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit Glass Code of 517642. | sring
| 1 |  | PolarizationOptics | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 3 |  | PolarizationOptics | WavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞ |  |  |  |  | ADDED |  |  |  | Changed cardinality
THIRD | 1 | Y | Prism |  | Manufacturer Specs | This element describes a Prism positioned in the Light Path for one or more Channels available for this Instrument. A Prism is a block of glass or other transparent material with at least two intersecting flat, polished surfaces, used to disperse light or deviate it through an angle. A Prism can be viewed as an ensemble of different rigidly mounted Mirrors, with each face having a permanent orientation with respect to others. A Prism can be employed to deflect or deviate a light beam, rotate or invert an image, separate polarization states, or disperse light into its component Wavelengths. An Instrument may have several Prisms with different functionality such as Polarizing, Reflecting, or Refracting (or Dispersing). Polarizing Prisms are birefringent crystals that divide incident non-polarized light into separate components polarized orthogonally to each other. Reflecting Prisms are useful for redirecting light beams by total internal reflection. Refracting (or Dispersing) Prisms can be employed to bend and separate light into its component colors. This element can be associated with an optional Optics Turret. |  |  | 0 … ∞ |  |  |  |  | CHANGED DEFINITION TO INCORPORATE ISO 10934 |  |  |  | There is AnnotationRef, do we need more Free Text description?
| 1 |  | Prism | ID |  | A Unique Identifier for this component. | PrismID |  | R
| 1 |  | Prism | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | Prism | DIC | DICPrismType | This field records whether this Prism is used for Differential Interference Contrast (DIC) microscopy and, if so, what DIC-compatible type it is (i.e., deSenarmont, Nomarski or Wollaston). | enum | deSenarmont, Nomarski, Wollarston, None | [5 empty] | ADDED | [5 empty] | Some use a single prims for all objectives and some individual, as Zeiss developed the Nomarski Prism, but other now use it. Nikon uses de Senarmont DIC Prism (A Normarski prisms modification). Not sure if Nikon offers only this though.
| 2 |  | Prism | Function | PrismFunction | This field records the function of this Prism. Beam-shaping (i.e., Anamorphic) Prism pairs are used for reshaping the profile of a laser beam. For example, the elliptical beam from a laser diode can be transformed into a beam with circular cross-section, by magnifying the elliptical beam in one dimension. They can also be used to convert a circular beam into an elliptical beam. Dispersing prisms can be employed to bend and separate light into its component wavelengths. Polarizing prisms (i.e., Glan–Thompson, Nomarski, Wollaston) are birefringent crystals, which divide incident non-polarized light into separate components polarized orthogonally to each other. Reflecting prisms are useful for redirecting (i.e., flip, invert, rotate, etc.) light beams by total internal reflection or for dividing them into two or more separate beams (i.e., beam-splitting) one. | enum | Beam-shaping, Deflecting, Polarizing, Reflecting, Refracting (or Disperising), Other | R |  |  |  |  | CHANGED FROM TYPE TO FUNCTION; modified description; added Dispersive and Beam-shaping |  |  |  | Added Dispersive type to accomodate prism spectral separation which is used for Confocal Microscopy
| 2 |  | Prism | Condenser |  | This field records whether this Prism is within the Condenser (True) or the Objective (False). | boolean |  | R |  |  |  |  | ADDED |  |  |  | Added
| 2 |  | Prism | OilObjective |  | This field records whether this Prism is designed to work with oil Objectives (True) or not (False). | boolean |  | R |  |  |  |  | ADDED |  |  |  | Added
| 2 |  | Prism | Compound |  | This field records whether this Prism is a Compound Prism (i.e., a set of multiple triangular prism elements placed in contact, and often cemented together to form a solid assembly; True) or a Simple Prism (False). | boolean | [6 empty] | ADDED |  |  |  | ADDED
| 3 |  | Prism | RefractiveIndex |  | This field records the Refractive Index of the material this Prism is made of, as specified by the Manufacturer. | float with unit:none |  | R
| 3 |  | Prism | Reflectance |  | This field records the maximum amount of light this Prism is capable of transmitting at peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R
| 3 |  | Prism | AngleOfIncidence |  | This field records the Angle Of Incidence (AOI) is the angle between a beam of light incident on the Prism surface and the line perpendicular to the surface at the point of incidence, called the normal. | float with unit:IncidenceAngleUnit |  | R
| 3 |  | Prism | DeviationAngle |  | This field records the Deviation Angle of this Prism. A beam passing through an object like a prism is deflected twice: once entering, and again when exiting. The sum of these two deflections is called the deviation angle and is measured with respect to the incident ray. | float with unit:DeviationUnitAngle |  | R
| 3 |  | Prism | Design | PrismDesign | This field records the shape of the Prism with specific reference to the angles defined by its faces. | enum | Rhomboid, Wedge, Glan-Focault, Glan-Taylor, Glan-Thompson, Nicol, Nomarski, Rochon, Senarmont, Wollaston, Abbe-Konig, Amici-roof, Dove, Penta, Porro, Porro-Abbe, Right-Angle, Schmidt–Pechan, Abbe, Amici, Grism, Pellin–Broca, Triangular, Other | [5 empty] | ADDED
| 3 |  | Prism | PrismAngle |  | This field records the Angle between two surfaces of this Prism. | float with unit:PrismAngleUnit
| 3 |  | Prism | AbbeNumber |  | This field records the Abbe number of this Prism. The Abbe Number is a measure of the material's dispersion (variation of Refractive Index versus Wavelength), with high values indicating low dispersion. | float with unit:none
| 3 |  | Prism | MaterialName | PrismMaterialName | This field records the chemical composition or the nonproprietary trade designation of the Material this Prism is made of. | enum | Calcium Fluoride, F2 (Flint), N-BK7 (Borosilicate Crown), N-SF11 (Dense Flint), Rutile (TiO2), UV-Grade Fused Silica (UVFS), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 3 |  | Prism | GlassCode |  | If available, this field records the international Glass Code of the material that constitutes this Prism. The Glass Code designation is based on U.S. military standard MIL-G-174, and it consists of a six-digit number specifying the glass according to its Refractive Index nd at the Fraunhofer d- (or D3-) line, and its Abbe Number Vd also taken at that line. The resulting Glass Code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit Glass Code of 517642. | sring
| 3 |  | Prism | Density |  | This field records the Density of the material this Prism is made of. | float with unit:DensityUnit
| 1 |  | Prism | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 3 |  | Prism | WavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞ |  |  |  |  | Changed cardinality |  |  |  | Changed cardinality
THIRD | 1 | Y | Detector-APE |  | Manufacturer Specs | This element describes a Detector available for this Instrument. An Instrument may have more than one Detector of different types, such as a Photomultiplier, or a Camera. The Detector ID can be used as a reference within the Channel element. The values stored in Detector represent the fixed values, variable values modified during the Acquisition go in DetectorSettings. | Substitution : choose one child from the list below | Camera, Photomultiplier | 1 … ∞ | was simple Class |  | Element | [5 empty] | APD is listed below in Point Detector | Should the detector and paramters be required? I do not see the "R" in the cardinal column... |  | APD
| 1 |  | Detector-APE | ID |  | A Unique Identifier for this component. | DetectorID |  | R | same | R | LSID
Type
Gain
AmplificationGain
Offset
Voltage
Zoom
| 1 |  | Detector-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 |  | Detector-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | Detector-APE | Mount | DetectorMount | This field records the type of Mount used to connect this detector to the Microscope Stand. | enum | C-mount, D-SLR adapter, SLR adapter, T-mount, Other | R
| 2 |  | Detector-APE | MaxBitDepth | PixelType | This field records the Maximum Bit Depth provided by the Detector's Analog to Digital (A/D) converter. | enum | bit, complex, double, double-complex, float, int8, int16, int32, uint8, uint16, uint32, Other | R |  |  |  |  | ADDED
| 2 |  | Detector-APE | QuantumEfficiency |  | This field records the Quantum Efficiency observed at Peak Wavelength, according to the Manufacturer. Quantum Efficiency is generally defined as is the ratio of incident photons to converted electrons. Typically expressed as a number between 0 and 1. | float with unit:none |  | R | [8 empty] | Changed to make it shorter (was QuntumEfficiencyAtPeakWavelength)
| 2 |  | Detector-APE | ElectronConversionFactor |  | This field records the Electron-Conversion Factor for this Detector, as specified by the Manufacturer. The Electron Conversion Factor relates the digital numbers associated with each pixel with the actual number of electrons generated by the Detector. It is expressed in e-/Digital Number (DN). Note that DN can also be referred to as Analog-to-Digital-Unit (ADU) or Binary Number. It is sometimes referred to as Digital Gain or, in case of CCD cameras as CCD Gain. | float with unit:none |  | R |  |  |  |  | CHANGED | To be addressed in Settings table, not here? Maybe record Gain Range here? |  |  | PLEASE VERIFY
| 2 |  | Detector-APE | ReadOutNoise |  | This field captures the Read Out Noise associated with this Detector, as specified by the Manufacturer. Read Out Noise is a constant amount of noise and it includes thermal noise build-up and clock induced charge added by CCD chip, the readout amplifier and surrounding electronics. It is generally expressed in photo-electrons (i.e., e- root-mean-square). | float with unit:ReadNoiseUnit |  | R |  |  |  |  | CHANGED
| 2 |  | Detector-APE | DetectorNoiseModel | DetectorNoiseModel | This field records the Noise Model associated with this Detector. | enum | Gaussian, Poisson, Other | R
| 2 |  | Detector-APE | DarkCurrentRate |  | This field records the Dark Current Rate of this Detector as specified by the Manufacturer. The Dark Current Rate is the median current detected in the absence of light. Generally expressed in e-/pixel/sec. | float with unit:DarkCurrentRateUnit |  | R
| 2 |  | Detector-APE | DynamicRange |  | This field records the Dynamic Range of this Detector in deciBel (dB). | float with unit:DynamicRangeUnit; [default dB] | [6 empty] | ADDED | Added. Or maybe move this to Detector-APE?
| 3 |  | Detector-APE | AmbientOperatingTemperature |  | This field records the Ambient Temperature at which this Detector is designed to operate, according to the Manufacturer. | float with unit:TemperatureUnit; [default: Celsius ºC]
| 3 |  | Detector-APE | AmbientOperatingHumidity |  | This field records the Ambient Humidity at which this detector is designed to operate, according to the Manufacturer. | float with unit:none
| 1 | Y | Detector-APE | RelayLensRef |  | This field contains a reference to a Lens positioned at the Mount connecting this Detector to the Microscope Stand. | Extension of Reference |  | 0 … 1
| 2 |  | Detector-APE | WavelengthRange |  | This element records the Range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞
| 1 | Y | Camera-APE |  | Substitution : choose one child from the list below | This element describes a Camera available for this Instrument. A Camera is a device that converts optical brightness into an electrical signal using a grid of pixel sensor unit cells and then reproduces an image of the specimen. An Instrument may have more than one camera of different types, such as traditional CCD, EMCCD, Intensified EMCCD, sCMOS, or Analog Video. |  | CCD, EMCCD, AnalogVideo, CMOS, Intensified CCD, Intensified CMOS
| 1 | Y | Camera-APE | ArrayWidth |  | This field records the Width of this Camera chip or sensor area, expressed in Pixel. | integer with unit:none
| 1 | Y | Camera-APE | ArrayHeight |  | This field records the Height of this Camera chip or sensor area, expressed in Pixel. | integer with unit:none
| 1 | Y | Camera-APE | PixelWidth |  | This field records the Physical Width of each Camera Pixel. | float with unit:PixelWidthUnit; default:µm |  | R | [8 empty] | PROPOSE TO CHANGE DESCRIPTION
| 1 | Y | Camera-APE | PixelHeight |  | This field records the Physical Height of each Camera Pixel. | float with unit:PixelHeightUnit; default:µm |  | R | [8 empty] | PROPOSE TO CHANGE DESCRIPTION
| 2 |  | Camera-APE | Illumination | CameraIllumination | This field records the type of illumination the Camera chip receives. | enum | Back, Front, Other | R
| 2 |  | Camera-APE | ManufacturerOffset |  | This field reports the Camera Offset when it is pre-set by the Manufacturer. The Offset is an electronic baseline added to the output signal from the sensor to ensure that the displayed signal level is always a positive number of counts. | float with unit:none | [6 empty] | ADDED |  |  |  | CALLED PRESET OFFSET |  | Offset?
| 2 |  | Camera-APE | Color | ColorMode | This field records whether this is a Color Camera and what type it is. | enum | none, 3-Sensor, Bayer, CYGM, MultiLayer, Other | [5 empty] | ADDED
| 3 |  | Camera-APE | PixelWellCapacity |  | This field records the maximum amount of charge each Pixel Well can hold, measured in electron counts. | integer with unit:PixelWellCapacity |  | R
| 2 |  | Camera-APE | MaximumFrameRate |  | This field records the maximal rate at which this Camera can acquire a single full-frame. | float with unit:MaximumFrameRateUnit ; [default fps] |  | R
| 3 |  | Camera-APE | MaximumReadoutRate |  | This field records the maximal pixel Readout Rate allowed by this Camera. | float with unit:MaximalReadoutRateUnit; [default pixel per second] |  | R
| 3 |  | Camera-APE | VerticalClockSpeed |  | This field records the maximal speed at which each vertical column of pixels in the camera chip can be shifted down and read. | float with unit:VerticalClockSpeedUnit
| 1 | Y | AnalogVideo |  |  | This element represents a device used to capture Analog Videos.
| 1 | Y | CCD |  |  | This element represents a Charged Coupled Device Camera.
| 1 | Y | CCD | Type | CCDType | This field records the Type of CCD Camera used in this case. | enum | Conventional, EMCCD, Other
| 2 |  | CCD | CCDArchitecture | CCDArchitecture | This field captures the Architecture of this CCD Camera and its associated Transfer Mode. | enum | Frame transfer, Full-frame, Interline transfer
| 3 |  | CCD | RegisterWellCapacity |  | This field records the maximum amount of charge each read-out register well can hold, measured in electron counts. | integer with unit:RegisterWellCapacity | [6 empty] | MOVED to CCD | only CCD has this
| 1 | Y | CMOS |  |  | This field records a complementary metal–oxide–semiconductor camera | [13 empty] | This field records whether (True) or not (False) this Detector is a Multianode composed of an array of multiple Photomultiplier Tubes.
| 1 | Y | CMOS | Type | CMOSType | This field records the Type of CMOS Camera used. | enum | Conventional, Scientific, Other
| 1 | Y | Intensified |  |  | This field records an intensified Camera. | [7 empty] | Changed Name to just "Intensified" since they can be CCD or CMOS
| 2 |  | Intensified | SensorType | SensorType | This field records the Type of Sensor used in this case. | enum | CCD, CMOS, Other | R |  |  |  |  | ADDED | and added this line to distinguish
| 2 |  | Intensified | IntensifierType | IntensifiedCCDType | This field records the Type of Intensifier used for this Camera. | enum | Conventional, EBCCD, MicroChannelPlate, Other | R |  |  |  |  | CHANGED
| 3 |  | Intensified | RegisterWellCapacity |  | This field records the maximum amount of charge each read-out register well can hold, measured in electron counts. | integer with unit:RegisterWellCapacity | [6 empty] | MOVED to CCD | only CCD has this
| 1 | Y | PointDetector-APE |  |  | This element represents any Detector that collects light from a restricted point-like area of an image. These detectors are typically used in Laser Scanning Confocal Microscopy. |  | Photomultiplier, PhotoDiode, Hybrid Photodetector | [5 empty] | Changed description to incorporate ISO 10934-1-2002 standard definitions. |  |  |  | Changed description to incorporate ISO 10934-1-2002 standard definitions.
| 1 | Y | PointDetector-APE | Coating | CoatingMaterial | This field records the Coating Material used to increase the photoelectric properties of the metallic cathode used in this Detector. The coating releases electrons much more readily than the underlying metal, allowing it to detect the low-energy photons in infrared radiation. The lens transmits the radiation from the object being viewed to a layer of coated glass. The photons strike the metal surface and transfer electrons to its rear side. The freed electrons are then collected to produce the final image. | enum | Bi-alkali, Multi-alkali, Gallium arsenide - GaAs, Gallium arsenide phosphide - GaAsP, Other | R
| 1 | Y | PointDetector-APE | SignalProcessing | SignalProcessingMode | This field describes the method employed by this device to process the output signal and convert it into a measure of incident light intensity. In Photon Counting mode the system is set in a manner in which each photon hitting the head of the detector gives either no or 1 diskrete pulse. This way by counting the pulses one is sure to be counting individual photons. In Analog mode the system is set so the mean value of all signal pulses is used as a measure of intensity. | enum | Analog, Photon counting, Other
| 3 |  | PointDetector-APE | ResponseTime |  | This field records the Time required for the Detector to respond to an optical input. | float with unit:ResponseTimeUnit
| 3 |  | PointDetector-APE | DeadTime |  | This field records the Time required for the Detector to become responsive again after being hit by a photon. | float with unit:DeadTime
| 3 |  | PointDetector-APE | RadiantSensitivity |  | This field records the Radiant Sensitivity observed with this Detector at the peak Wavelength | float with unit:RadiantSensitivityUnit |  | R
| 1 | Y | PhotoMultiplier |  |  | A Photomultiplier Tube (PMT) is a photoemissive device in which the absorption of a photon results in the emission of an electron and it is used for the detection of very weak signals. These Detectors work by amplifying the electrons generated by a photocathode exposed to a photon flux. The higher the supplied voltage (PMTVoltage setting), the greater the amplification of the current (PMTGain setting) and thus sensitivity, but also the higher the noise in the acquired image. The amplifier offset (PMTOffset setting) defines black level setting of the image. A PMT has no inherent spatial information of the image.
| 2 |  | PhotoMultiplier | CollectionEfficiency |  | This field measures the probability that photoelectrons will land on the effective area of the first dynode of the photomultiplier stage. While there might be dispersion at the second and later dynodes, this tends to have less of an effect on the total collection efficiency as the number of emitted electrons increases greatly. | percentageFraction |  | R
| 2 |  | PhotoMultiplier | Multianode |  | This field records whether (True) or not (False) this is Photo Multiplier Tube is composed of an array of multiple | boolean |  | R
| 3 |  | PhotoMultiplier | MultianodeChannelNr |  | In case this Detector is a Multianode composed of an array of multiple Photomultiplier Tubes, this field records the number of Channels it consists of. | integer with unit:none | [6 empty] | ADDED | Yes good. How about also specifying the arrangement: linear, square, hexagon, round, other? |  |  | Added
| 3 |  | PhotoMultiplier | MultianodeArrangement | MultanodeArrangement | In case this Detector is a Multianode composed of an array of multiple Photomultiplier Tubes, this field records the arrangement of the individual Photomultiplier Tubes. | enum | Linear, Round, Square, Other | [5 empty] | ADDED
| 3 |  | PhotoMultiplier | HeadOn |  | This is True if the Photocathode is positioned at either end of the vacuum tube. This is False if the photocathode element is located on the side of the vacuum tube (side-on geometry). | boolean
| 1 | Y | PhotoDiode |  |  | This element represents a Photovoltaic Detector. Most photodiodes used in confocal microscopes are Avalanche Photodiode (APD) which because of the amplification gain have good sensitivities and less dark current.
| 2 |  | PhotoDiode | Avalanche |  | This field records whether (True) or not (False) this is an "Avalanche" Photodiode. | boolean
| 3 |  | PhotoDiode | PINJunction |  | This field records whether this photodiode is based on a Positive-Intermediate-Negative (PIN) or a Positive-Negative (PN) architecture. | boolean
| 1 | Y | HybridPhotoDetector |  |  | This element represents a hybrid Photomultiplier and Avalanche Photodiode (APD).
| 1 |  | GenericDetector |  |  | This element is used to represent a Detector as a collection of Key/Value pairs, which are stored in a Map. One of the other detector elements should always be used in preference to this if possible.
| 2 |  | GenericDetector | Map |  | This is a Mapping of Key/Value pairs used to store customized metadata that describes this element. | element
Note on A1: ====== ID#AAAAKgE2KN0 (2019-12-18 22:28:54) This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on D1: ====== ID#AAAAKgE2KNw Caterina Strambio De Castillia (2019-12-19 19:05:02) This essentially means that this Element is not actually used per se. It is defined as the Ancestor of Elements that are actually created when representing Hardware components. The advantage is this way all Children elements will inherit all Attributes of this Abstract ancestor and one avoids to have to repeat the same Attributes multiple times. It is a modeling trick that avoids data duplication.
Note on B3: ====== ID#AAAAEEcLF7o (2019-12-18 22:28:54) This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.

## 4DN-BINA OME-Basic Ext_Stn_v02
4DN-OME Model |  |  | APE = Abstract Parent Element
[14 empty] | BINA WORKING GROUP | [6 empty] | Methods checklist |  |  | QUAREP comments
| Tier | M&M (Montero, Brown, Cameron, Itano and Strambio-De-Castillia) | Element | Attribute/sub-Element | Complex type | Description | Data type | Allowed values | Cardinality | OME 6-2016 | OME Cardinality/Required? | OME Data type | DECISION | Damir Comments | Arturo comments | Pablo Comments | Caterina comments | Jim comments | Nathalie comments | Ulrike comments | Michelle Itano | Lisa Cameron | Paola Motero Llopin | WG7 Member 1 [Enter your name and Choose a color] | WG7 Member 2 [Enter your name and Choose a color] | WG7 Member n [Enter your name and Choose a color]
IMAGE ACQUISITION Settings | 1 | Y | Experiment |  |  | This element describes the experiment during which this Image was acquired. The optional Description element may contain free text to further describe the experiment. |  |  | 1 | same | [9 empty] | We should indicate the essential info we would like to hear in this section, otherwise it is just a notes section
| 1 |  | Experiment | ID |  | A Unique Identifier for this Experiment. |  |  | R | same
| 1 |  | Experiment | Purpose | ExperimentPurpose | This field records a term used the general purpose of the imaging Experiment performed in this case. | enum | Co-localization, Electrophysiology, HightThroughputScreen, Histology, Intensity Quantification, Localization, Morphometry, OpticalTrapping, PGI-Documentation, PhotoAblation, TimeLapse, Other | R | was Type | R | enum
| 1 |  | Experiment | ExperimenterRef |  | This empty element refers to Experimenter that performed this experiment. | Extension of Reference |  | 1 … ∞ | same | 0 ... 1 | Ref | [7 empty] | Dyes: Fluorescent Proteins, Synthetic Dyes, None, Others - LabelingMethods: Immunofluorescene/labelling (Antibody Labelling), SNAP, HALO, SingleMoleculeFISH, CASFISH, FluorescentProtein (Electroporation), FluorescentProtein (e.g. with Lipofectamin), FluorescentProtein (endogenous), Click Chemistry, None, Other ... I don't know these methods: ChemicalModification, ForcedIntercalation, Immunocytochemistry, MolecularBeacon, SelfLabellingEnzymeTag, AptamerMotif, BacteriophageDerivedRNATag, BiarsenicalDye,
| 1 | Y | Experiment/Description | Description |  | A free text multi-line optional description for this component. | Element |  | 0 … 1 | same | 0 ...1 | Element
|  | Y | [7 empty] | MicrobeamManipulation [Sub-element]
| 1 |  | Experiment | Name |  | A user assigned Name for this Experiment | Denomination |  | R
| 1 | Y | Experiment | LabellingMethod |  | This field records a term that describes the method used for the Labelling or Visualization of the structures (i.e. molecular components) under study in this Experiment. The use of a sub-class of the 'Visualization Method' (http://purl.obolibrary.org/obo/FBbi_00000031) term from the Biological Imaging Methods Ontology (FBbi; https://www.ebi.ac.uk/ols/ontologies/fbbi) is recommended. | string |  | R |  |  |  | ADDED | They really call this "VisualizationMethod"? That term sounds to me how it's being visualized on a computer screen. |  |  | PROPOSE TO ADD
| 1 |  | Experiment | LabellingMethodTermAccession |  | This field records the Accession ID for the term used to describe the Visualization Method (also often referred to as Labelling Method). The use of a sub-classes of 'Visualization Method' (http://purl.obolibrary.org/obo/FBbi_00000031) from the Biological Imaging Methods Ontology (FBbi; https://www.ebi.ac.uk/ols/ontologies/fbbi) is recommended. | anyURI | [5 empty] | ADDED |  |  |  | PROPOSE TO ADD
| 1 |  | Experiment | Tag |  | This field describes small molecules, peptides or full proteins that were used as label in this Experiment. The use of a sub-class of the 'Tag' (http://purl.obolibrary.org/obo/MI_0507) term from the Molecular Interactions Controlled Vocabulary (MI; http://purl.obolibrary.org/obo/mi.owl) is recommended. | string | [5 empty] | ADDED |  |  |  | PROPOSE to ADD
| 1 |  | Experiment | TagTermAccession |  | This field records the Accession ID for the term used to describe the Tag used in this Experiment. The use of a sub-class of http://purl.obolibrary.org/obo/MI_0507 is recommended. | anyURI | [5 empty] | ADDED |  |  |  | PROPOSE to ADD
| 1 |  | Experiment | HistochemistryType | HystochemistryType | This field records a term that describes the general category of Histochemistry staining used in this case. | enum | Acid fuschin, Alcian blue, Aniline Blue, Azan, Chrome alum/haemotoxylin, Eosin, Giemsa, Histochemistry stains, Isamin blue/eosin, Masson's trichrome, Methyl green, Methylene blue, Nissl and methylene blue, Orange G, Periodic acid-Schiff reaction (PAS), Pyronin G, Reticulin Stain, Silver and gold, Sudan Black and osmium, Toluidine blue, van Gieson, Other |  |  |  |  | ADDED |  |  |  | CONSIDER substituting with http://purl.obolibrary.org/obo/FBbi_00000567
| 1 |  | Experiment | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... 1
| 1 |  | Experiment | Sample |  | This field captures a description of the sample that was imaged in this experiment. | Element |  | 1 ... ∞
| 1 | Y | Sample |  |  | A description of the sample that was imaged in this experiment. |  |  | 1
| 1 |  | Sample | ID |  | A Unique Identifier for this Sample. |  |  | R
| 1 |  | Sample | Name |  | A user assigned Name for this Sample. | Denomination |  | R
| 1 |  | Sample | BiologicalOrigin |  | This field captures whether (True) or not (False) this Sample is of Biological Origin. | boolean |  | R |  |  |  | ADDED
| 1 |  | Sample | LiveSample |  | This field captures whether (True) or not (False) this Sample was living during Image Acquisition. | boolean |  | R |  |  |  | ADDED
| 2 |  | Sample | MetrologySample |  | This field captures whether (True) or not (False) this is a Metrology Sample (i.e., a Sample that is used specifcially to calibrate and assess the performance of the Instrument utilized for this experimental Image Acquisition. | boolean |  | R |  |  |  | ADDED
| 1 |  | Sample | Organism |  | If this sample is of biological origin, this field records the biological organisn this sample derives from. A biological is a biological entity that consists of one or more cells and is capable of genomic replication (independently or not). The use terms from NCBITaxon (http://purl.obolibrary.org/obo/NCBITaxon_1) is recommended. Note: if the sample is not of biological origin please enter "none". | string |  | R |  |  |  | ADDED |  |  |  | SUGGESTED
| 1 |  | Sample | OrganismTermAccession |  | This field records the Accession ID for the term used to describe the Organism. The use of the NCBITaxon ontology (http://purl.obolibrary.org/obo/NCBITaxon_1) is recommended. | anyURI | [5 empty] | ADDED |  |  |  | SUGGESTED
| 1 |  | Sample | SampleType | SampleType | This field records a term that describes the general category of this Sample. | enum | WholeOrganism, Organism part, Organ, Tissue, Cell, Sub-Cellular structure, Organoid, Other |  |  |  |  | ADDED
| 1 |  | Sample | OrganismPart |  | If applicable, this field records the part of organism's anatomy or substance arising from an organism from which the biomaterial was derived, excludes cells. E.g. tissue, organ, system, sperm, blood or body location (arm). The use of a sub-class of the 'Organism Part' (http://www.ebi.ac.uk/efo/EFO_0000635) term from the Experimental Factor Ontology (EFO; http://www.ebi.ac.uk/efo/) is recommended. | sting | [5 empty] | ADDED |  |  |  | SUGGESTED
| 1 |  | Sample | OrganismPartTermAccession |  | This field records the Accession ID for the term used to describe the Organism. A sub-classe of the EFO:0000635 (http://www.ebi.ac.uk/efo/EFO_0000635) term from the Experimental Factor Ontology (EFO; http://www.ebi.ac.uk/efo/) is recommended. | anyURI | [5 empty] | ADDED |  |  |  | SUGGESTED
| 1 | Y | Sample | SamplePreparationMethod |  | This field records the method that was used to prepare this sample. The use of a sub-class of the 'Sample Preparation Method' (http://purl.obolibrary.org/obo/FBbi_00000001) term from the Biological Imaging Methods Ontology (FBbi; https://www.ebi.ac.uk/ols/ontologies/fbbi) is recommended. | string | [5 empty] | ADDED |  |  |  | SUGGESTED
| 1 |  | Sample | SamplePreparationMethodTermAccession |  | This field records the the Accession ID for the term used to describe method that was used to prepare this sample. A subclass of the Fbbi:00000001 (http://purl.obolibrary.org/obo/FBbi_00000001) term from the Biological Imaging Methods Ontology (FBbi; https://www.ebi.ac.uk/ols/ontologies/fbbi) is recommended. | anyURI
| 2 |  | Sample | FacingTheObjective | FacingTheObjectiveChoice | This field captures whether the CoverGlass, the SampleHolder, or neither (None), faced the Objective during Acquisition of this Image. | enum | SampleHolder, CoverGlass, None |  |  |  |  | ADDED
| 1 |  | Sample | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... 1
| 1 |  | Sample | Description |  | A free text multi-line optional description for this component. | Element |  | 0 ... 1 | [9 empty] | type of sample: Biological, non-biological -Biological: tissue, cells, organoids, others , Live, fixed -non-biological:...
| 1 | Y | Sample | MountingMedium |  | This element refers to the Mounting Medium that was used during the acquisiton of this Image. According to ISO 10934-1:2002, a Mounting Medium is a liquid, synthetic resin or other medium in which this Specimen is placed during Image Acquisition. NOTE For fixed Samples observed by transmitted-light microscopy, this Medium is transparent, colourless and of specified refractive index, enclosed between the slide and the cover glass. For reflected-light microscopy, the mounting medium is normally a resin with which the Sample is impregnated so that a polished section may be made. For live Samples, the Medium is typically the tissue culture growth substrate. | Element |  | 0 ... ∞
| 1 | Y | Sample | CoverGlass |  | This element describes a rectangular or circular piece of thin glass used to cover a microscopical preparation. | Element |  | 0 ... ∞ |  |  |  | ADDED
| 1 | Y | Sample | SampleHolder |  | This element describes the device (often referred to as Sample Carrier, Container or Vessel) on which the specimen was mounted for microscopical examination. It typically is flat and it could be made of glass or plastic. Examples include Microscope slides, Petri dishes, and multi-well plates. | Element |  | 0 ... ∞ |  |  |  | ADDED
| 1 | Y | MountingMedium |  | ManufacturerSpecs | This element refers to the Mounting Medium that was used during the acquisiton of this Image. According to ISO 10934-1:2002, a Mounting Medium is a liquid, synthetic resin or other medium in which this Specimen is placed during Image Acquisition. NOTE For fixed Samples observed by transmitted-light microscopy, this Medium is transparent, colourless and of specified refractive index, enclosed between the slide and the cover glass. For reflected-light microscopy, the mounting medium is normally a resin with which the Sample is impregnated so that a polished section may be made. For live Samples, the Medium is typically the tissue culture growth substrate. |  |  | 0 ... ∞
| 1 |  | MountingMedium | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 |  | MountingMedium | Type | MountingMediumType | This field records the general category to which this Mounting Medium belongs. | enum | Alcohol based, Glycerol based, Growth medium, Impregnating resin, Other | R
| 2 |  | MountingMedium | RefractiveIndex |  | This field reocrds the nominal Refractive Index of this Mounting Medium, as specified by the Manufacturer. | float with unit:none |  | R
| 1 | Y | CoverGlass |  | ManufacturerSpecs | Rectangular or circular piece of thin glass used to cover a microscopical preparation. |  |  | 0 ... ∞ |  |  |  | ADDED |  |  |  | ADDED as separate Element
| 1 |  | CoverGlass | ID |  | A Unique Identifier for this component. |  |  | R |  |  |  | ADDED
| 1 | Y | CoverGlass | Coating |  | This field reports a description for any Coating that was layered on the Cover Glass, either by the Manufacturer or after purchase (e.g. Poly-L-Lysine, Collagen). | string |  | R
| 1 | Y | CoverGlass | CoverGlassNo | CoverGlassGradeList | This fields reports the numerical designation that describes the expected thickness of this Coverslip (e.g., #1.5) and the relative expected thickness range expressed in mm. | enum | #0 (0.085 -0.13), #1 (0.13 -0.16), #1.5 (0.16 -0.19), #1.5H (0.17 -0.18), #2 (0.19 -0.23), #3 (0.25 -0.35), #4 (0.43 -0.64) | R |  |  |  | ADDED | [5 empty] | I like this!
| 2 |  | CoverGlass | Thickness |  | This fields reports the measured thickness of this Coverslip expressed in mm. | float with unit:none | [5 empty] | ADDED
Y
| 1 | Y | SampleHolder |  | ManufacturerSpecs | This term describes the device (often referred to as Sample Carrier, Container or Vessel) on which the specimen was mounted for microscopical examination. It typically is flat and it could be made of glass or plastic. Examples include Microscope slides, Petri dishes, and multi-well plates. |  |  | 0 ... ∞ |  |  |  | ADDED |  |  |  | ADDED as separate Element
| 1 |  | SampleHolder | ID |  | A Unique Identifier for this component. |  |  | R |  |  |  | ADDED
| 1 | Y | SampleHolder | Type |  | This field records the Type of this Sample Holder. Options include, microscope slide, Petri dish, tissue culture plate, micro-well plate, etc. (e.g. 35 mm MatTek Glass Bottom Dish). | string |  | R |  |  |  | ADDED
| 1 |  | Image |  |  | This element describes the actual image that was acquired and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. |  |  |  | same | [6 empty] | ADD BACK STAGE LABEL
| 1 | Y | Image | ID |  | A Unique Identifier for this Image. | ImageID |  | R | same
| 1 | Y | Image | Name |  | A user assigned Name for this Image | Denomination |  | R | same
| 1 | Y | Image | InstrumentRef |  | This empty element refers to the Instrument that was utilized to acquire this image. | Extension of Reference |  | 1 | same | 0 ... 1
| 1 | Y | Image | ExperimentRef |  | This empty element refers to the Experiment during which this Image was acquired. | Extension of Reference |  | 1 | same | 0 ... 1
| 1 | Y | Image | ExperimenterRef |  | This empty element refers to the Experimenter that acquired this image. | Extension of Reference |  | 0 … ∞ | same | 0 ... 1
| 1 | Y | Image | ExperimenterGroupRef |  | This empty element refers to the ExperimenterGroup that owns this Image. | Extension of Reference |  | 0 ... 1 | same | 0 ... 1
| 2 |  | Image | ROIRef |  | This empty element refers to any predefined four dimensional ROI (Region of Interest) that is associated with this Image. | Extension of Reference |  | 0 … ∞ | same | 0 … ∞ | [7 empty] | we record x-y stage coordinate for each ROI - use to map on lower mag overview images.???
| 1 | Y | Image | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same | 0 ... ∞
| 1 | Y | Image | AcquisitionDate |  | The Acquisition date of the Image. | Element |  | 0 ... 1 | same | 0 ... 1
| 1 |  | Image | Description |  | A free text multi-line optional description for this component. | Element |  | 0 ... 1 | same | 0 ... 1
| 1 |  | Image | Pixels |  | Pixels is going to be removed in the future, but it is still required. [...]. | Element |  | 1 | same | 1
| 2 |  | Image | ImagingEnvironment |  | This describes the environment that the biological sample was in, during the experiment. | Element |  | 0 … 1 | same | 0 … 1
ObjectiveSettings moved to Plane
| 1 | Y | Image | AcquisitionSoftwareRef |  | This empty element refers to any Acquisition Softwre that was during the Acquisition of this Image. | Extension of Reference
| 1 | Y | Image | MicroscopeSettings | Extension of Settings | This element refers to the Microscope used to acquire this image and it holds the settings that were applied to the Microscope during this Acquisition. | Element |  | 0 ... 1
| 2 |  | Image | MicroscopeTableSettings | Extension of Settings | This element refers to the MicroscopeTable used to support the Microscope used to acquire this Image and holds the settings that were applied to the Microscope during this Acquisition. | Element |  | 0 …1 |  |  |  | Added MicroscopezTableSettings | How about moving this to the Experiment? It feels weird to have it per Image as if it may change for the next Image taken with this system? |  |  | The reason to have it per Image is exactly because it can change from one Image to the next.
| 1 | Y | Pixels (should merge with Image) |  |  | Pixels is going to be removed in the future, but it is still required. [...]. |  |  | 1 | same | 1 | 1 | KEPT PIXELS in alignment to current version of OME 2016 | Let's remove the Pixels element in this round and merge it into Image as you had it before.
| 1 |  | Pixels (should merge with Image) | ID |  | A Unique Identifier for this Image. | LSID |  | R | same | R | LSID
| 1 | Y | Pixels (should merge with Image) | DimensionOrder | PixelsDimensionOrder | This field records the order in which the individual planes of data are interleaved in the Image. | enum | XYZCT, XYZTC, XYCTZ, XYCZT, XYTCZ, XYTZC | R | same | R | enum | ADDED to align with Material and Methods guidelines
| 1 | Y | Pixels (should merge with Image) | SizeX |  | This field records the X size of the pixel data array. | PositiveInteger |  | R | SizeX | R | PositiveInteger
| 1 | Y | Pixels (should merge with Image) | SizeY |  | This field records the Y size of the pixel data array. | PositiveInteger |  | R | SizeY | R | PositiveInteger
| 1 | Y | Pixels (should merge with Image) | SizeZ |  | This field records the number of Z-planes acquired for this Image. | PositiveInteger |  | R | SizeZ | R | PositiveInteger
| 1 | Y | Pixels (should merge with Image) | SizeC |  | This field records the number of Channels used in this Image | PositiveInteger |  | R | SizeC | R | PositiveInteger
| 1 | Y | Pixels (should merge with Image) | SizeT |  | This field records the number of Timepoints used in this Image. | PositiveInteger |  | R | SizeT | R | PositiveInteger
| 1 | Y | Pixels (should merge with Image) | PhysicalSizeX |  | This field records the real world Physical Width represented by a Pixel in this Image. This measure is often referred to as X Pixel Size. | PositiveFloat with unit:PhysicalSizeXUnit; default µm |  | R | PhysicalSizeX |  | PositiveFloat | CHANGED DESCRIPTION
| 1 | Y | Pixels (should merge with Image) | PhysicaSizeY |  | This field records the real world Physical Height represented by a Pixel in this Image. This measure is often referred to as Y Pixel Size. | PositiveFloat with unit:PhysicalSizeYUnit; default µm |  | R | PhysicalSizeY |  | PositiveFloat | CHANGED DESCRIPTION
| 1 | Y | Pixels (should merge with Image) | PhysicalSizeZ |  | This field records the real world Physical distance between the individual Z-planes recorded for this Image. This measure is sometimes referred to as Z Pixel or Voxel Size. | PositiveFloat with unit:PhysicalSizeZUnit; default µm |  | R | PhysicalSizeZ |  | PositiveFloat | CHANGED DESCRIPTION
| 1 | Y | Pixels (should merge with Image) | TimeIncrement |  | In cases in which a time series is acquired with a global timing specification setting, this field describes the time between each sequential plane (or frame) in the series (i.e., from the start of time point 1 to the start of time point 2). Examples include, a video stream or a multi-dimensional Acquisition routine. Note that, this field is not intended to substitute the Timestamp attribute recorded in Plane. Also note that in case this field is not relevant e.g. a time lapse Acquisition with irregular increments, the value -1 should be used to indicate that only the Timestamps (in the Plane element) are valid. | PositiveFloat with unit:TimeIncrementUnit |  | R | TimeIncrement |  | Float | CHANGED DEFINITION TO MAKE IT CLEAR THAT THIS IS A SETTING and it should not be substituted to the TimeStamp. | How to indicate it's a unevenly spaced time series? Value = 0.0? That would indicate time points are to be derived from Plane's TimeStamps.
| 2 |  | Pixels (should merge with Image) | PixelType (was Type) | PixelType | This field records the variable type used to represent the Pixel Intensity in the Image. | enum | bit, complex, double, double-complex, float, int8, int16, int32, uint8, uint16, uint32, Other | R | was Type | R | enum
| 2 |  | Pixels (should merge with Image) | Interleaved |  | This field records the order in which the channels are arranged within the data block: True if channels are stored RGBRGBRGB...; False if channels are stored RRR...GGG...BBB… | boolean |  |  | Interleaved |  | boolean
| 3 |  | Pixels (should merge with Image) | SignificantBits |  | The field records the number of bits within the type storing each pixel that are significant. e.g. you can store 12 bit data within a 16 bit type. This does not reduce the storage requirements but can be a useful indicator when processing or viewing the image data. | PositiveInteger |  |  | SignifanctBits |  | PositiveInt
| 3 |  | Pixels (should merge with Image) | BigEndian |  | This field is True if the pixels data was written in BigEndian order. If this value is present it should match the value used in BinData or TiffData. If it does not a reader should honour the value used in the BinData or TiffData. This values is useful for MetadataOnly files and is to allow for future storage solutions. | boolean |  |  | BigEndian |  | boolean
| 1 |  | Pixels (should merge with Image) | Plane |  | Each Image has to have at least one Plane and one Channel. The Plane element holds Microscope stage and image timing data for a given channel/z-section/timepoint. The total number of Plane elements for a given Image must equal to SizeZ x SizeC x SizeT. | Element |  | 1 … ∞ | same
| 1 |  | Pixels (should merge with Image) | Channel |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. | Element |  | 1 … ∞ | same
|  |  | Pixels (should merge with Image) | CHOOSE ONE OF THE FOLLOWING | [5 empty] | same
| 1 |  | Pixels (should merge with Image) | TiffData |  | This field describes the location of the pixel data in a tiff file. | Element | choose one of these x elements | 1 … ∞ | same
| 1 |  | Pixels (should merge with Image) | BinData |  | The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream. | Element |  | or | same
| 1 |  | Pixels (should merge with Image) | MetadataOnly |  | This place holder means there is no pixel data in this file. | Element |  | or | same |  |  | Fixed description type |  |  |  | Description typo
| 1 | Y | Pixels (should merge with Image) | SequentialAcquisition |  | In the case in which this Image has multiple Channels, this field records whether or not the Acquisition of those Channels happened in a Sequential (True) or Simultanous (False) manner. | boolean |  | R
| 1 |  | Pixels (should merge with Image) | NGFFData |  | This element describes the location of the pixel data in a NGFF storage set. | Element |  | or |  |  |  | CHANGED | Support for other containers |  |  | CHANGED TO GENERIC NGFF (Next Generation File Format to follow current nomencalture used by OME)
| 2 |  | Pixels (should merge with Image) | AdditionalDimensionMap |  | This element defines any additional dimension beyond the canonical 5 dimensions XYZCT. | Element |  | 0 … ∞ |  |  |  | ADDED | To avoid this kludge: https://docs.openmicroscopy.org/ome-Model/6.0.0/developers/6d-7d-and-8d-storage.html |  |  | Changed description
| 2 |  | AdditionalDimensionMap |  |  | This map defines any additional dimension beyond the canonical 5 dimensions XYZCT. It contains a number of elements for each Plane of data referring to this Additional Dimension. |  |  | 0 ... ∞ |  |  |  |  | One possible approach for an arbitrary number of additional dimensions |  |  | SHOULD THIS BE IN ADVANCED? | Not sure where this belongs - on-the-fly-processing. Maybe a can of worms we don't want to open?
| 2 |  | AdditionalDimensionMap | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 |  | AdditionalDimensionMap | Name |  | A user assigned Name for this component. | string |  | R
| 2 |  | AdditionalDimensionMap | Size |  | Number of sample points along this dimension | integer with unit:none |  | R |  |  |  | KEPT | Why not keep it conform the 5 canonical dimensions and use Size? |  |  | Substituted with Start and End
| 2 |  | AdditionalDimensionMap | Type | AdditionalDimensionType | This describes the Type of this Additional Dimension. | enum | Angle, Phase, Tile, Lifetime, Lambda, M/Z, MultiView, Other | R |  |  |  |  | I wasn't too sure what Tile means either but it's somewhat explained in: https://docs.openmicroscopy.org/ome-Model/6.0.0/developers/6d-7d-and-8d-storage.html |  |  | What is Tile?
| 2 | Y | AdditionalDimensionMap | TypeDescription |  | This is a description of the Type used to indicate to an application or user how the data should be interpreted. | string
| 2 |  | AdditionalDimensionMap | Order |  | The order in which the individual Planes of data referring to this dimension are interleaved in the Image past the canonical 5 dimension XYZCT. The number should be zero based, e.g., 0, 1, 2 etc. | integer with unit:none |  | R
| 2 |  | AdditionalDimensionMap | PresetStepIncrement |  | This field records the preset size of Step Increment between Planes of data along this dimension. The Increment is numbered from 0, with default value of 1. Setting a value for this attribute to -1 indicates non-uniform sampling along this Additional Dimension. In this case the actual Step Increment value associated with each Plane along this Additional Dimension will have to be stored in TheIncrementStamp attribute of TheAdditionalDimension. | float with unit:none | [5 empty] | KEPT; changed Name | changed description and data type |  |  | CHANGED NAME, DESCRIPTION and UNIT
| 2 |  | AdditionalDimensionMap | Unit |  | The unit used for the metric recorded along this dimension | string |  | R
| 2 |  | AdditionalDimensionMap | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞
| 2 |  | AdditionalDimensionMap | TheAdditionalDimension |  | This element refers to each individual Plane of data pertaining to a given Additional Dimension. | Element |  | 0 ... 1
| 2 | Y | AdditionalDimensionMap | Description |  | A free text multi-line optional description for this component. | Element |  | 0 ... 1
| 2 |  | TheAdditionalDimension |  |  | This element refers to each individual Plane of data pertaining to a given Additional Dimension. | [10 empty] | Added; Changed Plane to Plane (or Chunk) in description
| 2 |  | TheAdditionalDimension | ID |  | A Unique Identifier for this component. | LSID |  | R | [7 empty] | Changed Name; Changed Plane to Plane (or Chunk) in description
| 2 | Y | TheAdditionalDimension | TheAdditionalDimensionIndex |  | This records the position of this specific Plane of Additional Dimension data. | integer with unit:none |  | R | [7 empty] | Changed Plane to Plane (or Chunk) in description
| 2 |  | TheAdditionalDimension | TheIncrementStamp |  | This field records the actual Step Increment value associated with each Plane of data along this Additional Dimension for this Image. While, this field is only required with non-uniform sampling along this Additional Dimension (as indicated by a value of -1 stored in AdditionalDImension/@ActualStepIncrement), it is always allowed. | float with unit:AdditionalDimension->Units | [5 empty] | ADDED | added - analogous to Timestamps for the Time dimension. |  |  | CHANGED NAME, DESCRIPTION and UNIT
| 2 |  | Settings (extension base) |  | Reference | Settings is an empty complex type that is contained and extended by all the *Settings elements. Each *Settings element defines an attribute Named ID of simple type *ID and the other information that is needed. Each simple type *ID is restricted to the base type LSID with an appropriate pattern |  |  |  | Same |  |  | ADDED Additional vendor information
| 3 |  | Settings (extension base) | MapAnnotation |  | A Map Annotation. The contents of this is a list of Key/Value pairs. | Element |  | 0 … ∞ |  |  |  | ADDED: used existing MapAnnotation element from OME-XML
| 3 | Y | Settings (extension base) | AdditionalVendorInformation |  | Any additional information the Manufacturer of this component might want to provide to describe the component which is not covered as part of this standard including possibly proprietary attributes. The vendor is fully responsible for this information including but not limited to making sure this information is properly recorded and kept up to date. | Element |  | 0 … ∞ |  |  |  | ADDED. This AdditionalVendorInformation element is going to be a subelement of ManufacturerSpec which is the Extension base for all Microscope components described in this Model. | Will that also extend to the Settings entries? Vendors will also want to have a place for this type of stuff there. |  |  |  |  | I second Damir
| 3 | Y | MapAnnotation |  | Annotation | A Map Annotation. The contents of this is a list of Key/Value pairs. | [6 empty] | ADDED: used existing MapAnnotation element from OME-XML
| 3 | Y | MapAnnotation | ID | AnnotationID | A Unique Identifier for this component. | LSID |  | R
| 3 | Y | MapAnnotation | Namespace |  | We recommend the inclusion of a Namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type. | anyURI |  | R
| 3 |  | MapAnnotation | Annotator (ExperimenterID) | ExperimenterID | The Annotator is the person who attached this annotation.e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator. | LSID
| 3 |  | MapAnnotation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 3 |  | MapAnnotation | Description |  | A free text multi-line optional description for this component. | Element |  | 0 … 1
| 3 |  | MapAnnotation | Value |  | This is a Mapping of Key/Value pairs. | Element |  | 1
| 3 |  | Value |  |  | This is a Mapping of Key/Value pairs.
| 3 |  | Value | M |  | This is a Key/Value pair used to build up a Mapping. The Element and Attribute Name are kept to single letters to minimize the length at the expense of readability as they are likely to occur many times. | Element |  | 0 … ∞
| 3 |  | M |  |  | This is a Key/Value pair used to build up a Mapping. The Element and Attribute Name are kept to single letters to minimize the length at the expense of readability as they are likely to occur many times.
| 3 | Y | M | K |  | This is the key of this Key/Value pair. | string |  | R
| 3 |  | AdditionalVendorInformation |  | AdditionalVendorInformation | Any additional information the Manufacturer of this component might want to provide to describe the component which is not covered as part of this standard including possibly proprietary attributes. The vendor is fully responsible for this information including but not limited to making sure this information is properly recorded and kept up to date. | [6 empty] | ADDED. This AdditionalVendorInformation element is going to be a subelement of ManufacturerSpec which is the Extension base for all Microscope components described in this Model. | [5 empty] | But core components have to be specified; i.e. a vendor standard has to be defined
| 3 |  | AdditionalVendorInformation | ID | AdditionalVendorInformation | A Unique Identifier for this component. | LSID |  | R
| 3 |  | AdditionalVendorInformation | VendorName | AdditionalVendorInformation | The Name of the vendor responsible for this Information | string
| 3 |  | AdditionalVendorInformation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞ |  |  |  | ADDED | not sure how this can be codified in ER diagram |  |  | I think this should be a STRING ANNOTATION and then there should be a reference to a DATA block the same way Pixel refers to TIFF-DATA in the Model.
| 3 |  | AdditionalVendorInformation | VendorBinData | AdditionalVendorInformation | The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream. | Element | [9 empty] | Propose to ADD
| 3 | Y | VendorBinData |  | base64Binary | The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream.
| 3 | Y | VendorBinData | Compression |  | Specifies the compression scheme used to encode the data. | string
| 3 |  | VendorBinData | BigEndian |  | This is true if the binary data was written in BigEndian order. This is dependent on the system architecture of the machine that wrote the pixels. True for essentially all modern CPUs other than Intel and Alpha. All Binary data must be written in the same endian order. | boolean |  | R |  |  |  |  | not necessary for base64 encoded data
| 3 | Y | VendorBinData | Length |  | Character count attribute for the BinData field. This is the length of the base-64 encoded block. It allows easy skipping of the block when parsing the file. [unit:bytes] | NonNegativeLong with unit:bytes |  | R
| 1 | Y | ImagingEnvironment |  | Settings | This element describes the environmental conditions in which the biological sample was kept during the Acquisition of this Image. |  |  | 0 … 1 | same | 0 ... 1 |  | ADD ENVIRONMENT CONTROL DEVICE TO HARDWARE with REF here | [5 empty] | type of environment chamber (i.e. type of enclosure: sample, stage, entire frame, ect...) | System to establish the environmental control (manufacter, Serial Number, Model etc)
| 1 | Y | ImagingEnvironment | Temperature |  | The field records the Temperature the Acquisition was performed at. | float with unit:TemperatureUnit; [default: Celsius ºC] |  | R | same |  | float with unit:TemperatureUnit; [default: Celsius ºC]
| 1 | Y | ImagingEnvironment | Humidity |  | The field records then Humidity the acqusition was performed at. | percentFraction |  | R | same |  | percentFraction
| 1 | Y | ImagingEnvironment | CO2Percent |  | This field records the Carbon Dioxide concentration around the sample [units:none] during Acquisition. Expressed as a percentage fraction as a value from 0.0 to 1.0. | percentFraction |  | R | same |  | percentFraction | CHANGED TO OPTIONAL
| 2 |  | ImagingEnvironment | AirPressure |  | The field records the AirPressure the Acquisition was performed at. | float with unit:AirPressureUnit |  |  | same |  | float with unit:AirPressureUnit
| 3 |  | ImagingEnvironment | Map | Map |  | Element |  | 0 ... 1 | same | 0 ... 1 | Map
| 1 |  | ImagingEnvironment | ID (EnvironmentalControlDeviceID) |  | This the Unique Identifier of the Environmental Control Device utilized during the acquisition of this Image. | EnvironmentalControlDeviceID |  | R
| 2 |  | ImagingEnvironment | O2Percent |  | This field records the Oxygen concentration around the sample [units:none] during Acquisition. Expressed as a percentage fraction as a value from 0.0 to 1.0. | percentFraction | [5 empty] | ADDED
| 1 |  | ImagingEnvironment | Description |  | A free text multi-line optional description for this component. | Element |  | 0 … 1 |  |  |  | ADDED
| 2 | Y | MicroscopeSettings |  | Settings | This element refers to the Microscope used to acquire this Image and it holds the Settings that were applied to the Microscope during this Acquisition. |  |  | 0 … ∞ | [10 empty] | It is strongly linked with the image settings above
| 2 |  | MicroscopeSettings | ID (MicroscopeID) |  | This is the Unique Identifier of the Microscope used in this case. | LSID |  | R
| 3 | Y | MicroscopeSettings | FieldOfView |  | The Field of View defines the area of the specimen that is viewable via the eyepiece. This field caputers the measures size of the Field of View. | float with unit:FieldOfViewUnit[default mm]
| 2 |  | MicroscopeSettings | TotalEffectiveMagnification |  | This fields records the Total Effective Magnification used in this case. This number can be obtained by multiplying Magnification Changer:Magnification times the Objective:NominalMagnification or even better the Objective:CalibratedMagnification if available. | float with unit:none
| 3 | Y | MicroscopeSettings | MicroscopeTubeRef |  | This empty element refers to the existing Microscope Tube component used in this case. | Extension of Reference |  | 0 ... 1 |  |  |  | ADD
| 2 | Y | MicroscopeSettings | MagnificationChangerRef |  | This empty element refers to the existing Magnification Changer component used in this case. | Extension of Reference |  | 0 ... 1 |  |  |  | CHANGED | Now MagnificationChanger and isn't this now in LightPath?
| 2 |  | MicroscopeSettings | ObjectiveRef |  | This empty element refers to the existing Objective component used in this case. | Extension of Reference |  | 1
| 2 |  | MicroscopeTableSettings |  | Settings | This element refers to the MicroscopeTable used to support the Microscope used to acquire this Image and holds the settings that were applied to the Microscope during this Acquisition. |  |  | 1 … ∞ |  |  |  | ADDED MicroscopeTableSettings
| 2 |  | MicroscopeTableSettings | ID (MicroscopeTableID) |  | This is the Unique Identifier of the Microscope used in this case. | LSID |  | R
| 2 |  | MicroscopeTableSettings | Floating |  | This field recods whether the MicroscopeTable was Floating (True) or not (False) during the acquistion of this Image. | boolean |  | R
| 1 | Y | Plane |  |  | Each Image has to have at least one Plane and one Channel. The Plane element holds Microscope stage and image timing data for a given channel/z-section/timepoint. The total number of Plane elements for a given Image must equal to SizeZ x SizeC x SizeT. |  |  | 1 … ∞ | same |  |  | CHANGE DESCRIPTION to include the term Frame and clairfy what this is;
| 1 |  | Plane | TheZ |  | This field records the Z-section to which this plane corresponds. This is numbered from 0. | NonNegativeInt with unit:none |  | R | same |  | R | CHANGED DESCRIPTION
| 1 |  | Plane | TheT |  | This field records th Timepoint to which this plane corresponds. This is numbered from 0. | NonNegativeInt with unit:none |  | R | same |  | R | CHANGED DESCRIPTION
| 1 |  | Plane | TheC |  | This field records th Channel to which this plane corresponds. This is numbered from 0. | NonNegativeInt with unit:none |  | R | same |  | R | CHANGED DESCRIPTION
| 1 | Y | Plane | Timestamp (was DeltaT) |  | This field records the time at which this Plane was acquired as recorded from the beginning of Image Acquisition. | float with unit:ExposureTimeUnit. |  | R | was DeltaT |  |  | CHANGED DESCRIPTION
| 1 | Y | Plane | ExposureTime |  | This field records the duration of the Exposure utilized to acquire this Plane. | float with unit:ExposureTimeUnit. |  | R | same |  |  | CHANGED DESCRIPTION
| 2 |  | Plane | PositionX |  | This field records the X position of the Plane within the Microscope's reference frame. This value can typically be obtained by recording the Sample Positioning Device and Stage Insert coordinates. | float with unit:PositionXUint | [5 empty] | CHANGE DEACRIPTION TO REFLECT DAMIR DRAWING AND SUGGESTIONS | This is typically the axial point in the lightpath which typically is the center of the camera/scan area given by the stage coordinates. |  |  |  |  | Stage position?
| 2 |  | Plane | PositionY |  | This field records the Y position of the Plane within the Microscope's reference frame. This value can typically be obtained by recording the Sample Positioning Device and Stage Insert coordinates. | float with unit:PositionYUint | [5 empty] | CHANGE DEACRIPTION TO REFLECT DAMIR DRAWING AND SUGGESTIONS | [5 empty] | Stage position?
| 2 |  | Plane | PositionZ |  | This field records the Z position of the Plane within the Microscope's reference frame. This value can typically be obtained by recording the Sample Positioning Device and Stage Insert coordinates. | float with unit:PositionZUint | [5 empty] | CHANGE DEACRIPTION TO REFLECT DAMIR DRAWING AND SUGGESTIONS | Will need to review how this interacts with the FocalPosition attributes.
| 1 |  | Plane | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same |  | 0 ... ∞
| 3 |  | Plane | HashSHA1 |  | This optional element is a hash of the plane's image data. It is a choice between all the supported hash types. Currently the only method supported is SHA1 | Element |  | 1 | same |  | 1
| 1 |  | Plane | ID |  | A Unique Identifier for this Plane. |  |  | R
| 1 | Y | Plane | PixelDwellTime |  | This value typically refers to the settings applied to the Confocal Scanner during Acquisition, and it assumes that Pixel Dwell Time was set uniformly throughout the Plane. The Pixel Dwell Time represents the time this Confocal Scanner was set to dwell upon each pixel during the acquisition of this Plane (e.g. 1.58 µsec/pixel). The higher the Pixel Dwell Time the slower is the scan speed, the higher the sensitivity but also the potential photobleaching. | float with unit:PixelDwellTimeUnit |  | R |  |  |  | ADDED | This field records dwell time for a single-pixel Acquisition system such as a laser scanner assuming dwell time doesn't change within a plane.
| 3 |  | Plane | PlaneOffsetX |  | This field records the X offset (expressed in Pixels) that specifies the position of this Plane's center with respect to the center of the full frame. | float with unit:none | [5 empty] | CHANGE DEACRIPTION TO REFLECT DAMIR DRAWING AND SUGGESTIONS | This handles when the image/plane is actually a ROI taken from a full camera frame where the ROI's center is an arbitrary offset (here encoded) from the camera's frame's center. |  |  | ROI is OME is specific entity, Is this the same? |  | Stage position?
| 3 |  | Plane | PlaneOffsetY |  | This field records the Y offset (expressed in Pixels) that specifies the position of this Plane's center with respect to the center of the full frame. | float with unit:none | [5 empty] | CHANGE DEACRIPTION TO REFLECT DAMIR DRAWING AND SUGGESTIONS | Alternatively, we could cover that in the PlaneTransformMatrix below but that seems non-intuitive. |  |  | ROI is OME is specific entity, Is this the same? |  | Stage position?
| 3 |  | Plane | PlaneOffsetZ |  | This field describes the Z-offset imposed by the Focal Positioning device. | float with unit:none
| 3 |  | Plane | TheAdditionalDimensionRef | Reference | This empty element refers to the Additional Dimension Element to which this Plane (or Chunk of image data) corresponds. | Extension of Reference |  | 0 ... ∞ |  |  |  | ADDED |  |  |  | Added (substitutes TheExtra1...n)
| 1 |  | Plane | StageInsertRef |  | This empty element refers to the optional Stage Insert device that was used for the Acquisition of this Image. | Extension of Reference |  | 0 ... ∞ |  |  |  | ADDED
| 1 | Y | Plane | ObjectiveSettings |  | This element refers to the objective used to acquire this image and it holds the settings that were applied to the objective during this Acquisition. | Element |  | 0 ... ∞ 1 | moved from Image |  |  | MOVED IT TO PLANE | Shouldn't ObjectiveSettings be per Channel or possibly per Plane?
| 1 |  | Plane | SamplePositioningSettings |  | This empty element refers to the Sample Holder that was used for the Acquisition of this image. | Element |  | 0 … ∞ |  |  |  | ADDED
| 3 |  | Plane | PlaneTransformMatrix | Map | This element describes a 4x4 matrix relating the position of this Plane typically with respect to other Planes within this Image or collection of Images (or alternatively to the Microscope's reference frame). This can correct for translations, rotations, scale differences, or other affine distortions between Planes from the same set or between Image Datasets. An example is when multiple cameras are used for e.g. to acquire different Channels for the same Image. See: https://www.openmicroscopy.org/community/viewtopic.php?p=17433. | Element |  | 0 ... ∞ 1 |  |  |  | ADD and implement as a MATRIX | To allow converting this plane's position to real-world coordinates accounting for image plane translation, rotation, scale, reflection, and shear. See: https://www.openmicroscopy.org/community/viewtopic.php?p=17433
| 1 |  | SamplePositioningSettings |  | Settings | This element holds the Setting applied to a Sample Positioning Device used during the Acquisiton of this experimental Image as well as a reference (the ID attribute) to the Sample Positioning Device in question. |  |  | 0 … ∞ |  |  |  | ADDED
| 1 |  | SamplePositioningSettings | ID (SamplePositioningID) |  | This is the Unique Identifier of the Sample Positioning Device used in this case. |  |  | R |  |  |  | ADDED
| 1 |  | SamplePositioningSettings | XYVelocitySetting |  | This field describes the XY Velocity that was set for this Sample Positioning Device during the Acquisition of this experimental Image. | float with unit:XYVelocity | [5 empty] | ADDED
| 1 |  | SamplePositioningSettings | ZVelocitySetting |  | This field describes the Z-axis Velocity that was set for this Sample Positioning Device during the Acquisition of this experimental Image. | float with unit:ZVelocity | [5 empty] | ADDED
| 1 |  | SamplePositioningSettings | XYReproducibilitySettings |  | This field describes the XY Precision that was set for this Sample Positioning Device during the Acquisition of this experimental Image. This metric refers to the capacity this Device of maintaining a given XY Position over time as specified by the Manufacturer. This metric is generally referred to as Reproducibility and it might also be referred to as Lateral Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:XYReproducibilitySettingsUnit | [5 empty] | ADDED
| 1 |  | SamplePositioningSettings | ZReproducibilitySettings |  | This field describes the Z Precision that was set for this Sample Positioning Device during the Acquisition of this experimental Image. This metric refers to the capacity this Device of maintaining a given Z Position over time as specified by the Manufacturer. This metric is generally referred to as Reproducibility and it might also be referred to as Focal Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:ZReproducibilitySettingsUnit | [5 empty] | ADDED
| 1 | Y | ObjectiveSettings |  | Settings | This element holds the Setting applied to the Objective used during the Acquisiton of this Image as well as a reference (the ID attribute) to the Objective in question. |  |  | 0 … ∞ | same |  | - | [7 empty] | Already present in the Specs table
| 1 |  | ObjectiveSettings | ID (ObjectiveID) |  | This is the Unique Identifier of the Objective used in this case. | ObjectiveID |  | R | same |  | LSID
| 2 |  | ObjectiveSettings | CorrectionCollarPosition (was CollarPosition) |  | The Correction Collar is normally an adjustable ring on the objective, whose position can be modified. This field records the position of the collar during the Acquisition of this image. Each correction collar has an arbitrary scale or it might be scale-less and instead have fixed positions with qualitative designations (i.e., "Oil", "Water" in case in which the Objective might be corrected for Immersion Liquid). | string |  |  | was CorrectionCollar |  | float with unit:none
Medium; moved to ImmersionLiquid and changed to Type
moved to ImmersionLiquid
| 2 |  | ObjectiveSettings | NumericalApertureSetting |  | In case the Numerical Aperture (NA) of this Objective is variable (i.e., as for example in case of Objectives that can be used with or without Cover Glass or Objectives that can be "dipped" in different media), this field capture the NA setting that was employed for Image Acquisition. NOTE: please note that often NA regulation in these types of Objectives occurs via a rotating collar whose exact position is difficult to measure and reproduce across different Experiments. | string
| 2 |  | ObjectiveSettings | MeasuredRefractiveIndex |  | The field records the Temperature dependent measured Refractive Index of the Immersion Liquid (or Dipping Medium) utilized in this case. | float with unit:none
| 2 |  | ObjectiveSettings | Temperature |  | This field records the Temperature of the Objective if different from the ImagingEnvironment/Temperature. | float with unit:TemperatureUnit; [default: Celsius ºC]
| 2 |  | ObjectiveSettings | IrisAperture |  | This field records the aperture of the Iris diaphagm, which determines determines the size of the objective rear aperture yielding a variable numerical aperture range between 0.5 and the objective's upper limit. | float with unit:none
| 1 | Y | ObjectiveSettings | ImmersionLiquid |  | This element refers to the Immersion Liquid that was used with this Objective during the acquisiton of this Image. According to ISO 10934-1:2002, an Immersion Liquid is a liquid (commonly immersion oil, water or glycerol) specified as suitable for use in the space between the front of an immersion lens and the object NOTE Because the immersion liquid is considered in the computing of corrections to be part of the lens, its refractive index and dispersion (or Abbe number) are critical. Please refer to ISO 8036-1 for more details on Immersion Oil. | Element |  | 0 ... 1
| 1 | Y | ImmersionLiquid |  | ManufacturerSpecs | This element refers to the Immersion Liquid that was used with this Objective during the Acquisition of this Image. According to ISO 10934-1:2002, an Immersion Liquid is a liquid (commonly immersion oil, water, glycerol or any media in which the Specimen is submerged during acquisition in case a Cover Glass is not used) specified as suitable for use in the space between the front of an Objective lens and the Cover Glass or the Object itself (i.e., such as is the case for Dipping Objectives). Refer to ISO 8036-1 for more details on Immersion Oil. Please NOTE: Because the Immersion Liquid is considered in the computing of corrections to be part of the lens, its refractive index and dispersion (or Abbe number) are critical. Also please NOTE: In case a Cover Glass is not used during Acquisition the Immersion Liquid is most often referred to as "Dipping Medium". Examples of Dipping Media include, water-based solutions (biological buffers, culture media, other water-based solutions) and organic-based [FocusClear (CLARITY); CUBIC-1, CUBIC-2, CUBIC-R1, CUBIC-R2; 2,2-thiodiethanol (TDE); ethyl cinnamate; Benzyl Benzoate and BB-PEG (PEGASOS); Benzyl Alcohol, Benzyl Benzoate (BABB); dibenzyl ether (DBE); Ethanol].
| 1 |  | ImmersionLiquid | ID |  | A Unique Identifier for this component. | LSID |  | R
| 1 | Y | ImmersionLiquid | Type (was Medium) | ImmersionLiquidType | This field records the general category to which this Immersion Liquid belongs. | enum | Glycerol, Multi, Mineral Oil, Silicone Oil, Water, Other | R | moved from ObjectiveSettings; was Medium |  | enum
| 1 | Y | ImmersionLiquid | RefractiveIndex |  | This field reocrds the nominal Refractive Index of this Immersion Liquid, as specified by the Manufacturer. | float with unit:none |  | R | moved from ObjectiveSettings |  | float with unit:none
| 2 |  | ImmersionLiquid | Viscosity |  | This field records the nominal Viscosity of this Immersion Liquid as specified by the Manufacturer. | float with unit:ViscosityUnit
| 2 |  | PlaneTransformationMatrix |  |  | This element describes a 4x4 matrix relating the position of this Plane typically with respect to other Planes within this Image or collection of Images (or alternatively to the Microscope's reference frame). This can correct for translations, rotations, scale differences, or other affine distortions between Planes from the same set or between Image Datasets. An example is when multiple cameras are used for e.g. to acquire different Channels for the same Image. See: https://www.openmicroscopy.org/community/viewtopic.php?p=17433. |  |  | 0 ... ∞ 1 |  |  |  |  | I may need a little explanation how this works.
| 2 |  | PlaneTransformationMatrix | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 |  | PlaneTransformationMatrix | Row |  |  | Element |  | 1 … ∞
| 2 |  | PlaneTransformationMatrix | Cell |  |  | Element |  | 1 … ∞
| 2 |  | Row |  | rowType |  |  |  | 1 … ∞
| 2 |  | Row | RowNr |  | The index of this Row. | PositiveInt with unit:none |  | R
| 2 |  | Cell |  | cellType |  |  |  | 1 … ∞
| 2 |  | Cell | CellNr |  | The index of this Cell. | PositiveInt with unit:none |  | R
| 2 |  | Cell | CellValue |  | The value of the metric contained in this Cell. | float with unit:CellValueUnit |  | R
| 1 | Y | Channel |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. |  |  | 1 … ∞ | same
| 1 |  | Channel | ID |  | A Unique Identifier for this Channel. | LSID |  | R | same | R | LSID
| 1 |  | Channel | Name |  | A Name for the Channel that is suitable for presentation to the user. | Denomination |  |  | same |  | string
| 1 |  | Channel | IlluminationType | ChannelIlluminationType | This field records the method of Illumination used to capture the Channel. | enum | Confocal, Epifluorescence, LightSheet, NonLinear, Oblique, TIRF, Transmitted, Other | R | same |  | enum
| 1 |  | Channel | ContrastMethod | ChannelContrastMethod | This field describes the technique used to achieve contrast for this Channel. | enum | Brightfield, Darkfield, Differential Intensity Contrast (DIC), Fluorescence, HoffmanModulation, ObliqueIllumination, PhaseContrast, PolarizedLight, Other | R | same |  | enum |  |  |  |  | HAVE TO COMPARE WITH EXPERIMENT/IMAGING METHOD
| 1 |  | Channel | ExcitationWavelength |  | This field records the Excitation Wavelength used for this Channel. | float with unit:ExcitationWavelengthUnit |  |  | same |  | float
| 1 |  | Channel | EmissionWavelength |  | This field records the Emission Wavelength used for this Channel. | float with unit:EmissionWavelengthUnit |  |  | same |  | float
| 1 |  | Channel | Color |  | This field records the the Color used to render this Channel - encoded as RGBA. The default value is "-1" corresponding #FFFFFFFF, so solid white (it is a signed 32 bit value). | Color |  |  | same |  | Color
| 3 |  | Channel | SamplesPerPixel |  | This field is prevelantly used for RGB images and records the number of samples the detector takes to form each pixel value. [units:none] | PositiveInt |  |  | same |  | PositiveInt | CHANGE DESCRIPTION TO REFLECT THAT THIS IS PRIMARILY USED FOR RGB ONLY; TALK WITH OME ABOUT IT | Is e.g. used to represent RGB images
| 1 |  | Channel | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same |  | Ref
| 1 | Y | Channel | Fluorophore |  | This element records the Fluorophore whose Emission is recorded by this Channel. | Element |  | 1 … ∞ | Became Element |  | string
| 1 |  | Channel | LightPath |  | A description of the Light Path associated with the Acquisition of this Channel. | Element |  | 1 | same |  | Element
[10 empty] | AcquisitionMode |  |  | REMOVED because redundant with Illumination Type and Imaging Method |  |  |  | HAVE TO COMPARE WITH EXPERIMENT/IMAGING METHOD
[10 empty] | PinHoleSize |  |  | Moved to PinHoleSettings (ADVANCED)
[10 empty] | NDFilter |  |  | NDFilter is now a specialization of Filter-APE
[10 empty] | PockelCellSettings |  |  | PockelCellSettings is now an Element called ElectronOpticalDeviceSettings
[10 empty] | FilterSetRef |  |  | Moved to LightPath
[10 empty] | LightSourceSettings |  |  | Moved to LightPath
[10 empty] | DetectorSettings |  |  | Moved to LightPath
| 1 |  | Channel | ImagingMethod | ImagingMethodList | This filed describes the type of Light Microscopy Imaging Method performed in this case. Allowable terms were obtained from the Biological Imaging Methods Ontology (Fbbi; https://www.ebi.ac.uk/ols/ontologies/fbbi) Light Microscopy term (http://purl.obolibrary.org/obo/FBbi_00000345). | enum | FourDPlus, Coherent Anti-Stokes Raman Scattering (CARS), Confocal_Fluorescence_array-raster-scan, Confocal_Fluorescence_single-spot-raster-scan, Confocal_Fluorescence_slit-raster-scan, Confocal_Fluorescence_spinning disk, Confocal_Fluorescence_swept-field, Fluorescence Correlation Spectroscopy (FCS), Light Sheet Fluorescence Microscopy (LSFM), Multi-photon, Near Field Scanning Optical Microscopy (NSOM), Optical coherence tomography (OCT), Second-Harmonic Imaging Microscopy (SHIM), Spectral Imaging, Super-Resolution_GSD (Ground State Depletion Scanning), Super-Resolution_DOPI (Defocussed Oorientation and Position Imaging), Super-Resolution_FIONA (Fluorescence Imaging With One Nanometer Accuracy), Super-Resolution_PAINT (Point Accumulation for Imaging In Nanoscale Topography), Super-Resolution_PALM (Photoactivation Localization Microscopy), Super-Resolution_STORM (Stochastic Optical Reconstruction Microscopy), Super-Resolution_STED (Stimulated Emission Depletion), Super-Resolution_SIM (Structured Illumination Microscopy), Widefield_Fluorescence, Widefield_Fluorescence_FLIM (Fluorescence Lifetime Imaging), Widefield_Fluorescence_FLIP (Fluorescence Loss in Photobleaching), Widefield_Fluorescence_FRAP (Fluorescence Recovery After Photobleaching), Widefield_Fluorescence_FRET (Förster Resonance Energy Transfer), Widefield_Fluorescence_SPIM (Selective or Single Plane Illumination Microscopy), Widefield_Fluorescence_TIRF (Total Internal Reflection Fluorescence), Other | R | Moved here from Experiment | R | enum | MOVED FROM EXPERIMENT |  |  |  | PROPOSE instead of Type
| 1 |  | Channel | ImagingMethodTermAcession |  | This field records the Accession ID for the term used to describe the Imaging Method. A sub-classes of EFO:0000635 (http://www.ebi.ac.uk/efo/EFO_0000635) is recommended. | anyURI | [5 empty] | ADDED |  |  |  | PROPOSE TO ADD
| 1 |  | Fluorophore |  | Manufacturer Specs | This element describes the Fluorescent Label whose Emission is recorded by this Channel. |  |  | 1 … ∞
| 1 |  | Fluorophore | ID |  | A Unique Identifier for this component. |  |  | R
| 1 | Y | Fluorophore | Name |  | This field records a descriptive term that designates this Fluorophore as specified by the Manufacturer. For example, it might be the name of a specific Fluorescence Protein variant, or the name used by the Manufactuer for an organic dye. | string |  | R
| 1 | Y | Fluorophore | Type | FluorescentLabelType | This field records a term that describes the general category of the Fluorescent Label used in this case. | enum | Nucleic Acid_Aptamer motif, Nucleic Acid_Bacteriophage-derived RNA tag, Nucleic Acid_Chemical Modification, Nucleic Acid_dCad9, Nucleic Acid_Forced Intercalation, Nucleic Acid_Molecular beacon, Protein_Biarsenical dye, Protein_Biological fluorophore, Protein_Fluorescent protein, Protein_Immunological probe, Protein_Organic dyes, Protein_Self-labelling enzyme tag, Quantum dots, None, Other | R |  |  |  | ADDED |  |  |  | CONSIDER substituting with one of the following: ttp://purl.obolibrary.org/obo/MI_0373; http://purl.obolibrary.org/obo/MI_0856; http://purl.obolibrary.org/obo/MI_0857
| 1 | Y | Fluorophore | ExcitationWavelength |  | This field records the Excitation Wavelength used for this Channel. | float with unit:ExcitationWavelengthUnit |  | R | moved from Channel |  | float
| 1 | Y | Fluorophore | EmissionWavelength |  | This field records the Emission Wavelength used for this Channel. | float with unit:EmissionWavelengthUnit |  | R | moved from Channel |  | float
| 1 | Y | Fluorophore | Concentration |  | This field records the Concentration of the Fluorophore that was imaged for this Channel. | float with unit:FluorophoreConcentrationUnit
| 1 | Y | Fluorophore | FPBaseID |  | In case the Fluorophore is a Fluorescent Protein, this field records its FPBaseID (FPBase.org). | string
| 1 |  | Fluorophore | PrimaryPublicationID |  | If applicable, this field captures the a Unique Identifier for the Primary Publication describing this Fluorophore. Examples include, Digital Object Identifier (DOI), PubMed Central Identifier (PMCID), ArXiv.org ID etc. | string
| 1 | Y | LightPath |  |  | A description of the Ligth Path. |  |  | 1 |  |  |  | ADD LIGHT PATH and address ID issue |  |  |  | Introduce LightPathMap with order of all LightPath elements
| 1 |  | LightPath | ID |  | A Unique Identifier for this component. | LSID |  | R |  |  |  | LOOK INTO DAMIR SUGGESTION FOR MERGING LIGHT PATH TO MAP
| 1 |  | LightPath | Name |  | A User-defined Name for this component. | Denomination
| 1 | Y | LightPath | ExcitationFilterRef | Reference | This empty element refers to an ExcitationFilter component used in this Ligth Path. | Extension of Reference |  | 0 … 1 | same
| 1 | Y | LightPath | StandardDichroicRef | Reference | This empty element refers to a StandardDichroic component used in this Ligth Path. | Extension of Reference |  | 1 | was DichroicRef
| 1 | Y | LightPath | EmissionFilterRef | Reference | This empty element refers to an EmissionFilter component used in this Ligth Path. | Extension of Reference |  | 0 … 1 | same
| 1 |  | LightPath | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | same
| 2 |  | LightPath | LightSourceCouplingRef | Reference | This empty element refers to a LightSourceCoupling component used in this Ligth Path. | Extension of Reference |  | 0 … ∞ | moved from Channel
| 1 | Y | LightPath | FilterSetRef | Reference | This empty element refers to a Filter Set component placed in this Ligth Path. | Extension of Reference |  | 0 … ∞
| 1 | Y | LightPath | AdditionalFilterRef | Reference | This empty element refers to an additional Generic Filter component placed in this Ligth Path. | Extension of Reference |  | 0 … ∞
| 1 | Y | LightPath | OpticalApertureRef | Reference | This empty element refers to an Optical Aperture component placed in this Ligth Path. | Extension of Reference |  | 0 … ∞
| 1 | Y | LightPath | OpticaAssemblyRef | Reference | This empty element refers to an Optical Assembly component placed in this Ligth Path. | Extension of Reference |  | 0 … ∞
| 1 | Y | LightPath | PolarizationOpticsRef | Reference | This empty element refers to a PolarizationOptics component placed in this Ligth Path. | Extension of Reference |  | 0 … ∞
| 1 | Y | LightPath | LightSourceSettings | Settings | This element holds the Setting applied to a Light Source used in the Light Path during the acquisiton of this Channel as well as a reference to the LightSource (the ID attribute) in question. | Element |  | 1 …∞ | moved from Channel |  |  | MOVED TO LIGHT PATH | Move to LightPath
| 1 | Y | LightPath | AdditionalLensSettings | Settings | This element holds the Setting applied to a Lens used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Lens in question. | Element |  | 0 … ∞ | [7 empty] | ADD FUNCTION HERE as SETTING
| 2 |  | LightPath | AdditionalMirrorSettings | Settings | This element holds the Setting applied to a Mirror used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Mirror in question. | Element |  | 0 … ∞
| 1 | Y | LightPath | PrismSettings | Settings | This element holds the Setting applied to a Prism used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Prism in question. | Element |  | 0 … ∞
| 1 | Y | LightPath | ShutterSettings | Settings | This element holds the Setting applied to a Shutter used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Shutter in question. | Element |  | 0 … ∞
| 1 | Y | LightPath | DetectorSettings-APE | Settings | This element holds the Setting applied to the Detector used for the Acquisition of this Channel as well as a reference to the Detector (the ID attribute) in question. | Element |  | 1… ∞ | moved from Channel |  |  | MOVED TO LIGHT PATH | Move to LightPath
| 2 |  | LightPath | LightPathMap | Settings | This element holds a Map that records the order in which the different optical component associated with this Ligh Path are physically laid out. | Element |  | 0 … 1 |  |  |  | ADDED
| 2 |  | LightPathMap |  |  | This element holds a Map that records the order in which the different optical component associated with this Ligh Path are physically laid out. |  |  | 0 … 1 |  |  |  | ADDED
| 2 |  | OpticalComponentPosition |  |  | This element records the Position of invididual optical components in this Light Path. | Element |  | 1 ... ∞
| 2 |  | OpticalComponentPosition |  |  | This element records the Position of invididual optical components in this Light Path. |  |  | 1 ... ∞ |  |  |  | ADDED
| 2 |  | OpticalComponentPosition | ID (OpticalComponentID) |  | This is the ID of an existing Optical Component that is present in this Light Path. | LSID |  | R
| 2 |  | OpticalComponentPosition | Order |  | This field records the Positionl order in which this optical componet is found in this Light Path during the Acquisition of this Image. | PostiveInt with unit: none |  | R
| 1 | Y | LightSourceSettings |  | Settings | This element holds the Setting applied to a Light Source used in the Light Path during the acquisiton of this Channel as well as a reference to the LightSource (the ID attribute) in question. |  |  | 0 ... 1 | same |  |  | ADD REFERENCE TO MODULATION DEVICE
| 1 |  | LightSourceSettings | ID (LightSourceID) |  | A Unique Identifier for this component. | LightSourceID |  | R | same | R | LSID
| 1 |  | LightSourceSettings | Wavelength |  | This field records the nominal Wavelength of the excitation Light utilized in this case. | PositiveFloat with unit:WavelengthUnit |  | R | same |  | PositiveFloat with unit:WavelengthUnit
| 2 |  | LightSourceSettings | Attenuation |  | This field records the actual Attenuation of the excitation light power that was applied during the Acquisition of this image, expressed as a fractional value from 0.0 to 1.0. | PercentageFraction |  |  | same |  | PercentageFraction |  |  |  |  | Reference to Laser Modulation of various types will be put in LightSourceSettings
| 1 | Y | LightSourceSettings | IlluminationIntensity |  | This field records the nominal Illumination Intensity (also known as Illumination Power or Power Output) that was set for the Acquisition of this Image or was reported by the Image Acquisition Software. | float with unit:IlluminationIntensityUnit |  | R |  |  |  | CHANGED TO ALIGN WITH SUGGESTION FROM Michelle Itano, Paula Montero-Lopin and Lisa Cameron |  |  |  | Propose to call it Illumination Intensity (and use Power Settings in the description).
| 2 |  | LightSourceSettings | IntensityReportingPosition | LightSourceIntensityReportingPosition | This field records the Position where the reported Illumination Intensity value was obtained as reported by the Manufacturer. | enum | Back of the Objective, Back of the Microscope Stand, Source, Specimen, Other | R |  |  |  | ADDED
| 3 |  | LightSourceSettings | AttenuationMethod |  | This field records the Method that was used to Attenuate the Laser beam during the acquisiton of this Channel. | string | [5 empty] | CHANGED TO ALIGN WITH SUGGESTION FROM Claire Brown Michelle Itano, Paula Montero-Lopin and Lisa Cameron |  |  |  | Propose to add in response to Nathalie's comment on Specs, line 96; should we reference the Laser Modulation Device?
ADD REFERENCE TO LASER MODULATION?
| 1 | Y | FilterSet |  |  | A Filter Set is typically composed of one Excitation Filter, one Emission Filter and a Dichroic. They might either be pre-assebled by the Manufacturer and pruchased as a whole or they can be custome assembled on site depending on the needs of the user. |  |  | 0 … ∞ | not a ManufacturerSpecs extension | [6 empty] | I think we need and ID for this
| 1 |  | FilterSet | ID (FilterSetID) |  | A Unique Identifier for this component. | FilterSetID |  | R | same |  | LSID | ADDED ID
| 1 |  | FilterSet | ExcitationFilterRef | Reference | This empty element refers to an ExcitationFilter component used in this Ligth Path. | Extension of Reference |  | 0 … 1 | same | 0 … ∞ | FilterRef
| 1 |  | FilterSet | StandardDichroicRef | Reference | This empty element refers to a StandardDichroic component used in this Ligth Path. | Extension of Reference |  | 1 | same | R | Rer
| 1 |  | FilterSet | EmissionFilterRef | Reference | This empty element refers to an EmissionFilter component used in this Ligth Path. | Extension of Reference |  | 0 … 1 | same | 0 … ∞ | FilterRef
| 1 |  | FilterSet | Name |  | A User-defined Name for this component. | Denomination |  | R
| 1 | Y | AdditionalLensSettings |  | Settings | This element holds the Setting applied to a Lens used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Lens in question. |  |  | 0 … ∞
| 1 |  | AdditionalLensSettings | ID (LensID) |  | This is the Unique Identifier of the device used in this case. | LensID |  | R
| 2 |  | AdditionalLensSettings | Function |  | This field records the purpose for which this Additinal Lens was used during the Acquisiton of this Channel. | string
| 2 |  | AdditionalMirrorSettings |  | Settings | This element holds the Setting applied to a Mirror used in the Light Path during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Mirror in question. |  |  | 0 … ∞
| 2 |  | AdditionalMirrorSettings | ID (MirrorID) |  | This is the Unique Identifier of the device used in this case. | MirrorID |  | R
| 3 |  | AdditionalMirrorSettings | ActualMountedAngle |  | This field records the Angle at which this mirror was mounted during the Acqusition of this Channel with respect to the optical axis. | float with unit:none
| 1 | Y | PrismSettings |  | Settings | This element holds the Setting applied to a Prism used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Prism in question. |  |  | 0 … ∞
| 1 |  | PrismSettings | ID (PrismID) |  | This is the Unique Identifier of the device used in this case. | PrismID |  | R
| 3 |  | PrismSettings | PositionSetting | PositionSettingsList | This refers to the Wollaston Prism position that was set during the Acquisition of this Image. The Position must match the Wollaston Prism settings found at the base of the objective lens. | enum | L, M, H, N1, N2 | R
| 3 |  | PrismSettings | TransmittanceProfileFile | FileAnnotation | This element records a File in .csv or other suitable formats containing the Transmission Wavelength profile of this hardware component. | Element |  | 0 … ∞
| 3 |  | PrismSettings | ReflectanceProfileFile | FileAnnotation | This element records a File in .csv or other suitable formats containing the Reflection Wavelength profile of this hardware component. | Element |  | 0 … ∞
| 3 |  | PrismSettings | TransmittanceRangeSettings | TransmittanceRangeType | This element records the range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 3 |  | PrismSettings | ReflectanceRangeSettings | ReflectanceRangeType | This element records the range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 1 | Y | ShutterSettings |  |  | This element holds the Setting applied to a Shutter used in the Light Path during the acquisiton of this Channel as well as a reference (the ID attribute) to the Shutter in question. | [6 empty] | ADDED MOVE TO POINT DETECTOR
| 1 |  | ShutterSettings | ID (OpticalApertureID) |  | This is the Unique Identifier of the device used in this case. | OpticalApertureID |  | R |  |  |  | ADDED MOVE TO POINT DETECTOR
| 1 | Y | ShutterSettings | OpenDuringZMovement |  | This field captures whether this Shutter was open (True) or closed (False) during Z-movement between successive Z-planes in this Image. | boolean |  | R |  |  |  | ADDED to align with Material and Method
| 2 |  | ShutterSettings | ApertureDiameter |  | This field records the actual physical Aperture Diameter that was used for this Shutter during the Acquisition of this Channel. | float with unit:ApertureDiameterUnit | [5 empty] | ADDED MOVE TO POINT DETECTOR
| 2 |  | ShutterSettings | Mode | ShutterMode | This field records the operation mode that was applied to this Shutter during the Acquisition of this Channel. The Fast mode provides the quickest open and close action sequence for which the shutter is capable. The Soft mode gradually opens and closes the shutter to reduce vibration. Control of light intensity without affecting wavelength is accommodated by a Neutral Density shutter mode (for more details see https://www.microscopyu.com/applications/live-cell-imaging/the-automatic-microscope). | enum | Fast, Neutral density, Slow, Other |  |  |  |  | ADDED MOVE TO POINT DETECTOR
| 1 | Y | TransmissionWavelengthRangeSettings |  |  | This element records the Range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. |  |  | 0 … ∞ |  |  |  | ALIGN WITH TRANSMITTANCE and REFLETANCE AS THEY ARE IN SPECS
| 1 | Y | TransmissionWavelengthRangeSettings | CutOn |  | This field records the Wavelength(s) below which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R
| 1 | Y | TransmissionWavelengthRangeSettings | CutOff |  | This field records the Wavelength(s) above which there is less than 50% Transmittance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R
| 2 |  | TransmissionWavelengthRangeSettings | Transmittance |  | This field records the maximum amount of light this optical component is capable of transmitting in this Wavelength Range, at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | percentFraction with unit:none
| 3 |  | ReflectionWavelengthRangeSettings |  |  | This element records the range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. |  |  | 0 … ∞ |  |  |  | ALIGN WITH TRANSMITTANCE and REFLETANCE AS THEY ARE IN SPECS
| 3 |  | ReflectionWavelengthRangeSettings | CutOn |  | This field records the Wavelength(s) below which there is less than 50% Reflectance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R
| 3 |  | ReflectionWavelengthRangeSettings | CutOff |  | This field records the Wavelength(s) above which there is less than 50% Reflectance for this optical element. | positiveFloat with unit:WavelengthUnit |  | R
| 3 |  | ReflectionWavelengthRangeSettings | Reflectance |  | This field records the maximum amount of light this optical component is capable of reflecting in this Wavelength Range, at the peak Wavelength. It is represented as a fractional value from 0.0 to 1.0. | percentFraction with unit:none
| 1 | Y | DetectorSettings-APE |  | Settings | This element holds the Setting applied to the Detector used for the Acquisition of this Image as well as a reference to the Detector (the ID attribute) in question. |  |  | 0 ... 1 | same | 0 ... 1 | [5 empty] | Do we need to provide a FrameSize when it is already recorded in the Image?
| 1 |  | DetectorSettings-APE | ID (DetetectorID) |  | A Unique Identifier for this component. | DetectorID |  | R | same | R | LSID
| 1 | Y | DetectorSettings-APE | AnalogGain |  | This field reports the Analog Gain (also known as Inherent Gain or simply as Gain) that was applied during the Acquisition of this Image. The Analog Gain is a property Inherent to the Detector and it specifies how many electrons correspond to 1 signal unit (photon) detected. Gain is reported in terms of electrons/ADU (analog-to-digital unit). A gain of 8 means that the Detector digitizes the light signal so that each ADU corresponds to 8 photoelectrons. NOTE: In the case of Point Detectors the value of Analog Gain can be set freely by the user. Conversely, in the case of Cameras, it is common for the Gain to be preset by the Manufacturer to minimize noise. In this case, either the Analog Gain cannot be changed or the user is given the choice of selecting between a few preset Analog Gains values upon Acquisition. Finally, several Cameras might have so-called Automatic Gain Control (AGC or AutoGain) that will automatically adjusts so the full range of the Analog-to-Digital Converter (ADC) is used. Unless you can turn off the AGC, the use of these Cameras are highly discouraged for quantitative work. | float with unit:none |  | R | was Gain |  | float | CHANGED DESCRIPTION |  |  |  | PROPOSE TO CHANGE BACK TO GAIN AND USE ANALOG GAIN IN DESCRIPTION
| 1 | Y | DetectorSettings-APE | Offset |  | If available, this field reports the actual Offset that was set for this Detector during the Acquisition of this Channel. Offset is an electronic baseline added to the output signal generated by the sensor to ensure that the resulting signal level always correspond to a positive number of counts. Offset defines the black level setting of the image and it should be adjusted so that the lowest signals are near the detection threshold. NOTE: Please note that most modern Cameras do not have the possibility of setting the Offset. | float with unit:none |  |  |  |  | float | MOVED from Point Detector | Make not required: modern detectors/cameras no longer need this. |  |  | IF NOLONGER REQUIRED FOR CAMERAS WE SHOULD MOVE IT POINT DETECTORS NO?
Zoom moved to ScannerSettings (Confocal extension)
| 3 |  | DetectorSettings-APE | OffsetMode | DetectorOffsetMode | This field describes the way in which the Voltage Offset of the Detector is reported. It typically is expressed as an Arbitrary-units number, a Scaled percent (-100 to 100), or a a Voltage. | enum | Arbitrary-units, Scaled percentage, Voltage, Other |  |  |  |  | MOVED from Point Detector |  |  |  | THIS IS PROBABLY NOT A SETTING RIGHT? ALSO IS SOUNDS LIKE SOMEWHAT OF AN OVERKILL RIGHT?
| 1 | Y | DetectorSettings-APE | AmplificationGain |  | This field reports the Amplification Gain that was applied during the Acquisition of this Image. Amplification Gain in general measures of the ability of a Detector to add energy to the input light signal therefore increasing its power or amplitude. Higher Amplification Gain levels amplify the signal and result in greater brightness, contrast and background signal. Examples include: 1) the Electron Multiplication (EM) gain (as opposed to the inherent Analog Gain) for an EMCCD camera. 2) The Gain introduced by the Micro-Channel Plate in an Intensified CCD or CMOS camera. 3) The amplification of the electron current provided by a Photomultiplier Tube, and measured as the ratio of the input photoelectron current emitted by the photocatode and the current reaching the anode. 4) The signal amplification provided by an Avalache Photo Diode (APD). | float with unit:none |  | R |  |  |  | CHANGED TO INCLUDE OTHER TYPES OF AMPLIFICATION, including EM GAIN (EMCCD), MICROCHANNEL PLATE GAIN (intensified CCD or CMOS) and PMT GAIN
| 1 | Y | CameraSettings |  |  | This element refers to the Camera used to acquire this Channel and it holds the Settings that were applied to the Camera during this Acquisition.
| 1 | Y | CameraSettings | Binning | Binning | This field represents the number of adjacent pixels whose electric charge is combined in the resulting Image (e.g. binning = 2x2). Generally, binning increases read-out rates, signal-to-noise ratio, and relative sensitivity, while it reduces the spatial resolution. | enum | 1x1, 2x2, 4x4, 8x8, Other | R | specialized to CameraSettings
| 1 | Y | CameraSettings | PixelReadOutRate |  | This represents the effective Pixel Readout Rate used for the Acquisition of this Image. This value is defined as the inverse of the time required to digitize a single pixel (also referred to as serial conversion time). Pixel readout rate can be considered similar to a baud rate as it depends on the bytes per second that can be read from the detector and written on disk. Readout rates are usually given in pixels/second (e.g., 500 kilopixels/second). | float with unit:ReadOutRateUnit |  | R | specialized to CameraSettings
| 2 |  | CameraSettings | Integration |  | This field represents the number of sequential frames that have been averaged to obtain this image, to improve the overall signal-to-noise ratio. | integer with unit:none |  |  | specialized to CameraSettings |  |  |  | Is this different from exposure time? |  | PROPOSE TO ADD (BACK) INTEGRATION BASE (line, frame etc.) and INTEGRATION METHOD (AVERAGE ETC.) | Moved integration from Detector; Different aggregation modalities other than Average?
| 1 | Y | CameraSettings | DigitizerBitDepth | DigitizerType | If selectable, this field reports the Bit Depth of the Digitizer used for the Acquisition of this Channel. | enum | 12bit, 16bit, Other | R
| 1 | Y | CameraSettings | SensorExposureMode | CMOSSensorExposureMode | In case of a CMOS Camera and if selectable, this field reports the sensor Exposure modality used for the Acquisition of this Channel. In the Rolling Exposure mode adjacent rows of the array are exposed at slightly different times as the readout 'waves' sweep through each half of the sensor. In the Global Exposure mode (i.e., 'Snapshot' exposure mode) all pixels of the array are exposed simultaneously, thus enabling 'freeze frame' capture of fast moving or fast changing events. In this modality the CMOS Camera behaves like an interline CCD sensor. | enum | Global shutter, Rolling shutter, Other | R
| 3 |  | CameraSettings | FrameRate |  | Frame Rate is the inverse of the time needed for the CCD to acquire a single Frame (or Plane) and then completely read that image out. Frame Rate is typically expressed in frames per second (fps). Often, the Frame Rate can be approximately calculated from the total number of pixels and the readout rate, combined with the total exposure time. This field represents the effective Frame Rate used for the Acquisition of this Image. | float with unit:FrameRateUnit |  | R
| 3 |  | CameraSettings | OperatingTemperature |  | This field represents the effective temperature the sensor is regulated to during the Acquisition of this image. | float with unit:TemperatureUnit; [default: Celsius ºC]
| 1 | Y | PointDetectorSettings |  |  | This element refers to the Point Detector used to acquire this Channel and it holds the Settings that were applied to the Point Detector during this Acquisition.
| 1 | Y | PhotoMultiplierSettings | Voltage |  | This field refers to the Voltage settings applied to the Photomultiplier Tube (PMT) for the Acquisition of this Channel (e.g., voltage = 244). The higher the supplied Voltage, the greater the amplification of the current generated by the PMT. While this increases the sensitivity, it also increases the background in the resulting image. | float with unit:HighVoltageUnit |  | R | specialized to PhotoMultiplierSettings |  | float | CHANGED
| 1 | Y | PointDetectorSettings | PhotonCounting |  | This boolean indicates whether this Photomultiplier Tube was operated in photon-counting mode (True) or not (False), during the Acquisition of this Chanel. | boolean | [5 empty] | ADDED MOVE TO POINT DETECTOR | Need a boolean whether the PMT is operated in photon-counting or gain mode |  |  | Can this be applied also to other point detectors?
[14 empty] | diskussion re: LightPath configuration: How about encoding the entire lightpath per Channel instance as a linked-list of refs to the components following the travel of the photons. E.g. for fluorescence: start with the lightsource, followed by any beam-shaping optics, excitation filter(s), dichroic, objective, dichroic (again), (possible secondary dichroic), emission filter(s), camera. E.g. for DIC: start with lightsource, polarizer, wollaston, condenser lens, objective, wollaston, analyzer, camera. Question how to deal with Channel instances that have a real forked or joined lightpath such as multi-lightsheet illumination or a True-color camera system. |  |  | IT SEEMS LIKE I IS ALREADY THIS WAY? Could you clarify?
~
Note on A1: ====== ID#AAAAEEcLF7s (2019-12-18 22:28:54) This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on D1: ====== ID#AAAAEEZUM88 Caterina Strambio De Castillia (2019-12-19 19:05:02) This essentially means that this Element is not actually used per se. It is defined as the Ancestor of Elements that are actually created when representing Hardware components. The advantage is this way all Children elements will inherit all Attributes of this Abstract ancestor and one avoids to have to repeat the same Attributes multiple times. It is a modeling trick that avoids data duplication.
Note on B3: ====== ID#AAAAEEcLF7k (2019-12-18 22:28:54) This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on U61: ====== ID#AAAAKFd-mq0 Nathalie Gaudreault (2020-08-11 01:25:04) was settled with Damir
Note on E192: ====== ID#AAAAKofQy_Q tc={F86F945F-8F9A-7141-A4A8-D3A44ECCFC18} (2020-11-07 03:00:40) [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This might need to be modified to take into account NA adjustments. Also in other cases it appears that the collar might have markings such as “Oil”, Water etc.

## 4DN-BINA Advanced+Confocal Ext.
4DN-OME Model |  |  | APE = Abstract Parent Element
[14 empty] | BINA WORKING GROUP | [6 empty] | Methods checklist |  |  | QUAREP comments
| Tier | M&M (Montero, Brown, Cameron, Itano and Strambio-De-Castillia) | Element | Attribute/sub-Element | Complex type | Description | Data type | Allowed values | Cardinality | OME 6-2016 | OME Cardinality/Required? | OME Data type | DECISION | Damir Comments | Arturo comments | Pablo Comments | Caterina comments | Jim comments | Nathalie comments | Ulrike comments | Michelle Itano | Lisa Cameron | Paola Motero Llopin | WG7 Member 1 [Enter your name and Choose a color] | WG7 Member 2 [Enter your name and Choose a color] | WG7 Member n [Enter your name and Choose a color]
SPECS | 1 |  | Instrument |  |  | This element describes the Instrument used to capture the Image. It is primarily a container referring to the Microscope's components list, which include the Microscope Stand, Microscope Table, Light Source, Light SourceCoupling, SampleHolder, Focusing devices, Objective, Optovar, Filter, Dichroic, AdditionalOptics, Detector and CalibrationTool. The Filter element can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. Within the Image itself, a reference is made to this one Filter element. There may be multiple Light Sources, detectors, Objectives and filters etc. in an Instrument. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the Ligth Path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. |  |  | 1
| 1 | Y | TIRFHardwareModule |  | ManufacturerSpecs | Total Internal Reflection Fluorescence (TIRF) microscopy is a modality in which fluorescence is excited in a thin layer by an evanescent wave produced by total internal reflection. This element describes a Hardware Module that is used to achieve TIRF microscopy.
| 1 |  | TIRFHardwareModule | ID |  | A Unique Identifier for this component. | LSID |  | R
| 1 | Y | TIRFHardwareModule | Method | TIRFMethod | This field captures the approach used by this TIRF Hardware Module to produce TIRF illumination. In Objective-based TIRF, the excitation lightpath and the emission channel share the same optics, which provides high signals but might result in high background. In Prism-based TIRF excitation light enters the dovetail prism, reflects at the glass/water interface and escapes through the opposite facet without interferring with the emission channel . In Lightguide-based TIRF, excitation light enters the lightguide, reflects multiple times from its top and bottom surfaces to generate wide-area evanescent waves and escapes from the opposite side of the lightguide without interferring with the emission channel. | enum | Lightguide, Objective, Prism, Other | R |  |  |  | CHANGED
| 1 | Y | TIRFHardwareModule | Ring |  | This field captures whether (True) or not (False) this TIRF Hardware Module employs Ring (also known as spinning-spot) TIRF. | boolean |  | R
| 2 |  | TIRFHardwareModule | Motorized |  | This field captures whether this TIRF Hardware Module employs a Motorized device to set the TIRF Angle (True) or not (False). | boolean
| 1 | Y | TIRFHardwareModule | PrismRef |  | This empty element refers to a Prism component placed used in this case to produce TIRF microscopy. | Extension of Reference
| 1 |  | TIRFHardwareModule | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 | Y | ConfocalScanner-APE |  | Manufacturer Specs | According to ISO-10934-2-2007, a confocal microscope is a microscope in which, ideally, a point in the object plane is illuminated by a diffraction-limited spot of light, and light emanating from this point is focused upon and detected from an area smaller than the central area of the diffraction disc. The confocal principle leads to improved axial resolution by suppression of light from out-of-focus planes. In order to utilize this modality, a Confocal Scanner is used to form an Image of an extended Specimen region by controlling the direction of an outgoing Laser beam, the position of the Specimen or the location of the illuminated and detected spots simultaneously. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. |  |  | 0 … 1 |  |  |  | CHANGED | I changed a few small things in order to accommodate spinning disk, swept-field, grid multi-point, etc. scanners
| 1 |  | ConfocalScanner-APE | ID (ConfocalScannerID) |  | A Unique Identifier for this component. | ConfocalScannerID |  | R
| 1 |  | ConfocalScanner-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | ConfocalScanner-APE | NrOfScanPoint |  | This field records the number of scanning points available for this scan head. | integer with unit:none |  | R
| 2 |  | ConfocalScanner-APE | MaxFrameRate |  | This field captures the Maximum full Frame Rate available for this device, as specificed by the Manufacturer. | float with unit:MaxFrameRateUnits | [5 empty] | ADDED | added for multi-spot scanners
| 3 |  | ConfocalScanner-APE | PointsConfiguration |  | This field records a description of the arrangement of scan points (can be spiral, linear, grid, other) | string | [5 empty] | ADD
| 1 |  | ConfocalScanner-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 | Y | RasterScanner |  |  | A Raster Scanner is typically used for systematic periodic scanning (raster scanning) of the Laser beam or of the Specimen position, either in one or in two dimensions. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods.
| 1 | Y | RasterScanner | Type | RasterScannerType | This field records the type of this Raster Scanner. In a Galvanometer-mirror based system, two galvanometer scan mirrors (one for each axis) direct the laser beam to any XY position in the field of view. In a Resonant system laser positioning in the x-axis is provided by a special resonant scan mirror which oscillates at a fixed frequency. | enum | Galvanometer-mirror, Resonant, Other | R
| 2 |  | RasterScanner | HorizontalScanningSpeed |  | This field records the maximal Horizontal Scanning Speed of this device. | float with unit:HorizontalScanningSpeed
| 2 |  | RasterScanner | VerticalScanningSpeed |  | This field records the maximal Vertical Scanning Speed of this device. | float with unit:VerticalScanningSpeed
| 1 | Y | SpinningDisk |  |  | A Spinning disk confocal scanner is a device in which the scanning of the illuminated and detected spots is performed simultaneously using an opaque disk called a Nipkow disk, which has many ideally identical small holes arranged in Archimedean spirals. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods.
| 1 | Y | SpinningDisk | SingleDisk |  | This records whether this Spinning Disk Confocal Scanner uses Single-Disk (True) or Dual-Disk (False) technology. | boolean
| 1 | Y | SpinningDisk | DiskSpeed |  | This field records the maximal Disk rotation Speed of this device, according to the Manufactuer. | float with unit:DiskSpeedUnit
| 1 |  | AcoustoOpticalDevice-APE |  | ManufacturerSpecs | This is a device that performs spatial, temporal, and spectral modulations of light on the basis of acousto-optic (AO) diffraction. AO diffraction is a phenomenon whereby an acoustic wave propagating in an optically transparent medium, produces a periodic modulation of the index of refraction of the medium via the elasto-optical effect and might therefore diffract portions of an incident light into one or more directions. | Substitution : choose one of the children elements from the list below | AcoustoOpticalBeamSplitter, AcoustoOpticalDeflector, AcoustoOpticalModulator, AcoustoOpticalTuneableFilter | 0 … ∞
| 1 |  | AcoustoOpticalDevice-APE | ID (AcoustoOpticalDeviceID) |  | A Unique Identifier for this component. | Acousto-OpticalDeviceID |  | R
| 1 |  | AcoustoOpticalDevice-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | AcoustoOpticalDevice-APE | AttenuationCoefficient |  | This field records the attenuation coeffieient of this optical element. The attenuation coefficient (aka blocking coefficient) is used as a measure of the out-of-band attenuation of an optical filter or dichroic mirror over an extended range of the spectrum. The attenuation coefficient is often defined in units of optical density (OD), whereas OD= -log10(T), and T=Transmittance (0≤T≤1). | float with unit:none |  | R
| 2 |  | AcoustoOpticalDevice-APE | AttenuationMethod | AttenuationMethodType | This field records the mean by which this optical element achieves out-of-band transmission attenuation. | enum | Absorption, Reflection, Rifraction, Other
| 2 |  | AcoustoOpticalDevice-APE | ContrastRatio |  | This field records the reports the ratio of maximum versus minimum transmittance. | PercentFraction
| 2 |  | AcoustoOpticalDevice-APE | Coating |  | An optical Coating consists of one or more thin layers of material deposited on an optical component such as a Lens or Mirror, which alters the way in which the optic reflects and transmits light. This field records the thin-film material used to coat the non linear crystal and reduce surface reflection. | string
| 3 |  | AcoustoOpticalDevice-APE | TuningAccuracy |  | This field measures the accuracy with which the Wavelength diskriminating capacity of this device can be "tuned" or changed. | float with unit:none |  | R
| 3 |  | AcoustoOpticalDevice-APE | Polarization | PolarizationMode | This field records the Polarization effect of this optical element. Dichroic beamsplitters (and, in fact, any thin-film interference coating that is used at non-normal angles-of-incidence) will cause some amount of polarization, the precise effect varying greatly with wavelength and with the particular coating design. | enum | Planar, Circular, Elliptical, Other | R
| 3 |  | AcoustoOpticalDevice-APE | Bragg |  | This field indicates whether this device is operated under Bragg regime (True) or not. In Bragg regime, at one particular incidence angle , only one diffraction order is produced - the others are annihilated by destructive interference. | boolean |  | R
| 3 |  | AcoustoOpticalDevice-APE | TuningResolution |  | This field measures the wavelength diskriminating capacity of this device and it is defined as the full width at half-maximum (FWHM) of the main lobe from the selected optical output, and is a function of both wavelength and device configuration. | float with uint:TuningResolutionUnit
| 3 |  | AcoustoOpticalDevice-APE | TuningRate |  | This field records the time it takes for this device to be tuned without degrading resolution. This time is determined by the velocity of the acoustic wage as it traverses the aperture of this device. | float with unit:TuningRateUnit | [10 empty] | Not very farmiliar with AO tech...is this same as possible pulse frequency?
| 3 |  | AcoustoOpticalDevice-APE | AngleOfIncidence |  | This field records the optimal angle of incidence (AOI) for this mirror. TheAOI is the angle between the optical axis of the incident light and the axis normal to the surface of the mirror. In specular reflection AOI is by definition equal to the Angle Of Reflection (AOR). | float with unit:AngleOfIncidenceUnit
| 3 |  | AcoustoOpticalDevice-APE | AngularAperture |  | This field records the angular aperture is the half-angle of the cone of light entering the this device. It is, therefore, a measure of how much of the incident light actually passes through the device. | float with unit:none | [10 empty] | Is this angle of incidence?
| 3 |  | AcoustoOpticalDevice-APE | Material | AODMaterial | This field describes the transparent crystal or glass material used to build this device, through which the light propagates and whose refractive properties can be modulated by acoustic waves of different frequencies. | enum | Chalcogenide glasses, Crystalline quartz, Gallium phosphide, Germanium, Fused silica, Indium phosphide, Lithium niobate, Tellurium dioxide (TeO2), Other
| 3 |  | AcoustoOpticalDevice-APE | Design | AODDesign | This field describes wheher the acoustic wave is absorbed or reflected by the distal surface of the crystal. | enum | Resonant, TravelingWave, Other
| 3 |  | AcoustoOpticalDevice-APE | Geometry | AODGeometry | This field describes the the direction of the applied electric field with respect to the light beam. Longitudinal devices have the electric field in the direction of the light beam, which passes through holes in the electrodes. Transverse devices have the electric field perpendicular to the light beam. | enum | Longitudinal, Transverse
| 3 |  | AcoustoOpticalDevice-APE | OperatingTemperature |  | This field records the ambient temperature at which this device is designed to operate, according to the Manufacturer. | float with unit:TemperatureUnit; [default: Celsius ºC]
| 1 |  | AcoustoOpticalDevice-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 2 |  | AcoustoOpticalDevice-APE | WavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 1 ... ∞
| 1 |  | AcoustoOpticalBeamSplitter |  |  | This is actually an AOTF used in an imaginative manner by Leica to replace the dichroic mirror usually used to separate the illumination and the detection paths in a confocal fluorescence Microscope. In practice the AOBS is programmed so that it leaves most of the incoming beam of light undeflected, and it only deflects light originating from specific laser lines. The AOBS allows the simultaneous deflection of multiple laser lines simultaneously towards the sample, while at the same time controlling the intensity of each line independently. | [12 empty] | Also refered to as an AOTF on Zeiss spinning disk systems.
| 2 |  | AcoustoOpticalDeflector |  |  | Acousto-optic deflectors are devices which can be used to deflect a laser beam in one direction by a variable angle, controlled by the frequency of an electrical signal. Essentially, such a device is an acousto-optic modulator which is operated with an electrical drive signal of constant power but a variable frequency. The direction of the diffracted beam is determined by a Bragg condition, which contains the sound wavelength and thus depends on the sound frequency. | [6 empty] | STET |  |  |  | PROPOSE TO REMOVE IT
| 2 |  | AcoustoOpticalModulator |  |  | An acousto-optic modulator (AOM) is an electronically-tunable device which can be used for controlling the direction and/or intensity of a laser beam by an acoustically-induced diffraction grating in a crystal. It is based on the acousto-optic effect, i.e. the modification of the refractive index of some crystal or glass material by the oscillating mechanical strain of a sound wave (photoelastic effect).
| 2 |  | AcoustoOpticalModulator | Type | AOMType | This field describes the category of this device. | enum | Frequency shifters, Intercavity, Simple intensity, Other
| 1 |  | AcoustoOpticalTuneableFilter |  |  | This is kind tuneable filter, whose optical filtering properties can be tuned by virtue of acousto-optical deflection. As such, different wavelengths exits the crystal at different angles depending on the applied radio wave frequency. In an optical system in which the input light beam strikes the crystal at a specific angle and can only re-enter the system at a specific angle, then one can choose the wavelength of the light diffracted at the output angle angle by varying the frequency of the acoustic wave.
| 1 |  | ElectroOpticalDevice-APE |  | Manufacturer Specs | This is is a device which can be used for controlling the power (→ intensity modulators), phase (→ phase modulators) or polarization of light with an electrical control signal. It typically contains one or two Pockels cells, and possibly additional optical elements such as polarizers. | Substitution : choose one of the children elements from the list below | DeformableMirror, ElectroOpticalModulator, LiquidCrystalTuneableFilter, SpatialLightModulator
| 1 |  | ElectroOpticalDevice-APE | ID (ElectroOpticalDeviceID) |  | A Unique Identifier for this component. | ElectroOpticalDeviceID |  | R
| 1 |  | ElectroOpticalDevice-APE | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | ElectroOpticalDevice-APE | AttenuationCoefficient |  | This field records the attenuation coeffieient of this optical element. The attenuation coefficient (aka blocking coefficient) is used as a measure of the out-of-band attenuation of an optical filter or dichroic mirror over an extended range of the spectrum. The attenuation coefficient is often defined in units of optical density (OD), wherease OD= -log10(T), and T=Transmittance (0≤T≤1). | float with unit:none |  | R
| 2 |  | ElectroOpticalDevice-APE | AttenuationMethod | AttenuationMethodType | This field records the mean by which this optical element achieves out-of-band transmission attenuation. | enum | Absorption, Reflection, Refraction, Other
| 2 |  | ElectroOpticalDevice-APE | ModulationMode | ModulationModeList | This field records the describes the type of modulation this devices produces on a beam of light. | enum | Intensity, Phase, Polarization, Wavelength, Other
| 2 |  | ElectroOpticalDevice-APE | ContrastRatio |  | This field records the reports on the ratio of maximum versus minimum transmittance. | float with unit:none
| 2 |  | ElectroOpticalDevice-APE | Coating |  | An optical Coating consists of one or more thin layers of material deposited on an optical component such as a Lens or Mirror, which alters the way in which the optic reflects and transmits light. This field records the thin-film material used to coat the non linear crystal and reduce surface reflection. | string
| 3 |  | ElectroOpticalDevice-APE | TuningAccuracy |  | This field measures the accuracy with which the wavelength diskriminating capacity of this device can be "tuned" or changed. | float with unit:TuningAccuracyUnit |  | R
| 3 |  | ElectroOpticalDevice-APE | Polarization | PolarizationMode | This field records the polarization effect of this optical element. Dichroic beamsplitters (and, in fact, any thin-film interference coating that is used at non-normal angles-of-incidence) will cause some amount of polarization, the precise effect varying greatly with wavelength and with the particular coating design. | enum | Planar, Circular, Elliptical, Other | R
| 3 |  | ElectroOpticalDevice-APE | TuningResolution |  | This field measures the wavelength diskriminating capacity of this device and it is defined as the full width at half-maximum (FWHM) of the main lobe from the selected optical output, and is a function of both wavelength and device configuration. | float with uint:TuningResolutionUnit
| 3 |  | ElectroOpticalDevice-APE | TuningRate |  | This field records the time it takes for this device to be tuned without degrading resolution. This time is determined by the velocity of the acoustic wage as it traverses the aperture of this device. | float with unit:TuningRateUnit
| 3 |  | ElectroOpticalDevice-APE | AngleOfIncidence |  | This field records the optimal angle of incidence (AOI) for this mirror. TheAOI is the angle between the optical axis of the incident light and the axis normal to the surface of the mirror. In specular reflection AOI is by definition equal to the Angle Of Reflection (AOR). | float with unit:AngleOfIncidenceUnit
| 3 |  | ElectroOpticalDevice-APE | AngularAperture |  | This field records the angular aperture is the half-angle of the cone of light entering the this device. It is, therefore, a measure of how much of the incident light actually passes through the device. | float with unit:none
| 3 |  | ElectroOpticalDevice-APE | OperatingTemperature |  | This field records the ambient temperature at which this device is designed to operate, according to the Manufacturer. | float with unit:TemperatureUnit; [default: Celsius ºC]
| 3 |  | ElectroOpticalDevice-APE | OperatingHumidity |  | This field records the ambient humidity at which this device is designed to operate, according to the Manufacturer. | PercentFraction
| 3 |  | ElectroOpticalDevice-APE | WarmUpTime |  | This field records the time it takes for this device to reach operating conditions according to the Manufacturer. | float with unit:WarmUpTimeUnit; [default min]
| 1 |  | ElectroOpticalDevice-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 2 |  | ElectroOpticalDevice-APE | WavelengthRange |  | This element records a range of Wavelengths this device is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 1 ... ∞
| 2 |  | DeformableMirror |  |  | A Deformable Mirror (DM) is a Mechano Electro-Optical Device that has a controllable spatially varying mirror surface to modulate the wavefront of a beam of light impinging on it. This type of device is typically used in adaptive optics. | [6 empty] | ADDED | added
| 2 |  | DeformableMirror | Type | DeformableMirrorType | This field records the type of this Deformable Mirror. | enum | Bimorph, FerrroFluid, MEMS, Magnetic, Membrane, Segmented, Other | R
| 2 |  | DeformableMirror | Actuators |  | This field records the number of Actuators available for this device. | integer with unit:none
| 3 |  | DeformableMirror | ActuatorPitch |  | This field records the distance between the center of each Actuator in this device. | float with unit:ActuatorPitchUnit
| 3 |  | DeformableMirror | ActuatorStroke |  | This field records the maximum possible Actuator displacement available for this device. | float with unit:ActuatorStrokeUnit
| 3 |  | DeformableMirror | ResponseTime |  | This field records the rate at which this Deformable Mirror can change patterns. | float with unit:ResponseTimeUnit
| 2 |  | ElectroOpticalModulator |  |  | An Electro-Optical Modulator (EOM) is a device which can be used for controlling the power (→ intensity modulators), phase (→ phase modulators) or polarization of light with an electrical control signal.
| 2 |  | ElectroOpticalModulator | PockelsCell |  | Pockels cells are the basic components of electro-optic modulators. A Pockels cell is a device consisting of an electro-optic crystal (with some electrodes attached to it) through which a light beam can propagate. The phase delay in the crystal (→ Pockels effect) can be modulated by applying a variable electric voltage. The Pockels cell thus acts as a voltage-controlled waveplate. | Element |  | 0 ... ∞
| 2 |  | PockelsCell |  |  | Pockels cells are the basic components of electro-optic modulators. A Pockels cell is a device consisting of an electro-optic crystal (with some electrodes attached to it) through which a light beam can propagate. The phase delay in the crystal (→ Pockels effect) can be modulated by applying a variable electric voltage. The Pockels cell thus acts as a voltage-controlled waveplate.
| 2 |  | PockelsCell | HalfWaveVoltage |  | This field record the Voltage required for inducing a phase change equivalent to a half an optical wavelength. In an amplitude modulator, the applied voltage has to be changed by this value in order to go from the operation point with minimum transmission to that with maximum transmission. The HalfWaveVoltage can be referred to as, Uπ, Uλ/2 or Vλ/2. | float with unit:HalfWaveVoltageUnit |  | R
| 3 |  | PockelsCell | Material | PockelsCellMaterial | This field describes the nonlinear crystal materials used to contruct this Pockels cell. | enum | Ammonium dihydrogen phosphate, β-barium borate, Lithium niobate, Lithium tantalate, Potassium di-deuterium phosphate, Potassium titanyl phosphate, Other
| 1 |  | LiquidCrystalTunableFilter |  |  | This element describes a liquid crystal tunable filter (LCTF) is an optical filter that uses electronically controlled liquid crystal (LC) elements to transmit a selectable wavelength of light and exclude others.
| 2 |  | LiquidCrystalTunableFilter | SelectableBandwidth |  | This field indicates whether the bandwidth of this tuneable filter can be selected (True) or it is fixed (False). | boolean |  | R
| 3 |  | LiquidCrystalTunableFilter | MaximumPower |  | This field describes the maximum Power this device can tolerate (typically expressed in Joules). | float with unit:MaximumPowerUnit | [5 empty] | ADDED |  |  |  |  | Do we want to list upper bound for light intensity limit?
| 2 |  | SpatialLightModulator |  |  | A Spatial Light Modulator (SLM) is device that imposes a changeable spatially varying modulation on a beam of light. The modulation can be in Intensity or in Phase or both. | [7 empty] | added
| 2 |  | SpatialLightModulator | Type | SLMType | This field records the category to which this Spatial Light Modulator belongs. | enum | Digital Micromirror Device, Electrically-addressed SLM, Liquid Crystal on Silicon, Liquid Crystal Display, Optically-addressed SLM, Other | R
| 2 |  | SpatialLightModulator | ArrayWidth |  | This field records the number of spatial elements (pixels) this Spatial Light Modulator contains in the X direction. | NonNegativeInt
| 2 |  | SpatialLightModulator | ArrayHeight |  | This field records the number of spatial elements (pixels) this Spatial Light Modulator contains in the Y direction. | NonNegativeInt
| 3 |  | SpatialLightModulator | ResponseTime |  | This field records the rate at which this Spatial Light Modulator can change patterns. | float with unit:ResponseTimeUnit
| 2 |  | Microlens Array |  | ManufacturerSpecs | Micro-lens arrays contain multiple micro-lenses forming a one- or two-dimensional array on a supporting substrate. | [6 empty] | MOVE IT TO ITS OWN CATEGORY
| 2 |  | Microlens Array | ID (MicrolensArrayID) |  | A Unique Identifier for this component. | MicrolensArrayID |  | R
| 2 |  | Microlens Array | Name |  | A User-defined Name for this component. | Denomination
| 2 |  | Microlens Array | LensletDiameter |  | This field records the physical Dameter of the Lenslets that compose this Micro Lens Array. | float with unit:DiameterUnit; [default mm] |  | R
| 2 |  | Microlens Array | LensletPitch |  | This field records the distance between Lenslets in this Micro Lens Array. | float with unit:DistanceUnit [default: mm] |  | R
| 2 |  | Microlens Array | LensletArrangement |  | This field records the arrangement of individual Lenslets on this Array. | string |  | R |  |  |  | ADDED
| 2 |  | Microlens Array | LensletShape | LensletShape | This field records the shape of the individual Lenslets composing this Micro Lens Array. | enum | Circular, Hexagonal, Octagonal, Square, Other
| 2 |  | Microlens Array | LensletType | LensType | This field describes the type of the individual Lenslets that compose this Micro Lens Array. | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 2 |  | Microlens Array | Coating |  | An optical Coating consists of one or more thin layers of material deposited on an optical component such as a Lens or Mirror, which alters the way in which the optic reflects and transmits light. This field records the thin-film material used to coat the non linear crystal and reduce surface reflection. | string
| 3 |  | Microlens Array | Reflectance |  | This field records the maximum amount of light this Micro Lens Array is capable of reflecting. It is represented as a fractional value from 0.0 to 1.0. | float with unit:none |  | R
| 3 |  | Microlens Array | SubstrateMaterial | SubstrateMaterialName | This field records the chemical composition or the nonproprietary trade designation of the substrate of this Micro Lens Array. The substrate is typically composed of a glass material with low autofluorescence, such as UV-grade fused silica. | enum | Calcium Fluoride, N-BK7 (Borosilicate Crown), UV-Grade Fused Silica (UVFS), Zerodur (Lithium Aluminum Ailicon Oxide), Zinc Selenide (ZnSe), Other
| 3 |  | Microlens Array | FrontFocalLength |  | This field records the distance (f) measured along the optical axis from the front principal plane of this Micro Lens Array to the front focal plane (F). | float with unit:OpticalDistanceUnit; [default mm]
| 3 |  | Microlens Array | BackFocalLenght |  | This field records the distance (f') measured along the optical axis from the rear principal plane of this Micro Lens Array to the rear focal plane (F'). | float with unit:OpticalDistanceUnit; [default mm]
| 2 |  | Microlens Array | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 2 |  | Microlens Array | WavelengthRange |  | This records a range of wavelengths this device is designed for, according to the Manufacturer. Multiple of this elements may be used. | Element |  | 1 ... ∞
| 1 | Y | OpticalAperture-APE |  | Manufacturer Specs | This element represents any mechanical, non optical device that is placed in the Light Path and spatially modifies the light beam. Examples include, Pin Holes, Phase Rings, Dark Field Stops, Hoffman Modulators, Diffraction Gratings etc. This element can be associated with an optional Optics Turret. | Substitutions: | DarkFieldStop, DiffractionGrating, FourierFilter, PhaseRing, PinHole, HoffmanPlate, IrisDiaphragm, VarelRing |  |  |  |  | Changed to Optical Aperture or Spatical Filter | Didn't see anywhere else to put things such as these. Develop this further? Who? |  |  | HAD PROPOSED OptoMechanicalDevice BUT THEY are not really MECHANICAL IN NATURE BECAUSE THEY DO NOT MOVE; PROPOSE TO CALL IT OPTICAL APERTURE; OPTICAL MASK; SPATIAL FILTER??
| 1 | Y | DiffractionGrating |  |  | A Diffraction Grating is an optical component with a periodic structure that splits and diffracts light into several beams travelling in different directions. | [6 empty] | ADDED
| 2 |  | DiffractionGrating | Period |  | This field records the spacing between the centers of two adjacent slits (d). | float with unit:PeriodUnit; [default nm] |  | R | [7 empty] | ADDED TO BE EVALUATED
| 2 |  | DiffractionGrating | SpatialFrequency |  | This field records the number of slits (or lines) per unit length (1/d). Typically it is measured in terms of number of slits per millimiter. | float with unit:none |  | R | [7 empty] | ADDED TO BE EVALUATED
| 2 |  | DiffractionGrating | Transmittive |  | This field records whether the periodic fine structure of this Diffraction Grating is Transmissive (True) and therefore refracts light, or is Reflective (False) and therefore reflects it. | boolean | [9 empty] | ADDED TO BE EVALUATED
| 2 |  | DiffractiveGrating | ModulationMode | ModulationModeList | This field indicates what aspect of the light beam is modulated by this device. | enum | Intensity, Phase, Polarization, Wavelength, Other | [8 empty] | ADDED TO BE EVALUATED
| 3 |  | DiffractionGrating | Ronchi |  | This field records whether this Diffraction Grating has a square wave profile, as defined by Ronchi (True) or a sine wave profile (False). | boolean | [9 empty] | ADDED TO BE EVALUATED
| 2 |  | MaskingPlate |  |  | This element describes an opaque Masking Plate positioned in the light path to block specific components of the light beam. | [6 empty] | ADDED
| 2 |  | MaskingPlate | Function | MaskingPlateFunction | This field describes the Function of this Spatial Mask. | enum | FourierFilter, PhaseMask, Other | R | [7 empty] | ADDED TO BE EVALUATED
| 2 |  | MaskingPlate | Excitation |  | This field records whether this Spatial Mask is placed in the Excitation (True) or the Emission (False) light path. | boolean |  | R | [7 empty] | ADDED TO BE EVALUATED
| 2 |  | MaskingPlate | ApertureSize |  | This field captures the maximum physical Width of individual Apertures present in this Spatial Mask. | float with unit:ApertureSizeUnit |  | R | [7 empty] | ADDED TO BE EVALUATED
| 2 |  | MaskingPlate | ApertureNr. |  | This field captures the Number of individual apertures present in this Spatial Mask. | integer with unit:none |  | R | [7 empty] | ADDED TO BE EVALUATED
| 3 |  | MaskingPlate | ApertureShape |  | This field captures the Shape of individual apertures present in this Spatial Mask. | string | [9 empty] | ADDED TO BE EVALUATED
| 3 |  | MaskingPlate | Material |  | This field captures the Material with which this Spatial Mask is constructed. | string | [9 empty] | ADDED TO BE EVALUATED
| 1 | Y | Pinhole |  |  | A pin-like aperture positioned axially to the Light Path, which can act as a spatial filter for a beam of light to "clean up" the output of lasers and / or to suppress out-of-focus light in a confocal Microscope. Specifically, in Confocal Microscopy, such Pinholes are diaphragm typically situated in a plane conjugate with the object, and they restrict the area in the object plane that is illuminated and/or from which light is collected. |  |  | 0 … ∞ |  |  |  | DECIDED TO MOVE TO ADVANCED; modified Description to align with ISO 10934-2:2007(E) |  |  |  | Move to Confocal extension? Changed description
| 2 |  | Pinhole | MaxDiameter |  | This field records the Maximum physical Diameter of this Pinhole. | float with unit:DiameterUnit [default mm] |  | R
| 3 |  | Pinhole | Shape | PinholeShapeList | This field records the Shape of this Pinhole. | enum | Round, Square, Diamond, Slit
| 1 |  | Detector-APE |  | Manufacturer Specs | This describes a detector available for this Instrument. An Instrument may have more than one detector of different types, such as photomultiplier, or camera. The Detector ID can be used as a reference within the Channel element in the Image element. The values stored in Detector represent the fixed values, variable values modified during the Acquisition go in DetectorSettings. | Substitutions: | Camera, Photomultiplier | 1 … ∞ | [7 empty] | APD is listed below in Point Detector | Should the detector and paramters be required? I do not see the "R" in the cardinal column... |  | APD
| 1 |  | Camera-APE |  | Substitution : choose one child from the list below | This describes a camera device available for this Instrument. A camera is device that converts optical brightness into an electrical signal using a CCD(a semiconductor chip, which converts the intensity of light falling onto its surface at different points into a charge) and then reproduces an image of the specimen. An Instrument may have more than one camera of different types, such as EMCCD, intensified EMCCD, CMOS, or analog video. |  | CCD, Intensified CCD, AnalogVideo, CMOS
| 2 |  | WavefrontSensor |  |  | A two dimensional sensor for measuring the location of the wavefront, typically used in Adaptive Optics. | [6 empty] | ADDED
| 2 |  | WavefrontSensor | Type | WFSType | This field records the type of this Wavefront Sensor. | enum | Shack-Hartmann, Interferometer, Other
| 2 |  | WavefrontSensor | Accuracy |  | This field records the specified Wavefront Accuracy of this sensor at Peak Wavelength (expressed as a fraction of the Wavelength). | PercentFraction
| 3 |  | WavefrontSensor | Sensitivity |  | This field records the specified Wavefront Sensitivity of this sensor at Peak Wavelength (expressed as a fraction of the Wavelength). | PercentFraction
IMAGE ACQUISITION Settings
| 1 |  | Experiment |  |  | Describes the experiment during which this Image was acquired. The optional Description element may contain free text to further describe the experiment. |  |  | 1 | [10 empty] | We should indicate the essential info we would like to hear in this section, otherwise it is just a notes section
| 2 |  | MicrobeamManipulation |  |  | Defines a microbeam operation and the region of the image it was applied to. The LightSourceSettingsRef element is a reference to the LightSourceSettings that were used in this case, including a reference to the light source that was used for this micromanipulation experiment (e.g. a laser used for photobleaching). |  |  | 0 … ∞ | same | 0 … ∞ |  | MOVE TO ADVANCED | How about relegating this to an extension (e.g. the advanced one)? This is very advanced stuff and not typical usage. |  |  |  | SHOULD WE MOVE THIS TO ADVANCED
| 2 |  | MicrobeamManipulation | ID (MicrobeamManipulationID) |  | A Unique Identifier for this MicrobeamManipulation hardware component. | MicrobeamManipulationID |  | R | same | R
| 2 |  | MicrobeamManipulation | Type | MicrobeamManipulationType | This field defines the type of micromanipulation performed. | enum | FLIP, FRAP, inverse FRAP, OpticalTrapping, Photoablation, Photoactivation, Uncaging, Other | R | same
| 2 |  | MicrobeamManipulation | ExperimenterRef |  | This empty element refers to Experimenter that performed this experiment. | Extension of Reference |  | 0 ... ∞ | same |  | 1
| 2 |  | MicrobeamManipulation | ROIRef |  | This empty element refers to the predefined four dimensional ROI (Region of Interest) that was used for this micromanipulation experiment. | Extension of Reference |  | 1 … ∞ | same |  | 1 … ∞
| 2 |  | MicrobeamManipulation | Description |  | This elemnt holds a description for the Microbeam Manipulation experiment. | Element |  | 0 … 1 | same |  | 0 … 1
| 2 |  | MicrobeamManipulation | LightSourceSettings |  | This element holds a reference to the LightSource (i.e. the ID attribute) that was used in this case as well as the setting that were applied to in in this case. | Element |  | 0 … ∞ | same |  | 0 … ∞
| 2 |  | MicrobeamManipulation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
| 1 |  | Image |  |  | This element describes the actual image that was acquired and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. |  |  | 1
| 2 |  | Image | MicroBeamManipulationRef |  | This empty element refers to an optional micro beam manipulation that was performed in connection with the Acquisition of this Image (e.g. a microbeam operation, such as a laser-induced photobleaching). | Extension of Reference |  | 0 … ∞ | same | 0 … ∞
| 1 |  | Pixels (should merge with Image) |  |  | Pixels is going to be removed in the future, but it is still required. [...]. |  |  | 1 … ∞
| 1 |  | Channel |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. |  |  | 1 … ∞
| 1 |  | Plane |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. |  |  | 1 … ∞
| 1 |  | LightPath |  |  | A description of the Ligth Path. |  |  | 1
| 1 |  | LightPath | AcoustoOpticalDeviceSettings-APE |  |  | Element |  | 0 … ∞
| 1 |  | LightPath | ElectroOpticalDeviceSettings-APE |  | This element refers to the electro optical device used to acquire this image and it holds the settings that were applied to the electro optical modulator during this Acquisition. | Element |  | 0 … ∞
| 1 |  | LightPath | OpticalApertureSettings-APE |  |  | Element |  | 0 … ∞
| 1 | Y | TIRFSettings |  | Settings | Total Internal Reflection Fluorescence (TIRF) microscopy is a modality in which fluorescence is excited in a thin layer by an evanescent wave produced by total internal reflection. This element defines the TIRF conditions that were used for this Experiment. |  |  | 0 … ∞ |  |  |  | MOVED TO ADVANCED | Shouldn't there be a TIRFDevice in the Specs sheet?
| 1 | Y | TIRFSettings | HILO |  | This field capture whether (True) or not (False) TIRF illumination for this Experiment was set up at or below the critical angle (63˚; "true" TIRF) or with oblique illumination (i.e., Highly Inclined and Laminated Optical - HILO). | boolean
| 3 |  | TIRFSettings | TIRFAngle |  | This field captures the incidence Angle used in this case. | float with unit:TIRFAngleUnit
| 1 | Y | TIRFSettings | TIRFHardwareModuleRef |  | In case a specific TIRF Hardware Module was used for this TIRF experiment, this empty element refers to the TIRF Module that was used. | Extension of Reference |  | 0 ... ∞
| 2 |  | TIRFSettings | TIRFLightPath | LightPath_type | A description of the Ligth Path utilized for this TIRF experiment. | Element |  | 0 ... 1
| 1 |  | TIRFSettings | Description |  | A description for the TIRF Settings | Element |  | 0 ... 1
| 1 | Y | ConfocalScannerSettings |  | Settings | This element holds the Setting applied to a Confocal Scanner used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Confocal Scanner in question. |  |  | 0 … ∞
| 1 |  | ConfocalScannerSettings | ID (ConfocalScannerID) |  | This is the Unique Identifier of the device used in this case. | ConfocalScannerID |  | R
| 1 | Y | ConfocalScannerSettings | PresetPinholeSize |  | In case this Confocal Scanner (i.e., Spinning disk, Swept-field, or other Multi-point scanners) has several fixed Pinhole Sizes to choose from, this field records the Pinhole Size that was set during the Acquisition of this Channel. | float with unit:PinholeSizeUnits |  | R |  |  |  |  | spinning disk, swept-field, etc. scanners have disks or plates with fixed-size pinholes and with some you can select which disk/plate you want, thus listed in Settings
| 1 | Y | ConfocalScannerSettings | PinholeSpacing |  | Ifapplicable, this field records the Spacing between individual Pinholes that was set during the Acquisition of this Channel. | float with unit:PinholeSpacingUnit
| 1 | Y | ConfocalScannerSettings | ScanDirectionality | ConfocalScanMode | This field records the ScanDirectionality that was set for this device during Acquisition. In Unidirectional scan mode, data are only acquired while the laser beam moves in one direction. In Bidirectional scan mode, one line is scanned in one direction, and the next line in the opposite direction. In other words, the return run, or flyback, of the laser beam is used for recording data as well, thus increasing the scanning frequency when using bidirectional scan mode. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | enum | Bidirectional, Unidirectional | R
| 1 | Y | ConfocalScannerSettings | IntegrationNumber |  | Signal integration is generally utilized to improve the overall signal-to-noise ratio of the resulting Image. This field records the number of sequential Scanning Units (Line, Frame or Stack) that have been integrated for the Acquisition of this Channel. | PositiveInteger with unit:none |  | R
| 1 | Y | ConfocalScannerSettings | IntegrationUnit | ScanIntegrationMode | Signal integration is generally utilized to improve the overall signal-to-noise ratio of the resulting Image. This field refers to the nature of the Scanning Units used as the basis of signal integration. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | enum | Frame, Line, Stack, Other | R
| 1 | Y | ConfocalScannerSettings | IntegrationMethod | ScanIntegrationMethod | Signal integration is generally utilized to improve the overall signal-to-noise ratio of the resulting Image. This field defines the signal Integration Method utilized for the Acquisition of this Channel. | enum | Average, Sum, Other | R
| 1 | Y | ConfocalScannerSettings | EffectiveZoom |  | This describes the effective zoom factor, aka "Confocal Zoom" or "Scan Zoom", utilized for the acquisition of this Image (e.g., zoom = 1.3). | float with unit:none |  | R
| 1 | Y | ConfocalScannerSettings | ZoomAreaWidth |  | This field records the Width of the Zoom Area that was scanned during the Acquisition of this Channel, expressed in Pixels. | PositiveInteger with unit:none
| 1 | Y | ConfocalScannerSettings | ZoomAreaHeight |  | This field records the Height of the Zoom Area that was scanned during the Acquisition of this Channel, expressed in Pixels. | PositiveInteger with unit:none
| 1 |  | ConfocalScannerSettings | Descanned |  | This field indicates whether or not the Acquisition of this Image involved descanning. | boolean |  | R | [7 empty] | QUESTION AROSE: is this a property of the SCANNER or of the DETECTOR?
| 2 |  | ConfocalScannerSettings | ScanFormat | ConfocalScanFormat | This field defines the format of the scanned element utilized for this Acquisition. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | enum | Frame, Line, Stack, Other | R
| 2 |  | ConfocalScannerSettings | MultiChannelMode | ConfocalMultiChannelMode | This field defines whether the the the Aquisiton of Multiple Channels for this Image happened in a Parallel (i.e., multiple channels acquired at the same time), or Sequentail (i.e., one channel after the other) manner. | enum | Parallel, Sequential, Other | R |  |  |  | CHANGED | since Multi Tracking is mostly a Zeiss terminology (Leica calls it Parallel vs. Sequential) and the definition looks confusing, how about we use the more generic Leica Names and call this MultiChannelMode?
| 2 |  | ConfocalScannerSettings | ScanTime |  | This records the time in seconds required to acquire a given number of lines, frames or stacks including the time required for repositioning the scanning device to the starting point. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | float with unit:ScanTimeUnit |  | R
| 3 |  | ConfocalScannerSettings | ConfocalFieldNumber |  | This field recors the Field Number of the Confocal scan optic is expressed as the objective magnification times the maximum diameter of the scanned field; [mm], obtained without shifting the position of the specimen and without stitching single images into a total larger image. Because images generally have a square format, the diameter of the scanned field is equals to the square root of 2 times the length of X dimension of the scanned field. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | float with unit:none
| 3 |  | ConfocalScannerSettings | Rotation |  | This field records the Rotation angle of the scan region. | float | [5 empty] | ADDED | added since many scan heads have this ability
| 3 |  | ConfocalScannerSettings | ScanningFrequency |  | This field records the Scanning Frequency is defined at the ratio between the Number of scanning element (Lines, Frames or Stacks) acquired and ScanTime. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | PercentFraction
| 1 |  | AcoustoOpticalDeviceSettings-APE |  | Settings | This element holds the Setting applied to an Acousto-Optical Device used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Acousto-Optical Device in question.
| 1 |  | AcoustoOpticalDeviceSettings-APE | ID (AcoustoOpticalDeviceID) |  | This is the Unique Identifier of the device used in this case. | AcoustoOpticalDeviceID |  | R
| 2 |  | AcoustoOpticalDeviceSettings-APE | DiffractionEfficiency |  | This field records a measure of the performance efficiency of this device. It's a measure of how much optical power is diffracted into a designated direction compared to the power incident onto the diffractive element. | PercentFraction
| 3 |  | AcoustoOpticalDeviceSettings-APE | AcousticVelocity |  | This field records the acoustic wave transit time across the crystal used during this Acquisition. | float with unit:AcousticVelocityUnit
| 3 |  | AcoustoOpticalDeviceSettings-APE | AcousticFrequency |  | This field records the frequency of the acoustic wave used to alter the diffraction properties of the crystal during this Acquisition. Wavelength tuning is very rapid and limited only by the acoustic transit time across the crystal. | float with unit:AcousticFreuencyUnit
| 3 |  | AcoustoOpticalDeviceSettings-APE | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0 … ∞
| 3 |  | AcoustoOpticalDeviceSettings-APE | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0 … ∞ | [11 empty] | This element records the range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used.
| 1 |  | AcoustoOpticalDeviceSettings-APE | TransmissionWavelengthRangeSettings |  | This element records the Range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 3 |  | AcoustoOpticalDeviceSettings-APE | ReflectionWavelengthRangeSettings |  | This element records the Range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞ | [11 empty] | This element records the range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used.
| 1 |  | AcoustoOpticalTuneableFilterSettings |  |  | This element holds the Setting applied to an Acousto Optical Tuneable Filter (AOTF) used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the AOTF in question.
| 1 |  | AcoustoOpticalBeamSplitterSettings |  |  | This element holds the Setting applied to an Acousto Optical Beam Splitter (AOBS) used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the AOBS in question.
| 1 |  | AcoustoOpticalDeflectorSettings |  |  | This element holds the Setting applied to an Acousto Optical Deflector used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Acousto Optical Deflector in question.
| 1 |  | AcoustoOpticalModulatorSettings |  |  | This element holds the Setting applied to an Acousto Optical Modulator used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Acousto Optical Modulator in question.
| 1 |  | ElectroOpticalDeviceSettings-APE |  | Settings | This element holds the Setting applied to an Electro-Optical Device used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Electro-Optical Device in question.
| 1 |  | ElectroOpticalDeviceSettings-APE | ID (ElectroOpticalDeviceID) |  | This is the Unique Identifier of the device used in this case. |  |  | R
| 3 |  | ElectroOpticalDeviceSettings-APE | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0 … ∞
| 3 |  | ElectroOpticalDeviceSettings-APE | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0 … ∞
| 1 |  | ElectroOpticalDeviceSettings-APE | TransmissionWavelengthRangeSettings |  | This element records the Range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 3 |  | ElectroOpticalDeviceSettings-APE | ReflectionWavelengthRangeSettings |  | This element records the Range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 2 |  | DeformableMirrorSettings |  |  | This element holds the Setting applied to a Deformable Mirror used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Deformable Mirror in question.
| 2 |  | ElectroOpticalModulatorSettings |  |  | This element holds the Setting applied to an Electro-Optical Modulator used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Electro-Optical Modulator in question.
| 1 |  | LiquidCrystalTuneableFilterSettings |  |  | This element holds the Setting applied to a Liquid Crystal Tuneable Filter (LCTF) used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the LCTF in question.
| 2 |  | SpatialLightModulatorSettings |  |  | This element holds the Setting applied to a Spatial Light Modulator used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Spatial Light Modulator in question.
| 1 | Y | OpticalApertureSettings-APE |  | Settings | This element holds the Setting applied to an Optical Aperture used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Optical Aperture in question.
| 1 |  | OpticalApertureSettings-APE | ID (OpticalApertureID) |  | This is the Unique Identifier of the device used in this case. |  |  | R
| 3 |  | OpticalApertureSettings-APE | TransmittanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Transmittance profile of this optical component. | Element |  | 0 … ∞
| 3 |  | OpticalApertureSettings-APE | ReflectanceProfileFile | FileAnnotation | This element represents a file in .csv or other suitable formats containing the Reflectance profile of this optical component. | Element |  | 0 … ∞
| 1 |  | OpticalApertureSettings-APE | TransmissionWavelengthRangeSettings |  | This element records the Range of Transmission Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 3 |  | OpticalApertureSettings-APE | ReflectionWavelengthRangeSettings |  | This element records the Range of Reflection Wavelengths that were set for this optical component during the Acquisition of this Image. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | Element |  | 0 … ∞
| 1 | Y | DiffractiveGratingSettings |  |  | This element holds the Setting applied to a Diffraction Grating used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Diffraction Grating in question.
| 2 |  | DiffractiveGratingSettings | DiffractionEfficiency |  | This is a measure of the performance efficiency of this device. It's a measure of how much optical power is diffracted into a designated direction compared to the power incident onto the diffractive element. | PercentFraction
| 2 |  | MaskingPlateSettings |  |  | This element holds the Setting applied to a Masking Plate used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Masking Plate in question.
| 1 | Y | PinHoleSettings |  |  | This element holds the Setting applied to a Pinhole used during the Acquisiton of this Channel as well as a reference (the ID attribute) to the Pinhole in question.
| 1 | Y | PinHoleSettings | Aperture |  | This refers to the Diameter of the Pin Hole's Aperture that was set for this Acquisition, expressed in Airy units. | float with unit:none |  | R
Note on A1: ====== ID#AAAAJ7OOiPU (2019-12-18 22:28:54) This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on D1: ====== ID#AAAAJ7OOiPY Caterina Strambio De Castillia (2019-12-19 19:05:02) This essentially means that this Element is not actually used per se. It is defined as the Ancestor of Elements that are actually created when representing Hardware components. The advantage is this way all Children elements will inherit all Attributes of this Abstract ancestor and one avoids to have to repeat the same Attributes multiple times. It is a modeling trick that avoids data duplication.
Note on B3: ====== ID#AAAAJ7OOiPc (2019-12-18 22:28:54) This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.

## 4DN-BINA Calibrat Ext.
APE = Abstract Parent Element
[9 empty] | BINA WORKING GROUP | [10 empty] | QUAREP comments
| Tier | Element | Attribute or Subelement | Complex type | Description | Data type | Allowed values | Cardinality | DECISION | Caterina additions | Alex C. | Alex L | Claire | Glyn | Judith | Orestis | Peter | Roland | RESPONSE | WG7 Member 1 [Enter your name and Choose a color] | WG7 Member 2 [Enter your name and Choose a color] | Marco | WG7 Member n [Enter your name and Choose a color]
Microscope Speecifications | 1 | Instrument |  |  | This element describes the Instrument used to capture the Image. It is primarily a container referring to the Microscope's components list, which include the Microscope Stand, Microscope Table, Light Source, Light SourceCoupling, SampleHolder, Focusing devices, Objective, Optovar, Filter, Dichroic, AdditionalOptics, Detector and CalibrationTool. The Filter element can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. Within the Image itself, a reference is made to this one Filter element. There may be multiple Light Sources, detectors, Objectives and filters etc. in an Instrument. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the Ligth Path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. |  |  | 1
| 2 | IntensityCalibrationTool |  | ManufacturerSpecs | A tool, such as the Meta-Max calibration toold developed by the 4D Nucleome Imaging Standards Working Group, used to perform an intensity calibration procedure |  |  | 0 … ∞ | [5 empty] | can't think of anything else needed for this
| 2 | IntensityCalibrationTool | ID |  | A Unique Identifier for this component. |  |  | R
| 2 | IntensityCalibrationTool | Name |  | A User-defined Name for this component. | Denomination |  | R
| 2 | IntensityCalibrationTool | Type |  | The type of calibration proceudure this tool is intended to perform | enum | Detector, Excitation, Other | R
| 2 | IntensityCalibrationTool | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 … ∞
IMAGE ACQUISITION Settings
| 1 | Image |  |  | This element describes the actual image that was acquired and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. |  |  | 1
| 2 | OpticalCalibration |  |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | Element |  | 0 … ∞
| 2 | DetectorCalibration |  |  | This procedure measures the performance of the dectector | Element |  | 0 ... ∞
| 1 | Pixels (should merge with Image) |  |  | Pixels is going to be removed in the future, but it is still required. [...]. |  |  | 1 … ∞
| 1 | Channel |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. |  |  | 1 … ∞
| 2 | Channel | ExcitationCalibration |  | This describes the procedure that was used for excitation calibration and the resulting calibration measurements. | Element |  | 0 … ∞
| 2 | Channel | FieldCalibration |  | This describes the procedure used to measure the field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | Element |  | 0 … ∞
| 1 | Plane |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. |  |  | 1 … ∞
| 2 | MechanicalCalibration |  |  | This describes the procedures that were used for mechanical calibration and the resulting calibration results. |  |  | 0 … ∞
[14 empty] | I think somewhere there should be some lens infomration- esp. NA- pixel size and wavelength are in the grey image boxes |  |  |  |  | This is already available in other parts of the Model
| 2 | OpticalCalibration |  |  | This describes the procedure that was used for Optical Calibration and the resulting Calibration measurements. |  |  | 0 … ∞ |  | Alternative Name: PSFCalibration? |  |  |  | Add Refractive Index for immersion type (and sample?) | I second Glen. Also, some lens have collars for setting immersion type: eg oil, water, glycerol. That info should be recorded. In addition the theoritical values are not the same for different RI. |  |  |  | This is already available in other parts of the Model
| 2 | OpticalCalibration | ID |  | A Unique Identifier for this component. | LSID |  | R | [5 empty] | would this be equivalent to a lens ID? or a Unique Identifier for this calibration dataset? |  |  |  |  | This is a unique ID for this specific Optical Calibration event
| 2 | OpticalCalibration | Date |  | Ths field records the Date in which this Calibration procedure was performed. | date |  | R
| 2 | OpticalCalibration | TheoreticalXYSpotSize_FWHM |  | This field captures the Theoretical (expected) lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using a Gaussian function, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. This value should be calculated in a manner as to be compared to the Median XY Spot Size_FWHM. | float with unit:FWHM_MinTheoretical |  |  | CHANGED NAME AND DEFINITION
| 2 | OpticalCalibration | MedianXYSpotSize_FWHM |  | An important way of evaluating the effective resolution of any Microscope, consists in measuring the size and shape of the observed Point Spread Function (PSF) generated by imaging one or more fluorescent Point-like objects (i.e., Point Emitters) under conditions that are identical to those used to Acquire a given experimental Image. More specifically and according to ISO 21073:2019, lateral Resolution is defined by the FWHM of the intensity signal along a lateral direction through the centre of a fluorescent Point-like object (i.e., Spot or Point Emitter) positioned at the center of the Field of View (FOV). This field captures the average Observed Median lateral (i.e., in the xy-plane) Spot Sze expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using an Elliptic Gaussian function. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:FWHM_MedianSpotSize |  |  | ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. | ADD STATEMENT IN THE DESCRIPTION TO TAKE THIS MEASUREMENT IN THE CENTER but ideally it should be explored across the FOV. | Just to confirm that you are fitting to a Gaussian function? Using a single pixel value to measure the maximum can be dangerous in noisy data. |  |  | do we need to define these values are from the central area of the lens?
| 2 | OpticalCalibration | MinXYSpotSize_FWHM |  | An important way of evaluating the effective resolution of any Microscope, consists in measuring the size and shape of the observed Point Spread Function (PSF) generated by imaging one or more fluorescent Point-like objects (i.e., Point Emitters) under conditions that are identical to those used to Acquire a given experimental Image. More specifically and according to ISO 21073:2019, lateral Resolution is defined by the FWHM of the intensity signal along a lateral direction through the centre of a fluorescent Point-like object (i.e., Spot or Point Emitter) positioned at the center of the Field of View (FOV). This field captures the average Observed Minimum lateral (i.e., in the xy-plane) Spot Sze expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using an Elliptic Gaussian function. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:FWHM_MinSpotSize |  |  | ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. | ADD STATEMENT IN THE DESCRIPTION TO TAKE THIS MEASUREMENT IN THE CENTER but ideally it should be explored across the FOV. | Do you extract one width from each point or two? Most points are slightly elliptical in XY |  |  | do we need an n number for the no. of beads/spots sampled? | Regarding Glyn's comment, how about mean and s.d.? |  |  |  | This is why we ask for both Min and Max
| 2 | OpticalCalibration | MaxXYSpotSize_FWHM |  | An important way of evaluating the effective resolution of any Microscope, consists in measuring the size and shape of the observed Point Spread Function (PSF) generated by imaging one or more fluorescent Point-like objects (i.e., Point Emitters) under conditions that are identical to those used to Acquire a given experimental Image. More specifically and according to ISO 21073:2019, lateral Resolution is defined by the FWHM of the intensity signal along a lateral direction through the centre of a fluorescent Point-like object (i.e., Spot or Point Emitter) positioned at the center of the Field of View (FOV). This field captures the average Observed Maximum lateral (i.e., in the xy-plane) Spot Sze expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using an Elliptic Gaussian function. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:FWHM_MaxSpotSize |  |  | ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. | ADD STATEMENT IN THE DESCRIPTION TO TAKE THIS MEASUREMENT IN THE CENTER but ideally it should be explored across the FOV.
| 2 | OpticalCalibration | TheoreticaZSpotSize_FWHM |  | This field captures the Theoretical (expected) z-axial spot size expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using a Gaussian function, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the z-plane. This value should be used calculated in a manner as to be compared to the median Z spot size. | float with unit:FWHM_Ztheoretical |  |  | CHANGED NAME AND DEFINITION; ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. | ADD STATEMENT IN THE DESCRIPTION TO TAKE THIS MEASUREMENT IN THE CENTER but ideally it should be explored across the FOV. | Again, this is a fairly complex calculation and can vary considerably with Z in the presence of even a modest index mismatch
| 2 | OpticalCalibration | MedianZSpotSize_FWHM |  | An important way of evaluating the effective resolution of any Microscope, consists in measuring the size and shape of the observed Point Spread Function (PSF) generated by imaging one or more fluorescent Point-like objects (i.e., Point Emitters) under conditions that are identical to those used to Acquire a given experimental Image. More specifically and according to ISO 21073:2019, axial Resolution is defined by the FWHM of the intensity signal along the axial direction through the centre of a fluorescent Point-like object (i.e., Spot or Point Emitter) positioned at the center of the Field of View (FOV). This field captures the average Observed Median Z-axial Spot Sze expressed as the Full Width at Half Maximum (FWHM) of the Point Spread Function (PSF) intensity profile fitted using an Elliptic Gaussian function. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:FWHM_ZSpotSize |  |  | CHANGED NAME; ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. | ADD STATEMENT IN THE DESCRIPTION TO TAKE THIS MEASUREMENT IN THE CENTER but ideally it should be explored across the FOV.
| 2 | OpticalCalibration | CheckedNA |  | In case the Numerical Aperture (NA) of the Objective utilized to perform this Optical Calibration is variable (i.e., as for example in case of Objectives that can be used with or without Cover Glass or Objectives that can be "dipped" in different media), this field capture whether (True) or not (False) a manual check was performed to confirm that the Numerical Aperture was set to the Maximum allowable value. This check is necessary because often NA regulation in these types of Objectives occurs via a rotating collar whose exact position is difficult to measure and reproduce across different Experiments. In this context, setting the collar to the Maximum NA position is desirable not only because it is more easy to reproduce and it best to estimate the maximal resolution of the system. | boolean |  |  | ADDED; MODIFIED DESCRIPTION TO CLARIFY WHAT SETTING SHOULD BE CHOSEN | How do we know which is the correct value? Is this the max or min or something else? Sorry I think Judith said this but I missed it.
| 2 | OpticalCalibration | Z-Dimension |  | This field records whether (True) or not (False) this Optical Calibration procedure was performed using XY or XYZ Images of the Optical Standard sample. | boolean |  |  | ADDED
| 2 | OpticalCalibration | NyquistRate |  | This field captures the ratio between the Pixel Size and the Nyquist Criterion. The Nyquist-Shannon sampling theorem establishes that "when sampling a signal (e.g., converting from an analog signal to digital), the sampling frequency must be greater than twice the Band Width of the input signal in order to be able to reconstruct the original perfectly from the sampled version" (Whittaker E.T., (1915). Proc. Roy. Soc. Edinburgh Sect. A, 25:181; Shannon C.E., (1949) Proc. IRE, 37:10). The Nyquist Criterion determines the minimal Sampling Density needed to capture ALL information from the Specimen into the Image. It can be shown that if the Sampling Density is higher than the Nyquist Rate (also referred to as Critical Sampling Distance) all information about the object is captured. It is important to note that the Nyquist Rate depends on the Objective Numerical Aperture (NA) and on the Excitation and Emission WaveLengths. The recommended Nyquist Rate is > 2 for conventional light microscopy, and it is 2.5 - 3 for high resolution imaging. | float with unit:none |  |  | ADD
| 2 | OpticalCalibration | SampleSize |  | This field captures the Sample size that was used to calculate the reported Optical Calibration metrics. For example, this might represent the number of Colored Beads that were utilized to gather data on the Optical performance of the Microscope. | integer with unit:none |  |  | ADD
| 2 | OpticalCalibration | PSFEquation |  | This field captures the Equation that was used to perform the Eliptical PSF Gaussian Fit. | string |  |  | ADD
| 3 | OpticalCalibration | Planarity |  | This is a measure of the flatness of the focal plane across the Field of View (FOV). The planarity of the focal plane is calculated from the observed z-position of the best focus of each Point Emitter. It indicates focal aberrations or tilt. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:PlanarityUnit |  |  | STET; ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV. |  | For most modern objectives, field curvature is typically a few 10s of nm at the edge of the FOV. I do not know of any tests sensitive enough to measure this.
| 3 | OpticalCalibration | Asymmetry |  | This field reports the mean +/- standard deviation over the Field of View (FOV) of the observed Point Spread Function (PSF) Asymmetry (i.e. the ratio of the minimum lateral FWMH over the maximum lateral FWHM). It indicates the presence of optical aberrations. Ideally this measurement should be reflective of multiple individual Point Emitter observations across the Field of View (FOV). However, if a single measurement is taken, it is essential that this be taken from the center of the Objective FOV. | float with unit:none |  |  | THIS MIGHT NEED TO BE REMOVED DEPENDING ON QUAREP WG review; for now STET; ADDED more details about how and where to take the measurements. The measurement has to be TAKEN AT THE CENTERor the FOV but ideally it should be explored across the FOV.
| 2 | OpticalCalibration | OpticalCalibrationImageRef |  | This is an empty element that refers to an Image (typically a z-stack) obtained by observing the Optical Standard used in this instance and Acquired using the same Instrument and Settings employed to acquire the experimental Image this calibration refers to. | Extension of Reference |  | 1 ...∞
| 2 | OpticalCalibration | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞ | [6 empty] | Some lens have variable NA (eg Leica), there should be some way to have the actual NA used confirmed and recorded.
| 2 | OpticalCalibration | PSFEquationFile | FileAnnotation | This element refers to any type of file describing the PSF Equation that was used in this case. | Element |  | 0 ...∞ | ADD
| 2 | OpticalCalibration | OpticalCalibrationReport | FileAnnotation | This element refers to any File containing a report generated by the Calibration Software referred to in Calibration procedure. | Element |  | 0 ...∞ | CHANGED
| 2 | OpticalCalibration | ChromaticRegistrationEvaluation |  | According to ISO 21073:2019, co-registration accuracy measures the capacity of a microscope system to image a fluorescent object with different excitation and emission Wavelengths at the same position in the Image. Hence, the spatial distance of the Image positions of a Point-like object (i.e., Spot or Point Emitter) located in the centre of the Field of View (FOV) for different excitation and detection Wavelength bands is denoted as co-registration accuracy. The co-registration accuracy shall be calculated from the Intensity maxima or the Intensity centre of gravity of the Image of the Point-like object for the different Wavelengths. The lateral co-registration accuracy is expressed as: ∆lateral = √(xλ1 - xλ2)^2 + (yλ1 - yλ2)^2; and the axial co-registration accuracy is expressed as: ∆axial = √(zλ1 - zλ2)^2. This element refers to the results of a specific Co-registration Accuracy evaluation assay associated with this Optical Calibration procedure. | Element |  | 0 ...∞ | ADDED
| 2 | OpticalCalibration | FieldUniformityEvaluation |  | This describes the procedure that was used to measure the Field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | Element |  | 0 ...∞
| 2 | OpticalCalibration | OpticalStandard-APE |  | This element refers to the Optical Calibration Standard used in this case. This might be a pre-manufactured slide, such as those produced by Chroma (https://www.chroma.com/products/accessories/92001-autofluorescent-plastic-slides), Argolight (http://argolight.com/), Geller MicroAnalytical Laboratory (Magnification Reference Standard MRS-4 and others; http://www.gellermicro.com/mag_standards/mrs.html), Zeiss and others. Alternatively, the standard might be a slide containing multi-colored beads of known size and emission characteristics. Finally, slides coated with specific DNA origami patterns are increasingly being used for this purpose as well. | Element |  | 1 ...∞
| 2 | OpticalCalibration | CalibrationSoftware | Software | The Name of the software tool utilized to generate this Calibration report. As an example, such a Software tool might analyzes microscopy images of fluorescent beads in order to determine quantitatively the resolution and chromatic aberrations of a fluorescence microsocope over the full field of view. An example of such a tool is PSFj (http://www.knoplab.de/psfj/sample-page/). Alternatively, a Software tool might be used to extract Intensity Calibration results obtained with a specific Intensity Calibration Tool, such as a Power Meter or MetaMax (https://www.umassmed.edu/grunwaldlab/). | Element |  | 0 ...∞ | ADDED
| 2 | ChromaticRegistrationEvaluation |  |  | According to ISO 21073:2019, co-registration accuracy measures the capacity of a microscope system to image a fluorescent object with different excitation and emission Wavelengths at the same position in the Image. Hence, the spatial distance of the Image positions of a Point-like object (i.e., Spot or Point Emitter) located in the centre of the Field of View (FOV) for different excitation and detection Wavelength bands is denoted as co-registration accuracy. The co-registration accuracy shall be calculated from the Intensity maxima or the Intensity centre of gravity of the Image of the Point-like object for the different Wavelengths. The lateral co-registration accuracy is expressed as: ∆lateral = √(xλ1 - xλ2)^2 + (yλ1 - yλ2)^2; and the axial co-registration accuracy is expressed as: ∆axial = √(zλ1 - zλ2)^2. This element refers to the results of a specific Co-registration Accuracy evaluation assay associated with this Optical Calibration procedure. |  |  | 0 ...∞ | ADDED
| 2 | ChromaticRegistrationEvaluation | ID (Chromatic AberrationID) |  | A Unique Identifier for this component. | LSID |  | R | ADDED
| 2 | ChromaticRegistrationEvaluation | ReferenceWavelength |  | This field captures the Wavelength of the Reference Channel against which the Chromatic Registration Accuracy is being calculated in this procedure. | float with unit:Wavelength_1Unit |  | R | ADDED
| 2 | ChromaticRegistrationEvaluation | TestWavelength |  | This field captures the Wavelength of the Test Channel whose Chromatic Registration Accuracy is being calculated in this procedure. | float with unit:Wavelength_2Unit |  | R | ADDED
| 2 | ChromaticRegistrationEvaluation | XChromaticShift_RT |  | This field captures the offset in the the observed X-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Reference vs. the Test Wavelengths. This shall be calculates as: √(xλT - xλR)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately. | Slight magnification differences between colour channels will produce equal and opposite offsets either side of the optic axis, the average will be zero. RMS offset might be better. |  |  | also need wavelengths compared- do we fix this in a similar way to lens ISO data? eg green v red?
| 2 | ChromaticRegistrationEvaluation | YChromaticShift_RT |  | This field captures the offset in the the observed Y-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Reference vs. the Test Wavelengths. This shall be calculates as: √(xλT - xλR)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately.
| 2 | ChromaticRegistrationEvaluation | XChromaticShift_TR |  | This field captures the offset in the the observed X-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Test vs. the Reference Wavelengths. This shall be calculates as: √(xλR - xλT)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately.
| 2 | ChromaticRegistrationEvaluation | YChromaticShift_TR |  | This field captures the offset in the the observed Y-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Test vs. the Reference Wavelengths. This shall be calculates as: √(xλR - xλT)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately.
| 2 | ChromaticRegistrationEvaluation | AxialChromaticShift_RT |  | This field captures the offset in the the observed Z-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Reference vs. the Test Wavelengths. This shall be calculates as: √(zλT - zλR)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately.
| 2 | ChromaticRegistrationEvaluation | AxialChromaticShift_TR |  | This field captures the offset in the the observed Z-coordinate of the Intensity maxima or the Intensity centre of gravity of the Image of a Point-like object (i.e., Spot or Point Emitter) when comparing the Test vs. the Reference Wavelengths. This shall be calculates as: √(zλT - zλR)^2, where This offset could be reported either in number of Pixels or in physical Distance, when a sub-Pixel offset needs to be calculated. | float with unit:ChromaticShiftUnit |  |  | ADDED | ADD Chromatic Shift in both directions Test vs. Reference and Reference vs. Test. Use Either Pixels or Lenght. Enter both X and Y separately.
START FROM HERE | 2 | FieldUniformityEvaluation |  |  | This describes the procedure that was used to measure the Field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. |  |  | 0 … ∞
| 2 | FieldUniformityEvaluation | ID |  | A Unique Identifier for this component. |  |  | R | Added the coordinates of the max intensity pixel in the Flat Field Image, the average +/-SD of the ratios obtained by comparing the max pixel intensity and the intensities of the four corner pixels
| 2 | FieldUniformityEvaluation | MaxIntensityPixelXCoordinate |  | This field reports the x-axis coordinate of the Pixel that has the Maximum Intensity in the Image used for this Flat Field Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | FieldUniformityEvaluation | MaxIntensityPiexlYCoordinate |  | This field reports the y-axis coordinate of the Pixel that has the Maximum Intensity in the Image used for this Flat Field Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | FieldUniformityEvaluation | MeanMaxtoCornerIntensityRatio |  | This field reports the Mean between the Ratios obtained by comparing the Maximum Pixel Intensity and the Intensities of the four corner Pixels of the Image used for this Flat Field Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | FieldUniformityEvaluation | SDMaxtoCornerIntensityRatio |  | This field reports the Standard Deviation between the Ratios obtained by comparing the Maximum Pixel Intensity and the Intensities of the four corner Pixels of the Image used for this Flat Field Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | FieldUniformityEvaluation | FlatFieldImageRef |  | This is an empty element that refers to an Image obtained from a Standard Slide designed to eavluate the Field of View (FOV) dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the Optical Path. | Extension of Reference |  | 1 … ∞
build this to work with an omero database?
| 2 | OpticalCalibrationStandard-APE |  | ManufactuerSpecs | This element refers to the Optical Standard used in this case. This might be a pre-manufactured slide, such as those produced by Argolight (http://argolight.com/), Geller MicroAnalytical Laboratory (http://www.gellermicro.com/mag_standards/mrs.html), PSFCheck (https://www.psfcheck.com/) Zeiss and others. Alternatively, the standard might be a slide containing multi-colored beads of known size and emission characteristics. Finally, slides coated with specific DNA origami patterns are increasingly being used for this purpose as well. |  |  | 1 … ∞
| 2 | OpticalCalibrationStandard-APE | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 | OpticalCalibrationStandard-APE | Name |  | A User-defined Name for this component. | Denomination |  | R
| 2 | OpticalCalibrationStandard-APE | Method |  | This field records a Description of Method that was employed to prepare this Optical Calibration Standard Sample. | string |  |  | ADDED
| 2 | OpticalCalibrationStandard-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞
| 2 | OpticalCalibrationStandard-APE | WavelengthRange |  | This element records a range of Wavelengths this component is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞ | ADDED
| 2 | FluorescentReferenceSlide |  |  | This element a Fluorescent Reference Slide. These Slides provide a uniform fluorescent field and therefore they enable to check how evenly a Excitation Light Source is filling the Field of View (FOV), and whether it is centered. These Slides are easy to use an cheap. However they are typically made of plastic and it is often difficult to ensure the reproducibility of their emission properties. These Slides are mostly known to be produced by Chroma, but similar products are available from Ted Pella or Thorlabs. |  |  |  | ADDED | ADD GENERALIZATION OF CHROMA SLIDE; Are there others?
| 2 | ColoredBeadsSlide |  |  | This element describes a Slide coated with Multicolor Fluorescent Microspheres (also known as Multicolored Beads) used for assessing the Optical properties of the Microscope that was used to acquire this experimental Image. These Slides can be used to measure Optical properties such as resolution, chromatic registration, and different kinds of optical distorsions. | [8 empty] | also need date made for a homemade bead slide? or do we avoid this due to RI differences when home-made? | Considering there is no consensus in the sample used (ie beads, slide, DNAOrigami, etc), perhaps there should be room for ample description of the sample used? Date of preparation, Manufacturer and lots number for coverlips, beads, mounting media RI, Argo-slide etc,
| 2 | ColoredBeadsSlide | Date |  | This field captures the Date in which this Colored Beads Slide was prepared. |  |  | R
| 2 | ColoredBeadsSlide | BeadVolume |  | This field captures the Volume of the Multi Color Microspheres preparation that was utilized to prepare this Slide. | float with unit:BeadVolumeUnit
| 2 | ColoredBeadsSlide | MultiColorMicrosphere |  | This element describes a Multicolor Fluorescent Beads preparation used to facilitate the optical calibration of conventional fluorescence microscopes, confocal laser scanning microscopes, and associated image-processing equipments. These Beads are glass or plastic microspheres stained with one or more fluorescent dyes. | Element |  | 1 … ∞
| 2 | CalibrationStandardSlide |  |  | This element describes a Standard Calibration Slide (also referred to as Standard Test Slide) used for assessing the Optical properties of the Microscope that was used to acquire this experimental Image. These Slides can be used to measure multiple different Optical properties including resolution, chromatic registration, magnification, distance, and different kinds of optical distorsions. | [8 empty] | Add pattern imaged?
| 2 | CalibrationStandardSlide | Type | CalibrationSlideType | The type of calibration slide used in this case. Examples include slides produced by Argolight, Geller MicroAnalytical Laboratory, PSFCheck, Zeiss and others. | enum | ARGO-HM, ARGO-LM, ARGO-SIM, Geller-MRS4, Zeiss universal D=0, PSFCheck, Other | R | ADDED |  |  |  |  | incl. PSFCheck
| 2 | CalibrationStandardSlide | CalibrationMeasure | CalibratedMeasure | This field captures the Optical characteristic this Slide is intended to calibrate, examples include Co-registration Accuracy, Detector linearity, Distance, Magnification, Optical distorsion, Resolution, or Multi. | enum | Co-registration Accuracy, Detector linearity, Distance, Magnification, Optical distorsion, Resolution, Other | R
| 2 | CalibrationStandardSlide | ImagedArea |  | This field describes the specific region or sets or regions of the Calibration Slide that have been used for this Optical Calibration procedure. | string |  |  | ADDED
| 2 | CalibrationStandardSlide | PatternImageRef |  | This is an empty element that refers to an Image depcting the pattern etched on this Calibration Slide. | Extension of Reference |  |  | ADDED
| 2 | DNAOrigami |  |  | This element describes an Standard Test Slide constructed using the DNA Origami technique.
| 2 | DNAOrigami | Size |  | This field capture the Full Width at Half Maximum (FWHM) of the fluorescence emitters associated with this DNA Origami-based microscopy standard slide. | float with unit:SizeUnit |  | R | ADDED
| 2 | DNAOrigami | Type |  | This field captures the type of fluorescence emitter(s) associated with this DNA Origami-based Optical Calibration Standard used in this case. Options include Point-Source and Nanoruler. | enum | Point-Source, Nanoruler, Other | R | ADDED |  |  |  |  | Add FWHM size?
| 2 | DNAOrigami | Description |  | A free text multi-line optional description for this component. | Element |  | 0 … 1 | ADDED
| 2 | MultiColorMicrospheres |  | ManufacturerSpecs | This element describes a preparation of Multicolor Fluorescent Microspheres (also often referred to as Beads) used to facilitate the optical calibration of fluorescence microscopes, and associated image-processing equipments. These Microspheres are usually made of glass or plastic and they stained with one or more fluorescent dyes. |  |  |  | ADDED BEADS as a separate ELEMENT; add Slide preparation description (multi-line comment),
| 2 | MultiColorMicrospheres | ID |  | A Unique Identifier for this component. | LSID |  | R
| 2 | MultiColorMicrospheres | Diameter |  | This field captures the Diameter of these Microspheres. | float with unit:BeadSizeUnit |  | R |  | ADD BEADS as separate ELEMENT. Refer to Coverslip and SampleHolder; add Slide preparation description (multi-line comment),
| 2 | MultiColorMicrospheres | Concentration |  | This field captures the concentration of these Microspheres. | float with unit:BeadConcentrationUnit |  | R
| 2 | MultiColorMicrospheres | Type | BeadType | This field captures the type of these Multicolor Microspheres. | enum | Single color, Multi-color, Other
| 2 | MultiColorMicrospheres | Material |  | This field capture the material these Microspheres are made of. | string
| 2 | MultiColorMicrospheres | WavelengthRange |  | This element records a range of Wavelengths this component is designed for, according to the Manufacturer. Multiple of this element may be used. | Element |  | 0 ... ∞
| 2 | IntensityCalibration-APE |  |  | This describes the procedure(s) that were used for excitation and detector calibration and the resulting calibration results. |  |  | 0 … ∞
| 2 | IntenstiyCalibration-APE | ID |  | A Unique Identifier for this component. |  |  | R
| 2 | IntenstiyCalibration-APE | Date |  | Ths field records the Date in which this Calibration procedure was performed. |  |  | R
| 2 | IntenstiyCalibration-APE | Method |  | This field records a Description of Method used for this calibration procedure. | string |  | R | [6 empty] | What are these methods?
| 2 | IntenstiyCalibration-APE | IntensityCalibrationToolRef |  | This element is a reference to the tool, if any, used to perform this calibratin procdure | Extension of Reference |  | 0 … ∞
| 2 | IntenstiyCalibration-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞
| 2 | IntenstiyCalibration-APE | IntensityCalibrationReport | FileAnnotation | This element refers to any File containing a report generated by the Calibration Software referred to in Calibration procedure. | Element |  | 0 ...∞ | ADDED
| 2 | IntenstiyCalibration-APE | CalibrationSoftware | Software | The Name of the software tool utilized to generate the OpticalCalibration report. As an example, such a tool might analyzes microscopy images of fluorescent beads in order to determine quantitatively the resolution and chromatic aberrations of a fluorescence microsocope over the full field of view. An example of such a tool is PSFj (http://www.knoplab.de/psfj/sample-page/). | Element |  | 0 ...∞ | ADDED
| 2 | ExcitationCalibration |  |  | This element captures a procedure used to measure and calibrate the Wavelength and Irradiance of the Excitation Light Source used to Acquire this experimental Image. | [6 empty] | Should the type of detector be described? | [6 empty] | This is already available in other parts of the Model
| 2 | ExcitationCalibration | ObservedEx-Wavelength |  | This field records the observed Peak Excitation Wavelength at the Back Objective as measured by using this Calibration Procedure. | float with unit:ObservedExcitationWavelengthAtObjectiveUnit |  | R | Changed |  |  |  |  | Also needs the power settings value set in software(eg % laser for AOTF)
| 2 | ExcitationCalibration | ObservedPowerAtObjective |  | This field records the observed Excitation Light Power at the Back Objective as measured by using this Calibration Procedure. | float with unit:ObservedPowerAtObjectiveUnit |  | R | Changed
| 2 | ExcitationCalibration | ObseredPowerAtSample |  | This field records the observed Excitation Light Power at the Sample (i.e., Front of the Objective) as measured by using this Calibration Procedure. | float with unit:ObservedPowerAtSampleUnit |  | R | ADDED
| 2 | ExcitationCalibration | ShortTermIlluminationStability |  | According to ISO 21073:2019, Illumination Stability is defined as the percentage deviation of Minimum and Maximum Illumination Powers from the Average Power during the time interval considered. Specifically, SATB [%] = 100x {1-[(Pmax-Pmin)/Pmax+Pmin]}, where Pmax is the maximum laser power recorded during the time interval; and Pmin is the minimum laser power recorded during the time interval. This field records the Short Term Illumination Stability recorded during this Calibration Procedure by measuring Illumination Power every second for a time interval of 5 min of continuous illumination. | float with unit:none |  |  | ADDED | 5 min
| 2 | ExcitationCalibration | MediumTermIlluminationStability |  | According to ISO 21073:2019, Illumination Stability is defined as the percentage deviation of Minimum and Maximum Illumination Powers from the Average Power during the time interval considered. Specifically, SATB [%] = 100x {1-[(Pmax-Pmin)/Pmax+Pmin]}, where Pmax is the maximum laser power recorded during the time interval; and Pmin is the minimum laser power recorded during the time interval. This field records the Medium Term Illumination Stability recorded during this Calibration Procedure by measuring Illumination Power every 30 s for one second for a time interval of 120 min of continuous illumination. | float with unit:none |  |  | ADDED | 2 hours
| 2 | ExcitationCalibration | LongTermIlluminationStability |  | According to ISO 21073:2019, Illumination Stability is defined as the percentage deviation of Minimum and Maximum Illumination Powers from the Average Power during the time interval considered. Specifically, SATB [%] = 100x {1-[(Pmax-Pmin)/Pmax+Pmin]}, where Pmax is the maximum laser power recorded during the time interval; and Pmin is the minimum laser power recorded during the time interval. This field records the Long Term Illumination Stability recorded during this Calibration Procedure by measuring Illumination Power for one second for a time interval of several weeks or months. | float with unit:none |  |  | ADDED | Months
| 2 | ExcitationCalibration | PowerLinearity |  | This field captures the Linear Correlation Coefficient of observed Output Illumination Power response to different input Power settings, as measured at the back of the Objective. | float with unit:none |  |  | ADDED
| 3 | ExcitationCalibration | ObservedExcitationIrradiance |  | Observed mean excitation light irradiance (radiant power per unit area) as measured at the sample plane over the Field of View (FOV). | float with unit:ObservedExcitationIrradianceUnit |  | R | Tentatively left until further feedback from David Grunwald. |  |  | Specified for one or all objectives? |  |  | As measured with a power meter? |  |  |  | The way this works the Calibration procedure is associated with the use of a specific set of Acquisiton parameters including the Objective used etc. The Intensity Calibration Tool is referenced.
| 3 | ExcitationCalibration | ObservedExcitationVariance |  | Observed excitation light irradiance (radiant power per unit area) variation over the Field of View (FOV). | float with unit:none |  | R | Tentatively left until further feedback from David Grunwald.
| 2 | DetectorCalibration |  |  | This procedure measures the performance of the dectector | [9 empty] | To me the best DetectorCalibration strategy is to measure the photo-electron conversion factor, fairly complex to do, probably as complex to report?
| 2 | DetectorCalibration | DetectorCalibrationImageRef |  | This is an empty element that refers to an Image (typically a z-stack) obtained as part of this Detector Calibration Procedure used in this instance and Acquired using the same Instrument and Settings employed to acquire the experimental Image this calibration refers to. | Extension of Reference |  | 0 ...∞ | ADDED OPTIONAL IMAGE REF
| 2 | DetectorCalibration | CalibrationMap |  | This element describes a 2D matrix containing values computed as a resulut of the calibration procedure. This matrix is the same size of the experimental Image this calibration refers to. | Element |  | 0 … ∞ | MADE THIS OPTIONAL
| 2 | CalibrationMap |  |  | This element describes a 2D matrix containing values computed as a resulut of the calibration procedure. This matrix is the same size of the experimental Image this calibration refers to. |  |  | 0 … ∞
| 2 | CalibrationMap | MetricType |  | This field captures the type of metric that was computed as part of the calibratin procedure in this instance. | enum | Gain, LinearOffset, DetectionLimit, LinearRangeMax, Variance, Mean, Other | R
| 2 | CalibrationMap | IlluminationIntensity |  | This attribute only applies to matrices of type Variance and Mean. It represents the illumination intensity that was used to acquire the images used to contruct this matrix. The value can range from 0 (i.e. dark image) to any intenisty value experssed either as percentage or as absolute power. | float with unit:none |  | R
| 2 | CalibrationMap | NrOfImages |  | This field represents the number of images that were employed to generate this calibration matrix. | integer with unit:none |  | R | [5 empty] | do you envisage one image with different Voltages/gains per timepoint? this is awkward/ impossible to do in some softwares. Or do you only envisage a tool output eg metamax?
| 2 | CalibrationMap | Description |  | A free text multi-line optional description for the Calibration Metric reported as part of this Detector Calibration procedure. | Element |  | 0 …1 | ADDED OPTIONAL MULTI-LINE DESCRIPTION
| 2 | CalibrationMap | Row |  |  | Element |  | 1 … ∞
| 2 | Row | [5 empty] | 1 … ∞ | [5 empty] | I presume this is for camera based intensity QC?
| 2 | Row | RowNr |  | The index of this row | integer with unit:none |  | R
|  | Row | Cell |  |  | Element |  | 1 … ∞
| 2 | Cell | [5 empty] | 1 … ∞
| 2 | Cell | CellNr |  | The index of this cell. | integer with unit:none |  | R
| 2 | Cell | CellValue |  | The value of the metric contained in this cell. | float with unit:CellValueUnit |  | R
| 2 | MechanicalCalibration |  |  | This describes the procedures that were used for mechanical calibration and the resulting calibration results. |  |  | 0 … ∞ | Separated X, Y and Z accuracy (repetability) + precision (reproducibility) make it clear in description + SD for each measure (optional) |  |  |  |  | some systems can run the stage at different speeds which impacts the accuracy. It would be good to have the stage speed recorded - I don't knwo if this is captured in current metadata (I doubt it).
| 2 | MechanicalCalibration | ID |  | A Unique Identifier for this component. |  |  | R
| 2 | MechanicalCalibration | Date |  | Ths field records the Date in which this Calibration procedure was performed. |  |  | R
| 2 | MechanicalCalibration | MeasuredXReproducibility |  | This field captures the observed Precision with which this Instrument is capable of maintaining a given X-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. This metric is generally referred to as Reproducibility and it might also be referred to as X Lateral Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:MeasuredXReprducibilityUnit |  | R | CHANGED NAME
| 2 | MechanicalCalibration | MeasuredYReproducibility |  | This field captures the observed Precision with which this Instrument is capable of maintaining a given Y-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. This metric is generally referred to as Reproducibility and it might also be referred to as Y Lateral Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:MeasuredYReprducibilityUnit |  | R | ADDED
| 2 | MechanicalCalibration | MeasuredFocalReproducibility |  | This field captures the observed Precision with which this Instrument is capable of maintaining a given Z-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. This metric is generally referred to as Reproducibility and it might also be referred to as Focal Drift. This metric is particularly relevant for Stereology systems and for the calibration of Sample Positioning devices. | float with unit:MeasuredFocalReprducibilityUnit |  | R | CHANGED NAME
| 2 | MechanicalCalibration | MeasuredXRepeatability |  | This field captures the observed X-axis Accuracy with which this Instrument is capable of returning to a previously visited position during Image Acquisition and as measured by this Mechanical Calibration Procedure. This metric is important for example for Tiling. | float with unit:MeasuredXRepeatablityUnit |  | R | CHANGED NAME
| 2 | MechanicalCalibration | MeasuredYRepeatability |  | This field captures the observed Y-axis Accuracy with which this Instrument is capable of returning to a previously visited position during Image Acquisition and as measured by this Mechanical Calibration Procedure. This metric is important for example for Tiling. | float with unit:MeasuredYRepeatablityUnit |  | R | ADDED |  |  | x,y drift, what about x,y accuracy as required Stereology systems? |  | data s means? if so, n number and SD would be nice.
| 2 | MechanicalCalibration | MeasuredFocalRepeatability |  | This field captures the observed Z-axis Accuracy with which this Instrument is capable of returning to a previously visited position during Image Acquisition and as measured by this Mechanical Calibration Procedure. This metric is important for example for Tiling. | float with unit:MeasuredFocalRepeatablityUnit |  | R | CHANGED NAME |  |  | z-drift, what about z- accuracy for piezo or stand calibration?
| 2 | MechanicalCalibration | MeasuredXSettlingTime |  | This field captures the observed Time it takes for this Instrument to stabilize at the target X-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:MeasuredXSettlingTimeUnit |  |  | CHANGED NAME
| 2 | MechanicalCalibration | MeasuredYSettlingTime |  | This field captures the observed Time it takes for this Instrument to stabilize at the target Y-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:MeasuredYSettlingTimeUnit |  |  | ADDED |  |  | repositioning accuracy should be measured in x,y & z.
| 2 | MechanicalCalibration | MeasuredFocalSettlingTime |  | This field captures the observed Time it takes for this Instrument to stabilize at the target Z-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:MeasuredFocalSettlingTimeUnit |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredXReproducibility_SD |  | This field captures the Standard Deviation associated with the observed Precision with which this Instrument is capable of maintaining a given X-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredYReproducibility_SD |  | This field captures the Standard Deviation associated with the observed Precision with which this Instrument is capable of maintaining a given Y-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredFocalReproducibility_SD |  | This field captures the Standard Deviation associated with the observed Precision with which this Instrument is capable of maintaining a given Z-axis position over time during Image Acquisition and as measured by this Mechanical Calibration procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredXRepeatability_SD |  | This field captures the Standard Deviation associated with the observed X-axis Accuracy with which this Instrument is capable of returning to a previously visited position (useful for tiling) during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredYRepeatability_SD |  | This field captures the Standard Deviation associated with the observed Y-axis Accuracy with which this Instrument is capable of returning to a previously visited position (useful for tiling) during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredFocalRepeatability_SD |  | This field captures the Standard Deviation associated with the observed Z-axis Accuracy with which this Instrument is capable of returning to a previously visited position (useful for tiling) during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredXSettlingTime_SD |  | This field captures the Standard Deviation associated with the observed Time it takes for this Instrument to stabilize at the target X-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredYSettlingTime_SD |  | This field captures the Standard Deviation associated with the observed Time it takes for this Instrument to stabilize at the target Y-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | MeasuredFocalSettlingTime_SD |  | This field captures the Standard Deviation associated with the observed Time it takes for this Instrument to stabilize at the target Z-axis position during Image Acquisition and as measured by this Mechanical Calibration Procedure. | float with unit:none |  |  | ADDED
| 2 | MechanicalCalibration | SamplePositioningRef |  | This empty element refers to the Sample Positioning device that was used for the Acquisition of the experimental Image this Mechanical Calibration procedure refers to. | Extension of Reference |  | 0 ... 1
| 2 | MechanicalCalibration | FocusStabilizationDeviceRef |  | This empty element refers to the optional Focus Stabilization Device that was used for the Acquisition of the experimental Image this Mechanical Calibration procedure refers to. | Extension of Reference |  | 0 ... 1
| 2 | MechanicalCalibration | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | Extension of Reference |  | 0 ... ∞

## In prog_4DN-BINA SPW Revision
4DN-BINA-OME Model |  | APE = Abstract Parent Element | [10 empty] | Damir additions | Pablo additions | Caterina additions | Jim additions | Nathalie additions | Ulrike additions
| Tier | Element | Attribute/sub-Element | Complex type | Description | Cardinality | Data type | Allowed values | OME 6-2016 | OME Cardinality/Required? | OME Data type |  | Damir Comments | Pablo Comments | Caterina comments | Jim comments | Nathalie comments | Ulrike comments
SPW | 3 | Screen |  |  | This describes a set of plates or other multiwell sampleholders that belong to a single experiment |  |  |  | same
| 3 | Screen | ID |  | A Unique Identifier for this Screen | R | ScreenID |  | same
| 3 | Screen | Name |  | A user assigned Name for this Screen |  | Denomination |  | same
| 3 | Screen | Description |  | A long description for this Screen |  | string |  | same
| 3 | Screen | ProtocolIdentifier |  | A pointer to an externally defined protocol, usually in a screening database. |  | string |  | same |  |  |  | These 4 are just external references. I don't think we're intending to capture "biology" metadata here, right?
| 3 | Screen | ProtocolDescription |  | A description of the screen protocol; may contain very detailed information to reproduce some of that found in a screening database. |  | string |  | same
| 3 | Screen | ReagentSetIdentifier |  | A pointer to an externally defined set of reagents, usually in a screening database/automation database. |  | string |  | same
| 3 | Screen | ReagentSetDescription |  | A description of the set of reagents; may contain very detailed information to reproduce some of that information found in a screening database. |  | string |  | same
| 3 | Screen | Type |  | A human readable identifier for the screen type; e.g. RNAi, cDNA, SiRNA, etc. This string is likely to become an enumeration in future releases. |  | string |  | same
| 3 | Screen | SampleHolderRef |  | A reference to zero or more samp,e holders in this Screen | 0 … ∞ | Element |  | ref to the more generic SampleHolder instead of PlateRef
| 3 | Screen | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference |  | same
| 3 | SampleHolder-APE |  |  | This describes a generic sample holder which can be a plate, slide, dish, etc. This sample holder can be part of a Screen (of multiple sample holders) or singly part of an Experiment |  |  |  | new APE to allow other types of SampleHolders besides Plate |  |  | ADDED; MAKE THIS A SUB-TYPE OF SAMPLE HOLDER
| 3 | SampleHolder-APE | ID |  | A Unique Identifier for this SampleHolder | R | SampleHolderID |  | moved up from Plate
| 3 | SampleHolder-APE | Name |  | A user assigned Name for this SampleHolder |  | Denomination |  | moved up from Plate
| 3 | SampleHolder-APE | Description |  | A long description for this SampleHolder |  | string |  | moved up from Plate
| 3 | SampleHolder-APE | OriginX |  | This defines the X position (relative to the Microscope's reference frame) to use for the origin of the fields (all individual images) taken from this SampleHolder. It is used with the X in the WellSample or Scene to display the fields in the correct position relative to each other. |  | float with unit:UnitsLength |  | moved up from Plate
| 3 | SampleHolder-APE | OriginY |  | This defines the Y position (relative to the Microscope's reference frame) to use for the origin of the fields (all individual images) taken from this SampleHolder. It is used with the Y in the WellSample or Scene to display the fields in the correct position relative to each other. |  | float with unit:UnitsLength |  | moved up from Plate
| 3 | SampleHolder-APE | OriginZ |  | This defines the Z position (relative to the Microscope's reference frame) to use for the origin of the fields (all individual images) taken from this SampleHolder. It is used with the Z in the WellSample or Scene to display the fields in the correct position relative to each other. |  | float with unit:UnitsLength |  | moved up from Plate
| 4 | SampleHolder-APE | SpatialTransformMatrix |  | A 4x4 matrix relating this SampleHolder's position/orientation/scaling to real-world coordinates (e.g. the Microscope's reference frame). This handles all alignment issues of the sample holder due to its positioning in the Microscope's stage. | R | 4x4 matrix of float with unit:none |  | new |  |  | ALIGN WITH WHAT WE HAVE FOR PLANE
| 5 | SampleHolder-APE | ExternalIdentifier |  | The ExternalIdentifier attribute may contain a reference to an external database |  | string |  | moved up from Plate
| 3 | SampleHolder-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference |  | moved up from Plate
| 3 | Plate |  |  | This describes a plate or similar multiwell sample holder with a rigid and defined spatial geometry that belongs to an experiment or a screen |  |  |  | same
| 3 | Plate | ColumnNamingConvention | NamingConvention | The convention used to denote columns in this Plate |  | enum | letter, number | same
| 3 | Plate | Columns |  | Number of columns in this Plate |  | PositiveInteger |  | same
| 3 | Plate | RowNamingConvention | NamingConvention | The convention used to denote rows in this Plate |  | enum | letter, number | same
| 3 | Plate | Rows |  | Number of rows in this Plate |  | PositiveInteger |  | same
| 3 | Plate | Status |  | A textual annotation of the current state of the plate with respect to the experiment work-flow |  | string |  | same
| 3 | Plate | PlateAcquisitionRef |  | A reference to zero or more plates Acquisitions for this Plate | 0 … ∞ | Element |  | same
| 3 | Plate | WellRef |  | A reference to the wells in this Plate | 1 … ∞ | Element |  | same
| 3 | GenericHolder |  |  | This describes a non-structured sample holder such as a slide |  |  |  | new to allow other types of holders |  |  |  | Better Name for OtherHolder?
| 3 | GenericHolder | SizeX |  | The usable extent in X of this Holder from the SampleHolder's origin. | R | float with unit:UnitsLength |  | new to allow other types of holders
| 3 | GenericHolder | SizeY |  | The usable extent in Y of this Holder from the SampleHolder's origin. | R | float with unit:UnitsLength |  | new to allow other types of holders
| 3 | GenericHolder | SceneRef |  | A reference to zero or more Scene Acquisitions for this OtherHolder | 0 … ∞ | Element |  | new to allow other types of holders
| 3 | PlateAcquisition |  |  | PlateAcquisition is used to describe a single Acquisition run for a plate. This object is used to record the set of images acquired in a single Acquisition run. The Images for this run are linked to PlateAcquisition through WellSample. |  |  |  | same
| 3 | PlateAcquisition | ID |  | A Unique Identifier for this PlateAcquisition | R | PlateAcquisitionID |  | same
| 3 | PlateAcquisition | Name |  | A user assigned Name for this PlateAcquisition |  | string |  | same
| 3 | PlateAcquisition | Description |  | A long description for this PlateAcquisition |  | string |  | same
| 3 | PlateAcquisition | StartTime |  | Time when the first image of this Acquisition was collected |  | float with unit:TimestampUnit |  | same
| 3 | PlateAcquisition | EndTime |  | Time when the last image of this Acquisition was collected |  | float with unit:TimestampUnit |  | same
| 3 | PlateAcquisition | MaximumFieldCount |  | The maximum number of fields (well samples) in any well in this PlateAcquisition. This is only used to speed up user interaction by stopping the reading of every well sample. |  | PositiveInteger |  | same
| 3 | PlateAcquisition | WellSampleRef |  | A reference to zero or more well samples for this PlateAcquisition | 0 … ∞ | Element |  | same
| 3 | PlateAcquisition | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference |  | same
| 3 | Well |  |  | This describes a well in a Plate |  |  |  | same
| 3 | Well | ID |  | A Unique Identifier for this Well | R | WellID |  | same
| 3 | Well | Name |  | A user assigned Name for this Well |  | Denomination |  | same
| 3 | Well | Description |  | A long description for this Well |  | string |  | same
| 3 | Well | Column |  | Column number of this Well in its Plate |  | PositiveInteger |  | same
| 3 | Well | Row |  | Row number of this Well in its Plate |  | PositiveInteger |  | same
| 3 | Well | CenterX |  | The X position of the center of the well relative to the origin defined on the Plate. |  | float with unit:UnitsLength |  | added to allow placing wellsample into the well extent
| 3 | Well | CenterY |  | The X position of the center of the well relative to the origin defined on the Plate. |  | float with unit:UnitsLength |  | added to allow placing wellsample into the well extent
| 3 | Well | WellWidth |  | The width of the well. |  | float with unit:UnitsLength |  | added to allow placing wellsample into the well extent
| 3 | Well | WellHeight |  | The height of the well. |  | float with unit:UnitsLength |  | added to allow placing wellsample into the well extent
| 3 | Well | Shape |  | The shape of the well |  | enum | Round, Square, RoundedSquare, Other | added to allow placing wellsample into the well extent
| 3 | Well | Type |  | A human readable identifier for the screening status. e.g. empty, positive control, negative control, control, experimental, etc. |  | string |  | same
| 3 | Well | Color |  | A marker color used to highlight the well - encoded as RGBA |  | string |  | same
| 5 | Well | ExternalIdentifier |  | The ExternalIdentifier attribute may contain a reference to an external database. |  | string |  | same
| 5 | Well | ExternalDescription |  | A description of the externally defined identifier for this well |  | string |  | same
| 3 | Well | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference |  | same
| 3 | WellSample |  |  | This describes a well sample in a PlateAcquisition |  |  |  | same
| 3 | WellSample | ID |  | A Unique Identifier for this WellSample | R | WellSampleID |  | same
| 3 | WellSample | Index |  | This records the order of the well samples. Each index should be unique for a given plate but they do not have to be sequential, there may be gaps if part of the dataset is missing. In the user interface the displayed value of the index will be calculated modulo the number of PlateAcquisitions for the plate. | R | PositiveInteger |  | same
| 3 | WellSample | PositionX |  | The X position of the field (image) within the well relative to the origin defined on the Plate. |  | float with unit:UnitsLength |  | same
| 3 | WellSample | PositionY |  | The Y position of the field (image) within the well relative to the origin defined on the Plate. |  | float with unit:UnitsLength |  | same
| 3 | WellSample | PositionZ |  | The Z position of the field (image) within the well relative to the origin defined on the Plate. |  | float with unit:UnitsLength |  | added
| 3 | WellSample | Timepoint |  | The time-point at which the image started to be collected |  | float with unit:TimestampUnit |  | same
| 3 | WellSample | ImageRef |  | The reference to the Image acquired at this WellSample | R | Element |  | same
| 3 | WellSample | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference |  | added
| 3 | Scene |  |  | This describes an unstructured Scene in a multi-scene Acquisition in a free-form SampleHolder such as a slide |  |  |  | new |  |  | ADDED
| 3 | Scene | ID |  | A Unique Identifier for this Scene | R | SceneID
| 3 | Scene | Description |  | A long description for this Scene |  | string
| 3 | Scene | Index |  | This records the order of the scenes. Each index should be unique for a given Sample but they do not have to be sequential, there may be gaps if part of the dataset is missing. | R | PositiveInteger
| 3 | Scene | PositionX |  | The X position of this Scene (image) within the SampleHolder relative to the origin defined on the SampleHolder. |  | float with unit:UnitsLength
| 3 | Scene | PositionY |  | The Y position of this Scene (image) within the SampleHolder relative to the origin defined on the SampleHolder. |  | float with unit:UnitsLength
| 3 | Scene | PositionZ |  | The Z position of this Scene (image) within the SampleHolder relative to the origin defined on the SampleHolder. |  | float with unit:UnitsLength
| 3 | Scene | Timepoint |  | The time-point at which the image started to be collected |  | float with unit:TimestampUnit
| 3 | Scene | ImageRef |  | The reference to the Image acquired at this Scene | R | Element
| 3 | Scene | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference
Note on C17: ====== ID#AAAAGwhVmUw Caterina Strambio De Castillia (2020-07-28 17:19:12) MAKE IT INTO A SETTINGS

## DEPR_4DN-BINA Confocal Ext.
4DN-BINA-OME Model |  | APE = Abstract Parent Element | [6 empty] | DECISION | Damir additions | Pablo additions | Caterina additions | Jim additions | Nathalie additions | Ulrike additions
| Tier | Element | Attribute or Subelement | Complex type | Description | Cardinality | Data type | Allowed values |  | Damir additions | Pablo additions | Caterina additions | Jim additions | Nathalie additions | Ulrike additions
SPECS | 1 | Instrument |  |  | This element describes the Instrument used to capture the Image. It is primarily a container referring to the Microscope's components list, which include the Microscope Stand, Microscope Table, Light Source, Light SourceCoupling, SampleHolder, Focusing devices, Objective, Optovar, Filter, Dichroic, AdditionalOptics, Detector and CalibrationTool. The Filter element can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. Within the Image itself, a reference is made to this one Filter element. There may be multiple Light Sources, detectors, Objectives and filters etc. in an Instrument. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the Ligth Path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. | 1
Suggestion to change the Name to ConfocalScanner to allow for non-laser confocals
| 2 | ConfocalScanner |  | Manufacturer Specs | A laser scanner is a device which can be used for controlling the direction of an outgoing laser beam. Typically, it is used for systematic periodic scanning (raster scanning), either in one dimension or in two dimensions |  |  |  | DAMIR WILL REVIEW | I changed a few small things in order to accommodate spinning disk, swept-field, grid multi-point, etc. scanners
| 2 | ConfocalScanner | ID |  | A Unique Identifier for this component. | R | MicroscopeID
| 2 | ConfocalScanner | Name |  | A User-defined Name for this component. | R | Denomination
| 3 | ConfocalScanner | HorizontalScanningSpeed |  | Maximal horizontal scanning speed. |  | float with unit:HorizontalScanningSpeed
| 3 | ConfocalScanner | VerticalScanningSpeed |  | Maximal vertical scanning speed |  | float with unit:VerticalScanningSpeed
| 3 | ConfocalScanner | MaxFrameRate |  | Maximum full frame rate |  | float with unit:MaxFrameRateUnits |  |  | added for multi-spot scanners
| 3 | ConfocalScanner | Resonant |  | This records whether or not this is a resonant scanner. |  | boolean
| 3 | ConfocalScanner | NrOfScanPoint |  | This records the number of scanning points available for this scan head. |  | integer with unit:none
| 3 | ConfocalScanner | PointsConfiguration |  | Description of the arrangement of scan points (can be spiral, linear, grid, other) |  |  |  | ADD
| 3 | ConfocalScanner | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … ∞ | Extension of Reference
IMAGE ACQUISITION Settings
| 1 | Image |  |  | This element describes the actual image that was acquired and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document. | 1
| 1 | Pixels (should merge with Image) |  |  | Pixels is going to be removed in the future, but it is still required. [...]. | 1 … ∞
| 1 | Channel |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. | 1 … ∞
| 1 | Plane |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. | 1 … ∞
| 2 | LightPath |  |  | A description of the Ligth Path. | 1
| 3 | LightPath | LaserScanningSettings |  | This element refers to the laser scanner used to acquire this image and it holds the settings that were applied to the laser scanner during this Acquisition. | 0 … ∞ | Element
| 3 | ConfocalScannerSettings |  |  | This element refers to the laser scanner used to acquire this image and it holds the settings that were applied to the laser scanner during this Acquisition. | 0 … ∞
| 3 | ConfocalScannerSettings | ID (ConfocalScannerID) |  | This is the Unique Identifier of the device used in this case.
| 3 | ConfocalScannerSettings | ConfocalFieldNumber |  | The field number of the confocal scan optic is expressed as the objective magnification times the maximum diameter of the scanned field; [mm], obtained without shifting the position of the specimen and without stitching single images into a total larger image. Because images generally have a square format, the diameter of the scanned field is equals to the square root of 2 times the length of X dimension of the scanned field. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. |  | float with unit:none
| 3 | ConfocalScannerSettings | PinholeSize |  | This records the selected pinhole size of the fixed-sized pinholes of the scanner (typical for spinning disk, swept-field, or other multi-point scanners to have multiple to choose from) |  | integer with unit:PinholeSizeUnits |  |  | spinning disk, swept-field, etc. scanners have disks or plates with fixed-size pinholes and with some you can select which disk/plate you want, thus listed in Settings
| 3 | ConfocalScannerSettings | ScanMode |  | In unidirectional scan mode, data are only acquired while the laser beam moves in one direction. In bidirectional scan mode, one line is scanned in one direction, and the next line in the opposite direction. In other words, the return run, or flyback, of the laser beam is used for recording data as well, thus increasing the scanning frequency when using bidirectional scan mode. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | R | enum | Unidirectional, Bidirectional, Other
| 3 | ConfocalScannerSettings | ScanFormat |  | This defines the format of the scanned element utilized for this Acquisition. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | R | enum | Line, Frame, Stack, Other
| 3 | ConfocalScannerSettings | MultiChannelMode |  | This defines whether the the the Aquisiton of Multiple Channels for this Image happened in a Parallel (i.e., multiple channels acquired at the same time), or Sequentail (i.e., one channel after the other) manner. | R | enum | Parallel, Sequential, Other | CHANGED | since Multi Tracking is mostly a Zeiss terminology (Leica calls it Parallel vs. Sequential) and the definition looks confusing, how about we use the more generic Leica Names and call this MultiChannelMode?
| 3 | ConfocalScannerSettings | Descanned |  | This indicates whether or not this acquistion involved descanning. |  | boolean
| 3 | ConfocalScannerSettings | Rotation |  | This records the rotation angle of the scan region |  | float |  | ADDED | added since many scan heads have this ability
| 3 | ConfocalScannerSettings | ScanTime |  | This is the time in seconds required to acquire a given number of lines, frames or stacks including the time required for repositioning the scanning device to the starting point. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | R | float with unit:ScanTimeUnit
| 3 | ConfocalScannerSettings | ScanningFrequency |  | ScanningFrequency is defined at the ratio between the number of scanning element (lines, frames or stacks) acquired and ScanTime. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. |  | float with unit:ScanningFrequencyUnit
| 3 | ConfocalScannerSettings | IntegrationMode |  | This defines the scanning element used as the basis of signal integration. Refer to https://www.iso.org/standard/69820.html for further details and measurement methods. | R | enum | Line, Frame, Stack, Other
| 3 | ConfocalScannerSettings | IntegrationMethod |  | This defines the signal integration method utilized to obtain this image, to improve the overall signal-to-noise ratio. | R | enum | Average, Sum, Other
| 3 | ConfocalScannerSettings | IntegrationNumber |  | This is the number of sequential scanning elements (Line, Frame or Stack) that have been integrated (Sum or Average) to obtain this image, to improve the overall signal-to-noise ratio. | R | integer with unit:none
| 3 | ConfocalScannerSettings | EffectiveZoom |  | The zoom factor, aka "Confocal Zoom" or "Scan Zoom", utilized for the Acquisition of this image. | R | float with unit:none