## Element categories_v00-04
Nr. | Domain | Category | Sub-category | Element
1 | MicrosocpeSpecifications | AdditionalOptics |  | AdditionalFilter, AdditionalDichroic, AdditionalLens, CouplingLens, RelayLens, PolarizationOptics, Prism
2 | MicrosocpeSpecifications | Calibration |  | IntensityCalibrationTool
3 | MicrosocpeSpecifications | ChildElement |  | Pump, LEDModule, WavelengthRange, TransmittanceRange, ReflectanceRance, Map
4 | MicrosocpeSpecifications | Detector |  | PhotoMultiplier, GenricDetector
5 | MicrosocpeSpecifications | Detector | Detector.Camera | CCD, IntensifiedCCD, AnalogVideo, CMOS
6 | MicrosocpeSpecifications | LightPath |  | FilterSet, ExcitationFilter, StandardDichroic, EmissionFilter
7 | MicrosocpeSpecifications | LightSource |  | Laser, Arc, Filament, LightEmittingDiode, GenericExcitationSource
8 | MicrosocpeSpecifications | LightSourceCoupling |  | LightGuide, Direct
9 | MicrosocpeSpecifications | Magnification |  | Objective, Optovar
10 | MicrosocpeSpecifications | MicroscopicBasic |  | Instrument, MicroscopyBody, MicroscopeTable
11 | MicrosocpeSpecifications | Mirror |  | Reflector, Beamsplitter
12 | MicrosocpeSpecifications | SamplePositioning | SamplePositioning.Stage | MechanicalStage, PiezoElectricStage, SampleHolder
13 | MicrosocpeSpecifications | SamplePositioning | SamplePositioning.Focusing | ObjectiveTurretFocus, PiezoElectricObjectiveFocus, AutoFocus
14 | ImageAcquisitionSettings | CalibrationProcedure |  | CalibrationMap, (Cell, Row), CalibrationSlide, ColoredBeadsSlide, DetectorCalibration, DNAOrigami, FieldCalibration, MechanicalCalibration, OpticalCalibration
15 | ImageAcquisitionSettings | Channel |  | Channel, AdditionalOptics, LightPath
16 | Experimental | SamplePreparation |  | Experiment, Experimenter, Sample
17 | ImageAcquisitionSettings | Image |  | Image, Plane
18 | ImageAcquisitionSettings | SampleConditions |  | FocalPosition, ImagingEnvironment, SamplePosition
19 | ImageAcquisitionSettings | Settings |  | CameraSettings, LightSourceSettings, MicroscopeSettings, ObjectiveSettings, PhotoMultiplierSettings, TIRFSettings
20 | ImageAcquisitionSettings | SampleManipulation |  | MicrobeamManipulation

## 4DN proposal_Specs_v01-07
4DN-OME model
| Tier | Element | Attribute or Subelement | Complex type | Description | Cardinality | Data type | Allowed values
MICROSCOPE Speecifications | 1 | Instrument |  |  | This element describes the instrument used to capture the Image. It is primarily a container referring to the microscope's components list, which include the Microscope Body, Microscope Table, LightSource, LightSourceCoupling, SampleHolder, Focusing devices, Objective, Optovar, Filter, Dichroic, AdditionalOptics, Detector and CalibrationTool. The Filter element can be composed either of separate excitation, emission filters and a dichroic mirror or a single filter set. Within the Image itself, a reference is made to this one Filter element. There may be multiple light sources, detectors, objectives and filters etc. in an Instrument. Each of these has their own ID attribute, which can be referred to from Channel. It is understood that the light path configuration can be different for each channel, but cannot be different for each timepoint or each plane of an XYZ stack. | 1
| 1 | Instrument | ID |  | A unique identifier for this component. | R | InstrumentID
| 1 | Instrument | Name |  | A user defined name for this instrument. | R | Denomination
| 2 | Instrument | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 2 | AnnotationRef |  | CommentAnnotation | This is a simple multi-line comment or annotation describing this component.
| 2 | AnnotationRef | ID |  | A unique identifier for this component. | R | AnnotationID
| 2 | CommentAnnotation |  |  | An annotation that consists of a simple multi-line comment.
| 2 | CommentAnnotation | ID |  | A unique identifier for this component. |  | AnnotationID
| 5 | CommentAnnotation | Namespace |  | We recommend the inclusion of a namespace for annotations you define. If it is absent then we assume the annotation is to use our (OME's) default interpretation for this type.
| 2 | CommentAnnotation | Annotator |  | The Annotator is the person who attached this annotation. e.g. If UserA annotates something with TagB, owned by UserB, UserA is still the Annotator.
| 2 | CommentAnnotation/Description | Descrption |  | This element contains a description for the annotation.[plain-text multi-line string] |  | Element
| 2 | CommentAnnotation/Value | Value |  | This element contains the actual content of this annotation. [multi-line string] |  | Element
| 2 | CommentAnnotation/AnnotationRef | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. |  | Extension of Reference
| 1 | MicroscopeBody |  |  | A description of the body of this microscope. | 1 … ∞
| 1 | MicroscopeBody | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | MicroscopeBody | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | MicroscopeBody | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | MicroscopeBody | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | MicroscopeBody | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | MicroscopeBody | ID |  | A unique identifier for this component. | R | MicroscopeID
| 1 | MicroscopeBody | Name |  | A user defined name for this component. | R | Denomination
| 1 | MicroscopeBody | Type | MicroscopeType | The type of this microscope. | R | enum | Upright, Inverted, Custom Made, Dissection, Electrophysiology, Stereomicroscope, Other
| 5 | MicroscopeBody | EyepieceFieldNumber |  | The Field Number represents the diaphragm size of the microscope eyepiece expressed in mm. |  | float with unit:none
| 2 | MicroscopeBody | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 …1 | Extension of Reference
| 4 | MicroscopeTable |  |  | This element describes the table used to support this microscope | 1
| 4 | MicroscopeTable | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 4 | MicroscopeTable | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | MicroscopeTable | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | MicroscopeTable | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | MicroscopeTable | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 4 | MicroscopeTable | ID |  | A unique identifier for this component. | R | LSID
| 4 | MicroscopeTable | Name |  | A user defined name for this component. | R | Denomination
| 4 | MicroscopeTable | Type | MicroscopeTableTYpe | This element describes the type of this microscope table. | R | enum | Simple Bench, Non-isolating Frame, Active Air, Passive Air, Stone, Other
| 5 | MicroscopeTable | TableWidth |  | The width of this microscope table |  | float with unit:TableDepthUnit
| 5 | MicroscopeTable | TableDepth |  | The depth of this microscope table |  | float with unit:TableWidthUnit
| 4 | MicroscopeTable | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | LightSource-APE |  | LightSource | This describes the light source(s) available for this instrument. An instrument may have several light sources of different types (Laser, LightEmittingDiode, Filament, Arc etc.). A typical light-microscope has two light sources, one for transmitted light and one for fluorescence light. | 1 … ∞ | Substitution : choose one of the children elements from the list below | Laser, Filament, Arc, LightEmittingDiode, GenericLightSource
| 1 | LightSource-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | LightSource-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | LightSource-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | LightSource-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | LightSource-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | LightSource-APE | ID |  | A unique identifier for this component. | R | LightSourceID
| 1 | LightSource-APE | Name |  | A user defined name for this component. | R | Denomination
| 3 | LightSource-APE | NominalPowerOutput |  | The average nominal power output of this light-source. |  | float with unit:NominalPowerUnit
| 2 | LightSource-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | Laser |  |  | A Light Amplification by Stimulated Emission of Radiation (LASER) is a device that emits light through a process of optical amplification based on the stimulated emission of electromagnetic radiation.
| 5 | Laser | Type | LaserType | The general category of this laser. | R | enum | Gas, SolidState, MetalVapor, LaserDiode, SolidStateDiodePumped, LED, Dye, FreeElectron, Other
| 5 | Laser | Pump |  | Wheter or not this is a pumped laser | R | boolean
| 5 | Laser | LaserMedium | LaserMedium | This is the material this laser consists of. |  | enum | Cu, Ag, ArFl, ArCl, ArKr, KrFl, KrCl, XeFl, XeCl, XeBr, N, Ar, Kr, Xe, HeNe, HeCd, CO, CO2, H2O, HFl, NdGlass, NdYAG, ErGlass, ErYAG, HoYLF, HoYAG, Ruby, TiSapphire, Alexandrite, Rhodamine6G, CoumarinC30, GaAs, GaAlAs, EMinus, Other
| 2 | Laser | Wavelength |  | The center Wavelength of the light emitted by this laser. | R | float with unit:Wavelength
| 5 | Laser | FrequencyMultiplication |  | The optical frequency multiplication that may be specified for this laser. |  | positiveInt with unit:none
| 5 | Laser | Tuneable |  | Whether or not the laser is Tuneable. | R | boolean
| 5 | Laser | Pulse | LaserPulse | The Pulse mode of the laser. | R | enum | CW, SingleMode, QSwitched, Repetitive, ModeLocked, Other
| 5 | Laser | PockelCell |  | If true the laser has a PockelCell to rotate the polarization of the beam. |  | boolean
| 5 | Laser | RepetitionRate |  | The rate in Hz at which the laser pulses if the Pulse type is 'Repetitive'. |  | foat with unit:RepetitionRateUnit
| 1 | Filament |  |  | A Filament lamp is an incandescent light source with a wire filament heated to such a high temperature that it glows with visible light (incandescence).
| 3 | Filament | Type | FilamentType | The type of Filament lamp light source. | R | enum | Standard, Halogen, Other
| 1 | Arc |  |  | An Arc lamp is a lamp that produces light by an electric arc (also called a voltaic arc), which is typically generated between metal electrodes through an inert gas in a glass bulb.
| 3 | Arc | Type | ArcType | The type of Arc lamp light source. | R | enum | Hg, Xe, HgXe, Other
| 1 | LightEmittingDiode |  |  | A Light Emitting Diode (LED) is a semiconductor light source that emits light when current flows through it. Electrons in the semiconductor recombine with electron holes, releasing energy in the form of photons. The color of the light (corresponding to the energy of the photons) is determined by the energy required for electrons to cross the band gap of the semiconductor. White light may be obtained by using multiple semiconductors or a layer of light-emitting phosphor on the semiconductor device.
| 3 | LightEmittingDiode | Type | LightEmittingDiodeType | The type of LightEmittingDiode light source. | R | enum | Unmounted, Mounted, Collimated, Other
| 2 | LightEmittingDiode | LEDModule |  | A Light Emitting Diode (LED) light source is often composed of multiple modules each with its own wavelenght characteristics. This element describes an individual LED module. Multiple of this elements may be used. | 1 … ∞ | Element
| 2 | LEDModule | PeakWavelenght |  | The nominal wavelength at which the amount of light power output is at a maximum. | R | float with unit: PeakWavelengthUnit
| 3 | LEDModule | PeakPower |  | The nominal light power output generated at peak wavelenght. | R | float with unit:PeakPowerUnit
| 5 | LEDModule | ProjectionAngle |  | The angle at which this light source projects light |  | integer with unit:ProjectionAngleUnit
| 3 | LEDModule | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50% | R | integer with unit: CutInUnit
| 3 | LEDModule | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50% | R | integer with unit: CutOutUnit
| 5 | LEDModule | WavelengthProfileFile |  | The location of a file containing the wavelenght profile of this component. | R | anyURI
| 1 | GenericExcitationSource |  |  | The GenericExcitationSource element is used to represent a light source as a collection of key/value pairs, stored in a Map. One of the other light source objects should always be used in preference to this if possible.
| 3 | GenericExcitationSource | Map |  | This is a Mapping of key/value pairs used to store customized metadata that describes this element. |  | element
| 4 | LightSourceCoupling-APE |  | LightSourceCoupling | This element describes the device(s)/method(s) used to couple the available light source(s) with this microscope body in order to transfer light to the specimen. A typical light-microscope has a light coupling device for each available light source. | 1 … ∞ | Substitution : choose one child from the list below | LighthGuide, Direct
| 4 | LightSourceCoupling-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 4 | LightSourceCoupling-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | LightSourceCoupling-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | LightSourceCoupling-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | LightSourceCoupling-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 4 | LightSourceCoupling-APE | ID |  | A unique identifier for this component. | R | LightSourceCouplingID
| 4 | LightSourceCoupling-APE | Name |  | A user defined name for this component. | R | Denomination
| 5 | LightSourceCoupling-APE | Transmittance |  | The maximum amount of light this LightSourceCoupling apparatus is capable of transmitting at the peak wavelenght. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 5 | LightSourceCoupling-APE | CouplingLensRef |  | This is a reference to a lens positioned at the connection between this light source coupling element and the microscope body. | 0 … 1 | Extension of Reference
| 4 | LightSourceCoupling-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 4 | LightGuide |  |  | A Light Guide coupling between the light source and the microscope generally consists of a casing enclosing a medium capable of transferring light from the illuminator to the back of the microsocpe body.
| 5 | LightGuide | Type | LightGuideType | The type of light guide used in this case. | R | enum | Fiber optics, Liquid, Other
| 5 | LightGuide | PeakWavelength |  | The wavelength associated with the transmittance peak for this light guide. | R | float with unit:PeakWavelengthUnit
| 5 | LightGuide | RefractiveIndex |  | The refractive index of the material this light guide is made of. | R | float with unit:none
| 5 | LightGuide | NumericalAperture |  | The numerical aperture of this light guide expressed as a floating point (real) number. Expected range 0.02 - 1.5. | R | float with unit:none
| 5 | LightGuide | AcceptanceAngle |  | The angle within which the light beam produced by a light source may be coupled with the bound modes of this light guide. | R | float with unit:AcceptanceAngleUnit
| 5 | LightGuide | Geometry | LightGuideGeometry | The shape of this light guide. |  | enum | Spot Light, Line Light, Ring Light, Backlight, Diffuse Axial, Other
| 5 | LightGuide | MaterialName | LightGuideMaterialName | The chemical composition or the non-proprietary trade designation of the material used for this light guide. |  | enum | Quartz, Silica, Borosilicate, Other
| 5 | LightGuide | SingleMode |  | Whether or not this light guide is single-mode. If not, then the guide is assumend to be multi-modal. |  | boolean
| 5 | LightGuide | WaveguideMode | LightGuideWaveguideMode | The transverse mode (or superposition of such modes) that are being imposed by this light guide onto the light wave traversing it. These modes generally follow different propagation constants. |  | enum | Transverse electromagnetic, Transverse electric, Transverse magnetic, Hybrid, Other
| 5 | LightGuide | Diameter |  | The outer diameter of this light guide. |  | float with unit:DiameterUnit
| 5 | LightGuide | WavelengthRange |  | This records a range of wavelengths this device is designed for, according to the manufacturer. Multiple of this elements may be used. | 1 | Element
| 5 | WavelengthRange | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | WavelengthBounds | The location of a file containing the wavelenght profile of this component. | R | anyURI
| 4 | Direct |  |  | A Direct coupling between the light source and the microscope consists either of a direct mount or of an optical light path that directs light across space to the back of the microscope body.
| 5 | Direct | Type | DirectType | The type of direct coupling available for this Instrument. | R | enum | Mounted, Optical light path, Other
| 5 | Direct | Mount | DirectMount | The mount utilized to couple the available light sources to this microscope body. | R | string
| 3 | SamplePositioning-APE |  | Substitution : choose one child from the list below | This describes a objective lens or stage positioning device available for this instrument. An instrument may have more than one of these devices of different type, such as mechanical stage, piezoelectric stage, objective turret and piezoelectric objective. | 1 … ∞ |  | MechanicStage, PiezoelectricStage, PiezoelectricObjective
| 3 | SamplePositioning-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | SamplePositioning-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | SamplePositioning-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | SamplePositioning-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | SamplePositioning-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | SamplePositioning-APE | ID |  | A unique identifier for this component. | R | FocusingID
| 3 | SamplePositioning-APE | Name |  | A user defined name for this component. | R | Denomination
| 3 | SamplePositioning-APE | ZTravelRange |  | The Z axis range this device is capable of moving in | R | float with unit:ZTravelRangeUnit
| 5 | SamplePositioning-APE | Resolution |  | The expected motion resolution of this device. | R | float with unit:ResolutionUnit
| 5 | SamplePositioning-APE | PositionLinearityError |  | The expected positioning error associated with this device. Typically expressed as % error. | R | float with unit:none
| 4 | SamplePositioning-APE | Repetability |  | The expected precision with which this device is capable of returning to a previously visited position. | R | float with unit:RepetabilityUnit
| 4 | SamplePositioning-APE | SettlingTime |  | The expected time it takes for this device to stabilize at the target position. |  | float with unit:SettlingTimeUnit
| 5 | SamplePositioning-APE | MaxVelocity |  | The maximum velocity this device is capable of moving at |  | float with unit:MaximumVelocity
| 2 | SamplePositioning-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 3 | ObjectiveTurretFocus |  |  | An objective turret positoning device available for this instrument.
| 3 | PiezoElectricObjectiveFocus |  |  | An optional piezoelectric objective positioning device associated with this instrument.
| 3 | Stage-APE |  |  | This describes a stage x-y positioning device associated with this instrument.
| 3 | Stage-APE | MovementAxes |  | The axis along which this device moves | R | enum | X, Y, Z, XY, XYZ, Other
| 5 | Stage-APE | XTravelRange |  | The X axis range this device is capable of moving in. | R | float with unit:XTravelRangeUnit
| 5 | Stage-APE | YTravelRange |  | The Y axis range this device is capable of moving in | R | float with unit:YTravelRangeUnit
| 3 | SampleHolder |  |  | This describes an optional sample holder mounted on the microscope stage.
| 3 | MechanicalStage |  |  | This describes a mechanical stage positioning device associated with this instrument. This component is not very common but can be found in simple surveying microscopes.
| 3 | PiezoElectricStage |  |  | This describes an optional piezoelectric stage positioning device associated with this instrument.
| 3 | AutoFocus |  |  | This element describes one or more optional focus stabilization device(s) available for this instrument. | 0 … ∞
| 3 | AutoFocus | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | AutoFocus | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | AutoFocus | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | AutoFocus | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | AutoFocus | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | AutoFocus | ID |  | A unique identifier for this component. | R | AutoFocusID
| 3 | AutoFocus | Name |  | A user defined name for this component. | R | Denomination
| 3 | AutoFocus | Type | AutoFocusType | The type of this autofocus device. | R | enum | Image based, Optical, Other
| 3 | AutoFocus | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | Optovar |  |  | Optovars are magnification changing devices originally developed by Zeiss, which provide additional magnification that can be beneficial when a desired field of view is required for the imaging sensor. Magnification changers are placed before the tube lens within the infinity space of the optical system. Typical magnification increments are: 1.25x, 1.6x, and 2x. | 0 … ∞
| 1 | Optovar | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Optovar | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Optovar | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Optovar | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Optovar | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | Optovar | ID |  | A unique identifier for this component. | R | ID
| 1 | Optovar | Name |  | A user defined name for this component. | R | Denomination
| 1 | Optovar | Magnification |  | The magnification associated with this magnification changing lens, as specified by the manufacturer. | R | float with unit:none
| 5 | Optovar | LensType |  | The type of lens associated with this magnification changing device. |  | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 2 | Optovar | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | Objective |  |  | The microscope's objective lens is a compound lens consisting of several simple lenses (elements), usually arranged along a common axis. | 1 … ∞
| 1 | Objective | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Objective | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Objective | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Objective | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Objective | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 1 | Objective | ID |  | A unique identifier for this component. | R | ObjectiveID
| 1 | Objective | Name |  | A user defined name for this component. | R | Denomination
| 2 | Objective | Correction | ObjectiveCorrection | The type of optical correction (i.e. chromatict, flatness, etc.) associated with this objective. | R | enum | Achro, Achromat, Achroplan, Apo, Fl, Fluar, Fluor, Fluotar, Neofluar, Pl, Plan, Plano, PlanApo, PlanFluor, PlanNeoFluar, SuperFluor, UPlanApo, UPlanFl, UV, VioletCorrected, Other
| 1 | Objective | NominalMagnification |  | The magnification of the lens as specified by the manufacturer - i.e. 60 represents a 60X lens. | R | float with unit:none
| 5 | Objective | CalibratedMagnification |  | The magnification of the lens as measured via a calibration process- i.e. '59.987' for a 60X lens. |  | float with unit:none
| 1 | Objective | LensNA |  | The numerical aperture of this objective expressed as a floating point (real) number. Expected range 0.02 - 1.5. | R | float with unit:none
| 2 | Objective | NominalImmersionMedium | ObjectiveNominalImmmersionMedium | The immersion medium this objective is designed for. | R | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 3 | Objective | Infinity |  | Records whether or not the objective is infinity corrected. | R | boolean
| 2 | Objective | CoverslipTickness |  | The coverlip thickness this objective is recommended for. | R | float with unit:none
| 3 | Objective | WorkingDistance |  | The working distance of this objective, expressed as a floating point (real) number. | R | float with unit:WorkingDistanceUnit
| 4 | Objective | Iris |  | Records whether or not the objective was fitted with an Iris. | R | boolean
| 3 | Objective | ObjectiveFieldNumber |  | The maximal diameter of the viewfield allowed by this objective, expressed in mm and measured at the intermediate image plane. NB: The field of view is influenced by the magnification and diameter of the ocular (eyepiece) field diaphragm, but also by the limits imposed by the design of the objective. | R | float with unit:none
| 5 | Objective | ParfocalizingDistance |  | The distance between the mounting plane of this objective and the specimen plane. | R | float with unit:ParafocalizingDistanceUnit[default mm]
| 5 | Objective | BackFocalPlanePosition |  | The position of the rear focal plane of this objective with respect to the specimen plane. | R | float with unit:BackFocalPlanePositionUnit[default mm]
| 2 | Objective | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 2 | Lens |  |  | This describes a lens positioned in the light path for one or more channel available for this instrument. A lens is a component of glass or transparent plastic material, usually circular in diameter, which has two primary surfaces that are ground and polished in a specific manner designed to produce either a convergence or divergence of light passing through the material. An instrument may have several lenses with different fuctions such as, a coupling lens positioned between the light source and the microscope body, a relay lens positioned at the detector mount, or an additional lens positioned in the light path for various purposes. | 0 … ∞
| 2 | Lens | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Lens | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Lens | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Lens | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Lens | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Lens | ID |  | A unique identifier for this component. | R | LensID
| 2 | Lens | Name |  | A user defined name for this component. | R | Denomination
| 5 | Lens | Type | LensType | The general category this lens belongs to. | R | enum | Amici-Bertrand, Biconvex, Biconcave, Plano Convex, Plano Concave, Convex Meniscus, Concave Meniscus, Equiconvex, Other
| 2 | Lens | Magnification |  | The nominal magnification provided by this lens as specified by the manufacturer. | R | integer with unit:none
| 5 | Lens | Construction | LensConstruction | Whether or not this lens is simple or composite. | R | enum | Simple, Composite, Other
| 5 | Lens | Geometry | LensGeometry | The geometrical shape of this lens' surface(s). | R | enum | Spherical, Aspherical, Cylindrical, Ball, Hemispheric-ball, Cone, Drum, Rod, Toric
| 5 | Lens | RefractiveIndex |  | The refractive index of the material this lens is made of, as specified by the manufacturer. | R | float with unit:none
| 5 | Lens | AttenuationCoefficient |  | The attenuation coeffieient of this lens. The attenuation coefficient (aka blocking coefficient) is a material property that characterizes the degree of transparency of a lens. The attenuation coefficient is often defined in units of optical density (OD), wherease OD= -log10(T), and T=Transmittance (0≤T≤1). | R | float with unit:none
| 5 | Lens | FocalDistanceToObject |  | The distance between the front principal plane of this lens and the specimen plane. | R | float with unit:FocalDistanceToObjectUnit
| 5 | Lens | FocalDistanceToImage |  | The distance between the rear principal plane of this lens and the image plane. | R | float with unit:FocalDistanceToImageUnit
| 5 | Lens | FocalLenght |  | The distance from the rear principal plane of this lens and its focal plane. | R | float with unit:FocalLengthUnit
| 5 | Lens | Diameter |  | The physical diamter of this lens. |  | float with unit:none
| 5 | Lens | AbbeNumber |  | The Abbe number of this lens. The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion. |  | float with unit:none
| 5 | Lens | RadiusOfCurvature |  | The radius of curvature of this lens. A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusUnit
| 5 | Lens | MaterialName | LensMeterialName | This records the chemical composition or the nonproprietary trade designation of the material this lens is made of. |  | enum | alpha-Barium Borate, N-BK7 (Borosilicate Crown), Calcium Fluoride , UV-Grade Fused Silica (UVFS), Germanium, Magnesium Fluoride, N-SF11 (Dense Flint), Polytetrafluoroethylene (PFTE), Silicon (Si), Zinc Selenide (ZnSe), Other
| 5 | Lens | GlassCode |  | If available provide the international glass code. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642.[1] |  | sring
| 5 | Lens | Density |  | The density of the material this lens is made of. |  | float with unit:DensityUnit
| 2 | Lens | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 5 | CouplingLens |  |  | A lens positioned at connector between this light source coupling device and the microscope body.
| 3 | RelayLens |  |  | A lens positioned at the mount connecting a detector to the microscope body.
| 2 | AdditionalLens |  |  | Any additional lens positioned in the light path for a given channel.
| 2 | Prism |  |  | This describes a prism positioned in the light path for one or more channel available for this instrument. A prism is a block of glass or other transparent material with flat, polished surfaces, which can be viewed as an ensamble of different rigidly mounted mirrors, with each face having a permanent orientation with respect to all the others. A prism can be employed to deflect or deviate a light beam, rotate or invert an image, separate polarization states, or disperse light into its component wavelengths. An instrument may have several prisms with different factionality such as, reflecting, polarizing, and refracting or dispersion prisms. Reflecting prisms are useful for redirecting light beams by total internal reflection. Polarizing prisms are birefringent crystals that divide incident non-polarized light into separate components polarized orthogonally to each other. Refracting prisms can be employed to bend and separate light into its component colors. These prisms are employed to produce polarized light for optical instruments such as microscopes and polarimeters. | 0 … ∞
| 2 | Prism | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Prism | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Prism | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Prism | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Prism | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Prism | ID |  | A unique identifier for this component. | R | PrismID
| 2 | Prism | Name |  | A user defined name for this component. | R | Denomination
| 3 | Prism | Type | PrismType | The category this prism belongs to. | R | enum | Reflecting, Polarizing, Refracting, Other
| 5 | Prism | RefractiveIndex |  | The refractive index of the material this prism is made of, as specified by the manufacturer. | R | float with unit:none
| 5 | Prism | Reflectance |  | The maximum amount of light this Prism is capable of transmitting at the peak wavelenght. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 5 | Prism | AngleOfIncidence |  | The angle of incidence (AOI) is the angle between a beam of light incident on the prism surface and the line perpendicular to the surface at the point of incidence, called the normal. | R | float with unit:IncidenceAngleUnit
| 5 | Prism | DeviationAngle |  | A beam passing through an object like a prism is deflected twice: once entering, and again when exiting. The sum of these two deflections is called the deviation angle and is measured with respect to the incident ray. | R | float with unit:DeviationUnitAngle
| 5 | Prism | Geometry | PrismGeometry | The geometrical shape of this prism. |  | enum | Right-Angle, Porro, Dove, Amici, Penta, Romboid, Nicol, Glan-Thompson, Wollaston, Rochon, Senarmont, Triangular, Abbe prism, Pellin–Broca, Grim, Compound, Other
| 5 | Prism | AbbeNumber |  | The Abbe number of this prism. The Abbe Number is a measure of the material's dispersion (variation of refractive index versus wavelength), with high values indicating low dispersion. |  | float with unit:none
| 5 | Prism | PrismAngle |  | The angle between two surfaces of this prism. |  | float with unit:PrismAngleUnit
| 5 | Prism | MaterialName | PrismMaterialName | This records the chemical composition or the nonproprietary trade designation of the material this prism is made of. |  | enum | Calcium Fluoride, F2 (Flint), N-BK7 (Borosilicate Crown), N-SF11 (Dense Flint), Rutile (TiO2), UV-Grade Fused Silica (UVFS), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | Prism | GlassCode |  | If available provide the international glass code for the glass this prism is made of. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642. |  | sring
| 5 | Prism | Density |  | The density of the material the prism is made of. |  | float with unit:DensityUnit
| 5 | Prism | WavelengthRange |  | This records a range of wavelengths this device is designed for, according to the manufacturer. Multiple of this elements may be used. | 1 | Element
| 2 | Prism | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | PolarizationOptics |  |  | This describes an optical polarization device positioned in the light path for one or more channel available for this instrument. A polarization device is an optical filter that lets light waves of a specific polarization pass through while blocking light waves of other polarizations. | 0 … ∞
| 2 | PolarizationOptics | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | PolarizationOptics | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | PolarizationOptics | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | PolarizationOptics | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | PolarizationOptics | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | PolarizationOptics | ID |  | A unique identifier for this component. | R | PolarizationOpticsID
| 2 | PolarizationOptics | Name |  | A user defined name for this component. | R | Denomination
| 2 | PolarizationOptics | Function | PolarizationOpticsFunction | This field records whether this polarizer optics is used to generate or to analyze polarized light | R | enum | Polarizer, Analyzer
| 3 | PolarizationOptics | Construction | PolarizationOpticsConstruction | The modality with which this polarization device is constructed. |  | enum | Achromatic, Zero Order, Multiple Order, Other
| 5 | PolarizationOptics | Fabrication | PolarizationOpticsFabrication | The way in with which this polarization device is fabricated. |  | enum | Wire Grid, Film, Polarizing Beamsplitter, Polarizing BandPass, Variable, Birefringent, Polymer, Quarz, Other
| 3 | PolarizationOptics | Retardation | PolarizationOpticsRetardation | This describes the phase shift between the polarized light projected along the fast axis and the component projected along the slow axis. |  | enum | Half Wave, Quarter Wave, Full Wave, Other
| 5 | PolarizationOptics | MaterialName | PolarizationOpticsMaterialName | This records the chemical composition or the nonproprietary trade designation of the material this polarization device is made of. |  | enum | alpha-Barium Borate, Calcite, Calcium Fluoride, N-BK7 (Borosilicate Crown), Rutile (TiO2), Yttrium Orthovanadate (YVO4), Zinc Selenide (ZnSe), Other
| 5 | PolarizationOptics | GlassCode |  | If available provide the international glass code for the glass this polarization device is made of. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642. |  | sring
| 5 | PolarizationOptics | WavelengthRange |  | This records a range of wavelengths this device is designed for, according to the manufacturer. Multiple of this elements may be used. | 1 | Element
| 2 | PolarizationOptics | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelengthProfile | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | FilterSet |  |  | A filter set is typically composed of one excitation filter, one emission filter and a dichroic. They might either be pre-assebled by the manufacturer and pruchased as a whole or they can be custome assembled on site depending on the needs of the user. | 0 … ∞
| 2 | FilterSet | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | FilterSet | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | FilterSet | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | FilterSet | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | FilterSet | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | FilterSet | ID |  | A unique identifier for this component. | R | FilterSetID
| 2 | FilterSet | Name |  | A user defined name for this component. | R | Denomination
| 2 | FilterSet | ExcitationFilterRef | Reference | This refers to an available excitation filter, which is part of this filter set. | 0 … 1 | Extension of Reference
| 2 | FilterSet | StandardDichroicRef | Reference | This refers to an available dichroic mirror, which is part of this filter set. | 0 … 1 | Extension of Reference
| 2 | FilterSet | EmissionFilterRef | Reference | This refers to an available emission filter, which is part of this filter set. | 0 … 1 | Extension of Reference
| 2 | FilterSet | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 2 | Filter-APE |  |  | This describes an optical filter positioned in the light path for one or more channel available for this instrument. An optical filter is a device that selectively transmits light of different wavelengths. An instrument may have several filters of different types, such as an excitation, an emission filter or an additional filter positioned in the light path for different purposes. For each fluorescence Image there should be at least one filter element specified per Channel. This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet. | 1 … ∞
| 2 | Filter-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Filter-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Filter-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Filter-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Filter-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Filter-APE | ID |  | A unique identifier for this component. | R | FilterID
| 2 | Filter-APE | Name |  | A user defined name for this component. | R | Denomination
| 3 | Filter-APE | Type | FilterType | The type of filter used in this case. | R | enum | Bandpass, Multipass, Longpass, Shortpass, NeutralDensity, HeatFilter, Tunable, Other
| 5 | Filter-APE | ReflectionCoating |  | An optical coating is one or more thin layers of material deposited on an optical component such as a lens or mirror, which alters the way in which the optic reflects and transmits light. |  | string
| 3 | Filter-APE | FilterWheel |  | This can can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. |  | string
| 5 | Filter-APE | TransmittanceProfileFie |  | This contains the location of file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 2 | Filter-APE | TransmittanceRange |  | This records the range of wavelengths that are transmitted by the filter. It also records the maximum amount of light transmitted. Multiple of this elements can be used. | 1 … ∞ | Element
| 2 | Filter-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 2 | TransmittanceRange | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 2 | TransmittanceRange | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 3 | TransmittanceRange | Transmittance |  | The maximum amount of light this optical component is capable of transmitting at the peak wavelenght. It is represented as a fractional value from 0.0 to 1.0. | R | percentFraction with unit:none
| 2 | ExcitationFilter |  |  | An optical filter is a device that selectively transmits light of different wavelengths. This element refers to an excitation filter. This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | EmissionFilter |  |  | An optical filter is a device that selectively transmits light of different wavelengths. This element refers to an emission filter. This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | AdditionalFilter |  |  | An optical filter is a device that selectively transmits light of different wavelengths. This element refers to a filter positioned along the light path in addition to the excitation and the emission filters.This element can be associated with an optional FilterWheel - Note: FilterWheel is not the same as a FilterSet.
| 2 | Mirror-APE |  | Substitution : choose one child from the list below | This describes a mirror positioned in the light path for one or more channel available for this instrument. An instrument may have several mirrors of different types, such as dichroic, beamsplitter or reflector. | 1 … ∞ |  | Dichroic, Beamsplitter, Reflector
| 2 | Mirror-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | Mirror-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Mirror-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Mirror-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Mirror-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Mirror-APE | ID |  | A unique identifier for this component. | R | MirrorID
| 2 | Mirror-APE | Name |  | A user defined name for this component. | R | Denomination
| 5 | Mirror-APE | AngleOfIncidence |  | This field records the optimal angle of incidence (AOI) for this mirror. TheAOI is the angle between the optical axis of the incident light and the axis normal to the surface of the mirror. In specular reflection AOI is by definition equal to the Angle Of Reflection (AOR). | R | float with unit:AngleOfIncidenceUnit
| 5 | Mirror-APE | MirrorType | MirrorType | The general category this mirror belongs to. |  | enum | Convex, Concave, Planar, Other
| 5 | Mirror-APE | Geometry | MirrorGeometry | The geometrical shape of the surface(s) of this mirror. |  | enum | Spherical, Aspherical, Cylindrical, Ellipsoidal, Hyeprboloidal, Paraboloidal, Other
| 5 | Mirror-APE | RadiusOfCurvature |  | The radius of curvature of this mirror. A spherical lens or mirror surface has a center of curvature located either along or decentered from the system local optical axis. The vertex of the lens surface is located on the local optical axis. The distance from the vertex to the center of curvature is the radius of curvature of the surface. |  | float with unit:RadiusOfCurvatureUnit
| 5 | Mirror-APE | SubstrateType | SubstrateType | This field records the type of substrate that is used for the beamsplitter coating. It is typically a ground and polished piece of optical glass. |  | enum | Metallic, Dielectric, Other
| 5 | Mirror-APE | SubstrateMaterial | SubstrateMaterialName | This records the chemical composition or the nonproprietary trade designation of the mirror's optical substrate. For dichoric mirrors the substrate is typically composed of a glass material with low autofluorescence, such as UV-grade fused silica. |  | enum | Calcium Fluoride, N-BK7 (Borosilicate Crown), UV-Grade Fused Silica (UVFS), Zerodur (Lithium Aluminum Ailicon Oxide), Zinc Selenide (ZnSe), Other
| 5 | Mirror-APE | GlassCode |  | If available provide the international glass code for the glass this polarization device is made of. This is based on U.S. military standard MIL-G-174, and is a six-digit number specifying the glass according to its refractive index nd at the Fraunhofer d- (or D3-) line, and its Abbe number Vd also taken at that line. The resulting glass code is the value of nd-1 rounded to three digits, followed by Vd rounded to three digits, with all decimal points ignored. For example, BK7 has nd = 1.5168 and Vd = 64.17, giving a six-digit glass code of 517642. |  | sring
| 2 | Mirror-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 2 | Dichroic |  |  | A dichroic beamsplitter or dichroic mirror is a mirror that spectrally separate light by transmitting and reflecting light as a function of wavelength. In the context of fluorescence microscopy, they are typically used to reflect the excitation light, but transmit the emitted fluorescence. Dichroic mirrors can be categorized in different manners. However, the most important criteria for choosing a dichoric mirror are the transmittance and the reflectance wavelenght ranges.
| 3 | Dichroic | DichroicType |  | The type of dichroic mirror used in this case | R | enum | Bandpass, Multipass, Longpass, Shortpass, Other
| 3 | Dichroic | HardSputtered |  | This field records whether this coating of this dichroic is hard-sputtered or traditional. | R | boolean
| 3 | Dichroic | AttenuationCoefficient |  | The attenuation coeffieient of this dichroic The attenuation coefficient (aka blocking coefficient) is used as a measure of the out-of-band attenuation of an optical filter or dichroic mirror over an extended range of the spectrum. The attenuation coefficient is often defined in units of optical density (OD), wherease OD= -log10(T), and T=Transmittance (0≤T≤1). | R | float with unit:none
| 5 | Dichroic | AttenuationMethod |  | The mean by which this dichroic filter achieves out-of-band transmission attenuation. |  | enum | Absorption, Reflection, Rifraction, Other
| 5 | Dichroic | Technology |  | The filtering technology utilized by this dichroic. The main types of filter technology used in fluorescence analysis are absorption (aka colored) glass, thin-film coating, and acousto-optical tunable, liquid-crystal tunable filters. |  | enum | Absorption Glass, Metallic-coating, Thin-film interference-coating, Acousto-optical, Liquid-crystal, Other
| 5 | Dichroic | Polarization |  | This field records the polarization effect of this dichroic. Dichroic beamsplitters (and, in fact, any thin-film interference coating that is used at non-normal angles-of-incidence) will cause some amount of polarization. ), the precise effect varying greatly with wavelength and with the particular coating design. |  | enum | Planar, Circular, Elliptical, Other
| 3 | Dichroic | FilterWheel |  | This can can refer to any arrangement of filters in a filter holder of any shape. It could, for example, be a filter slider. |  | string
| 5 | Dichroic | TransmittanceProfileFile |  | This contains the location of a file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 5 | Dichroic | ReflectanceProfileFile | WavelengthBounds | This contains the location of a file in .csv or other suitable formats containing the wavelenght profile of this component. | R | anyURI
| 2 | Dichroic | TransmittanceRange |  | This records the range of wavelengths that are transmitted by the dichroic mirror. It also records the maximum amount of light transmitted | 1 … ∞ | Element
| 3 | Dichroic | ReflectanceRange |  | This records the range of wavelengths that are reflected by the mirror. | 0 … ∞ | Element
| 2 | TransmittanceRange | CutIn |  | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 2 | TransmittanceRange | CutOut |  | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 3 | TransmittanceRange | Transmittance |  | The maximum amount of light this optical component is capable of transmitting at the peak wavelenght. It is represented as a fractional value from 0.0 to 1.0. | R | percentFraction with unit:none
| 3 | ReflectanceRange | CutIn |  | CutIn is the wavelength below which the response rate is below 50% | R | positivefloat with unit:CutInUnit
| 3 | ReflectanceRange | CutOut |  | CutOut is the wavelength above which the response rate is below 50% | R | positivefloat with unit:CutOutUnit
| 3 | ReflectanceRange | Reflectance |  | The maximum amount of light this Mirror is capable of reflecting. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 2 | StandardDichroic |  |  | This describes a dichroic positioned in the standard light path for one or more channel available for this instrument.
| 2 | AdditionalDichroic |  |  | This describes an additional dichroic positioned in the light path for one or more channel available for this instrument.
| 2 | Beamsplitter |  |  | This is a mirror that reflects part of the light, regardless of wavelenght.
| 3 | Beamsplitter | Transmittance |  | The maximum amount of light this beamsplitteris capable of transmitting at the peak wavelenght. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 5 | Beamsplitter | TransmissionAngle |  | The angle at which light is transmitted by this beamsplitter. | R | float with unit:none
| 3 | Beamsplitter | Reflectance |  | The maximum amount of light this Mirror is capable of reflecting. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 2 | Reflector |  |  | This is a total reflection mirror.
| 3 | Reflector | Reflectance |  | The maximum amount of light this Mirror is capable of reflecting. It is represented as a fractional value from 0.0 to 1.0. | R | float with unit:none
| 2 | Detector-APE |  | Substitution : choose one child from the list below | This describes a detector available for this instrument. An instrument may have more than one detector of different types, such as photomultiplier, or camera. The Detector ID can be used as a reference within the Channel element in the Image element. The values stored in Detector represent the fixed values, variable values modified during the acquisition go in DetectorSettings. | 1 … ∞ |  | Camera, Photomultiplier
| 2 | Detector-APE | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 2 | Detector-APE | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | Detector-APE | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | Detector-APE | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | Detector-APE | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 2 | Detector-APE | ID |  | A unique identifier for this component. | R
| 2 | Detector-APE | Name |  | A user defined name for this component. | R | Denomination
| 3 | Detector-APE | Gain |  | The analog gain is the conversion unit for this detector, as specified by the manufacturer. It specifies how many electrons correspond to 1 signal unit. It is expressed in e-/Analog-to-Digit-Unit (ADU). | R | float with unit:none
| 5 | Detector-APE | Mount |  | The type of mount used to connect this detector to the microscope body. | R | enum | C-mount, T-mount, SLR adapter, D-SLR adapter, Other
| 5 | Detector-APE | ReadNoise |  | The read noise of this detector, as specified by the manufacturer. Read noise includes thermal noise build-up and clock induced charge. This is generally expressed in e- rms. | R | float with unit:ReadNoiseUnit
| 5 | Detector-APE | DetectorNoiseModel | DetectorNoiseModel | The type of noise associated with this detector. | R | enum | Poisson, Gaussian, Other
| 5 | Detector-APE | DarkCurrentRate |  | The dark current rate of this detector as specified by the manufacturer. The dark current rate is the median current detected in the absence of light. Generally expressed in e-/pixel/sec. | R | float with unit:DarkCurrentRateUnit
| 3 | Detector-APE | PeakWavelength |  | The wavelength associated with maximum light intensity detected by this device. | R | float with unit:PeakWavelengthUnit
| 3 | Detector-APE | QuantumEfficiencyatPeakWavelength |  | The quantum efficiency observed at peak wavelength, according to the manufacturer. Quantum Efficiency is generally defined as is the ratio of incident photon to converted electrons. Typically expressed as a number between 0 and 1. | R | float with unit:none
| 5 | Detector-APE | AmbientOperatingTemperature |  | The ambient temperature at which this detector is designed to operate, according to the manufactuer. |  | float with unit:TemperatureUnit; default: Celsius[ºC]
| 5 | Detector-APE | AbientOperatingHumidity |  | The ambient humidity at which this detector is designed to operate, according to the manufactuer. |  | float with unit:none
| 5 | Detector-APE | WavelengthRange |  | This records a range of wavelengths this device is designed for, according to the manufacturer. Multiple of this elements may be used. | 0 … 1 | Element
| 3 | Detector-APE | RelayLens |  | This is a reference to a lens positioned at the mount connecting this detector to the microscope body. | 0 … 1 | Extension of Reference
| 2 | Detector-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 5 | WavelengthRange | CutIn | WavelengthBounds | CutIn is the wavelength below which the response rate is below 50%. | R | positivefloat with unit:CutInUnit
| 5 | WavelengthRange | CutOut | WavelengthBounds | CutOut is the wavelength above which the response rate is below 50%. | R | positivefloat with unit:CutOutUnit
| 5 | WavelengthRange | WavelenghtProfile | WavelengthBounds | The location of a file containing the wavelenght profile of this component | R | anyURI
| 2 | Camera-APE |  | Substitution : choose one child from the list below | This describes a camera device available for this instrument. A camera is device that converts optical brightness into an electrical signal using a CCD(a semiconductor chip, which converts the intensity of light falling onto its surface at different points into a charge) and then reproduces an image of the specimen. An instrument may have more than one camera of different types, such as EMCCD, intensified EMCCD, CMOS, or analog video. |  |  | CCD, Intensified CCD, AnalogVideo, CMOS
| 3 | Camera-APE | Illumination | CameraIllumination | The type of illumination the camera chip receives. | R | enum | Front, Back, Other
| 3 | Camera-APE | ArrayWidth |  | The width of this camera chip expressed in pixel. | R | integer with unit:pixel
| 3 | Camera-APE | ArrayHeight |  | The height of this camera chip expressed in pixel. | R | integer with unit:pixel
| 3 | Camera-APE | PixelWidth |  | The physical width of each pixel. | R | float with unit:PixelWidthUnit; default:µm
| 3 | Camera-APE | PixelHeight |  | The physical height of each pixel. | R | float with unit:PixelHeightUnit; default:µm
| 3 | Camera-APE | ImageWidth |  | The width of the image generated by the camera. |  | float with unit:ImageWidthUnit
| 3 | Camera-APE | ImageHeight |  | The width of the image generated by the camera. |  | float with unit:ImageHeightUnit
| 5 | Camera-APE | PixelWellCapacity |  | The maximum amount of charge each pixel well can hold, measured in electron counts. | R | integer with unit:PixelWellCapacity
| 5 | Camera-APE | RegisterWellCapacity |  | The maximum amount of charge each register well can hold, measured in electron counts. |  | integer with unit:RegisterWellCapacity
| 5 | Camera-APE | VerticalClockSpeed |  | The maximal speed each vertical colunm of pixels in the camera chip can be shifted down and read. |  | float with unit:VerticalClockSpeedUnit
| 5 | Camera-APE | MaximumFrameRate |  | The maximal rate at which this camera can acquire a single frame. | R | float with unit:MaximumFrameRateUnit [default fps]
| 5 | Camera-APE | MaximumReadoutRate |  | The maximal pixel read out rate allowed by this camera | R | float with unit:MaximalReadoutRateUnit[default pixel per second]
| 2 | CCD |  |  | A charged coupled device camera.
| 3 | CCD | Type | CCDType | The type of CCD camera used. | R | enum | Conventional, EMCCD, Interline, Other
| 2 | IntensifiedCCD |  |  | An intensified charged coupled device camera.
| 3 | IntensifiedCCD | Type | IntensifiedCCDType | The type of Intensified CCD camera used. | R | enum | Conventional, EBCCD, Other
| 2 | AnalogVideo |  |  | A device used to capture analog videos.
| 2 | CMOS |  |  | A complementary metal–oxide–semiconductor camera
| 3 | CMOS | Type | CMOSType | The type of CMOS camera used. | R | enum | Conventional, Scientific, Other
| 2 | PhotoMultiplier |  |  | A photomultiplier tube, useful for light detection of very weak signals, is a photoemissive device in which the absorption of a photon results in the emission of an electron. These detectors work by amplifying the electrons generated by a photocathode exposed to a photon flux.
| 3 | PhotoMultiplier | Type | PhotoMultiplierType | The type of photomultiplier used. | R | enum | PMT, Photodiode, APD, Other
| 5 | PhotoMultiplier | RadiantSensitivity |  | The radiant sensitivity observed with this PhotoMultiplier at peak wavelength | R | float with unit:RadiantSensitivityAtPeakWavelengthUnit
| 2 | GenericDetector |  |  | This element is used to represent a detector as a collection of key/value pairs, which are stored in a Map. One of the other detector elements should always be used in preference to this if possible.
| 3 | GenericDetector | Map |  | This is a Mapping of key/value pairs used to store customized metadata that describes this element. |  | element
| 3 | IntensityCalibrationTool |  |  | A tool, such as the Meta-Max calibration toold developed by the 4D Nucleome Imaging Standards Working Group, used to perform an intensity calibration procedure | 0 … ∞
| 3 | IntensityCalibrationTool | Manufacturer | Manufacturer Spec | The manufacturer of this component. | R | string
| 3 | IntensityCalibrationTool | Model | Manufacturer Spec | The model of this component. | R | string
| 5 | IntensityCalibrationTool | SerialNumber | Manufacturer Spec | The serial number of this component. | R | string
| 5 | IntensityCalibrationTool | LotNumber | Manufacturer Spec | The lot number of this component. |  | string
| 5 | IntensityCalibrationTool | SpecsFile | Manufacturer Spec | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | anyURI
| 3 | IntensityCalibrationTool | ID |  | A unique identifier for this component. | R
| 3 | IntensityCalibrationTool | Name |  | A user defined name for this component. | R | Denomination
| 3 | IntensityCalibrationTool | Type |  | The type of calibration proceudure this tool is intended to perform | R | enum | Detector, Excitation, Other
| 3 | IntensityCalibrationTool | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which 1): BOXES represent real-life entity-sets, about which information need to be collected. One way to think about it is that ENTITIES correspond to NOUNS. 2) LINES represent RELATIONSHIPS that connect entities with each other. Namely, RELATIONSHIPS correspond to VERBS (i.e. actions/states/occurrences). 3) 3) ATTRIBUTES corresponds to ADJECTIVES, which capture the actual information that needs to be collected about each entity.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on C18: We do not have any place in which the microscope acquisition software is described -Caterina Strambio De Castillia
Note on C43: APE = Abstract Parent Entity In the .xsd file, "Group" is used instead of " -APE"
Note on C44: Duplicate LightSource Transmitted vs Fluorescence
Note on C90: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: DONE DIRECT, MISSING LIGHT GUIDE. After that move on to Sample Holder
Note on D279: Change to CutOn CutOff

## 4DN proposal_Settings_v01-07
4DN-OME model
| Tier | Element | Attribute | Complex type | Description | Cardinality | Data type | Allowed values
IMAGE ACQUISITION Settings | 1 | Experiment |  |  | Describes the experiment during which this Image was acquired. The optional Description element may contain free text to further describe the experiment. | 1
| 1 | Experiment | ID |  | A unique identifier for this Experiment. | R
| 1 | Experiment | Name |  | A user assigned name for this Experiment | R | Denomination
| 2 | Experiment | Type | ExperimentType | A term to describe the type of imaging experiment performed in this case. | R | enum | 4D+, Co-localization, DeterministicSuper-Resolution, Electrophysiology, FLIM, FRAP, FRET, HightThroughputScreen, IonImaging, LightSheet, OpticalTrapping, PALM, PGI-Documentation, PhotoAblation, PhotoActivation, SpectralImaging, SPIM, StochasticSuper-Resolution, StructuredIllumination, TimeLapse, Other
| 2 | Experiment | LabellingMethod | ExperimentLabellingMethod | A term that describes the method used for labelling the molecular component under study in this experiment. | R | enum | AptamerMotif, BacteriophageDerivedRNATag, BiarsenicalDye, CASFISH, ChemicalModification, FISH, FluorescentProtein, ForcedIntercalation, Immunocytochemistry, Immunofluorescene, MolecularBeacon, SelfLabellingEnzymeTag, SingleMoleculeFISH, None, Other
| 1 | Experiment | ExperimenterRef |  | This empty element refers to Experimenter that performed this experiment. | 1 … ∞ | Extension of Reference
| 1 | Experiment | Sample |  | A description of the sample that was imaged in this experiment. | 1
| 3 | Experiment | MicrobeamManipulation |  | Defines a microbeam operation and the region of the image it was applied to. The LightSourceSettingsRef element is a reference to the LightSourceSettings that were used in this case, including a reference to the light source that was used for this micromanipulation experiment (e.g. a laser used for photobleaching). | 0 … ∞
| 1 | Experiment/Description | Experiment/Description |  | A description for this imaging Experiment | 0 … 1 | Element
| 2 | Experiment | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | Sample |  |  | A description of the sample that was imaged in this experiment. | 1
| 1 | Sample | ID |  | A unique identifier for this Sample. | R
| 1 | Sample | Name |  | A user assigned name for this Sample. | R | Denomination
| 2 | Sample | MountingMedium |  | The mounting medium the sample is embedded in during acquisition | R | string
| 3 | Sample | MountingMediumRefractiveIndex |  | The refractive index (RI) of the mounting medium. If possible the RI should be measured. Alternatively, please refer to the nominal RI as provided by the manufacturer. | R | float with unit:MountingMediumRefractiveIndexUnit
| 1 | Sample | Sample/Description |  | A description of the sample that was imaged in this case. | 0 .. 1 | Eleement
| 2 | Sample | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | MicrobeamManipulation |  |  | Defines a microbeam operation and the region of the image it was applied to. The LightSourceSettingsRef element is a reference to the LightSourceSettings that were used in this case, including a reference to the light source that was used for this micromanipulation experiment (e.g. a laser used for photobleaching). | 0 … ∞
| 3 | MicrobeamManipulation | ID |  | A unique identifier for this MicrobeamManipulation component. | R
| 3 | MicrobeamManipulation | Type | MicrobeamManipulationType | Defines the type of micromanipulation performed. | R | enum | Fluorescence Recovery After Photobleaching, inverse FRAP, Fluorescence Loss In Photobleaching, Photo-Ablation, Photo-Activation, Uncaging, OpticalTrapping, Other
| 3 | MicrobeamManipulation | ExperimenterRef |  | This empty element refers to Experimenter that performed this experiment. | 0 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | ROIRef |  | This empty element refers to the predefined four dimensional ROI (Region of Interest) that was used for this micromanipulation experiment. | 1 … ∞ | Extension of Reference
| 3 | MicrobeamManipulation | LightSourceSettings |  | This holds a reference to the LightSource (i.e. the ID attribute) that was used in this case as well as the setting that were applied to in in this case. | 1 … ∞ | Element
| 3 | MicrobeamManipulation | MicrobeamManipulation/Description |  | A description for the Microbeam Manipulation experiment | 0 … 1 | plain-text multi-line string
| 3 | MicrobeamManipulation | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 … 1 | Extension of Reference
| 1 | Image |  |  | This element describes the actual image that was acquired and its meta-data. The elements that are references (ending in Ref or Settings) refer to elements defined outside of the Image element. Ref elements are simple links, while Settings elements are links with additional values. If any of the required Image attributes or elements are missing, its guaranteed to be an invalid document.
| 1 | Image | ID |  | A unique identifier for this Image. | R | ImageID
| 1 | Image | Name |  | A user assigned name for this Image | R | Denomination
| 1 | Image | AcquisitionDate |  | The acquisition date of the Image. | R | date
| 1 | Image (was Pixels) | SizeX |  | X size of the pixel data array. | R | integer with unit:none
| 1 | Image (was Pixels) | SizeY |  | Y size of the pixel data array. | R | integer with unit:none
| 1 | Image (was Pixels) | SizeZ |  | Number of Z-planes acquired for this Image. | R | integer with unit:none
| 1 | Image (was Pixels) | SizeC |  | Number of Channels used in this Image | R | integer with unit:none
| 4 | Image (was Pixels) | SizeT |  | Number of Timepoints used in this Image | R | integer with unit:none
| 1 | Image (was Pixels) | PhysicalSizeX |  | Physical width of a pixel expressed in µm. | R | float with unitis:PhysicalSizeXUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeY |  | Physical height of a pixel expressed in µm. | R | float with unitis:PhysicalSizeYUnit; default µm
| 1 | Image (was Pixels) | PhysicalSizeZ |  | Physical distance between optical planes. | R | float with unitis:PhysicalSizeZUnit; default µm
| 4 | Image (was Pixels) | TimeIncrement |  | This attribute is used for time series that have a global timing specification instead of per-timepoint timing info. For example in a video stream. | R | float with unit:TimeIncrementUnit
| 2 | Image (was Pixels) | DimensionOrder | PixelsDimensionOrder | The order in which the individual planes of data are interleaved in the Image |  | enum | XYZCT, XYZTC, XYCTZ, XYCZT, XYTCZ, XYTZC
| 2 | Image (was Pixels) | ChannelOrder |  | The order in which the channels were acquired and are stored |  | string
| 3 | Image (was Pixels) | PixelType | PixelType | The variable type used to represent the pixel intensity in the image | R | enum | int8, int16, int32, uint8, uint16, uint32, float, double, complex, double-complex, bit, Other
| 5 | Image (was Pixels) | SignificantBit |  | The number of bits within the type storing each pixel that are significant. e.g. you can store 12 bit data within a 16 bit type. This does not reduce the storage requirements but can be a useful indicator when processing or viewing the image data. |  | PositiveInteger
| 3 | Image (was Pixels) | Interleaved |  | How the channels are arranged within the data block |  | boolean | TRUE: if channels are stored RGBRGBRGB..; FALSE: if channels are stored RRR..GGG..BBB…
| 5 | Image (was Pixels) | BigEndian |  | This is true if the pixels data was written in BigEndian order. If this value is present it should match the value used in BinData or TiffData. If it does not a reader should honour the value used in the BinData or TiffData. This values is useful for MetadataOnly files and is to allow for future storage solutions. |  | boolean | TRUE: if pixel data was written in BigEndian mode
| 1 | Image | InstrumentRef |  | This empty element refers to the Instrument that was utilized to acquire this image. | 1 | Extension of Reference
| 1 | Image | ExperimentRef |  | This empty element refers to the Experiment during which this Image was acquired. | 1 | Extension of Reference
| 1 | Image | ExperimenterRef |  | This empty element refers to the Experimenter that acquired this image. | 0 … ∞ | Extension of Reference
| 1 | Image | ExperimentGroupRef |  | This empty element refers to the ExperimenterGroup that owns this Image. | 0 … ∞ | Extension of Reference
| 3 | Image | MicroBeamManipulationRef |  | This empty element refers to an optional micro beam manipulation that was performed in connection with the acquisition of this Image (e.g. a microbeam operation, such as a laser-induced photobleaching). | 0 … ∞ | Extension of Reference
| 2 | Image | ROIRef |  | This empty element refers to the predefined four dimensional ROI (Region of Interest) that was used for this micromanipulation experiment. | 0 … ∞ | Extension of Reference
| 3 | Image | TIRFSettings | Extension of Settings | Defines the TIRF settings used for this experiment. | 0 … ∞ | Element
| 4 | Image | ImagingEnvironment |  | This describes the environment that the biological sample was in, during the experiment. | 0 … ∞ | Element
| 1 | Image | MicroscopeSettings | Extension of Settings | This element refers to the microscope used to acquire this image and it holds the settings that were applied to the microscope during this acquisition. | 1 … ∞ | Element
| 2 | Image | ObjectiveSettings |  | This element refers to the objective used to acquire this image and it holds the settings that were applied to the objective during this acquisition. | 1 … ∞ | Element
| 3 | Image | OpticalCalibration |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | 0 … ∞ | Element
| 3 | Image | DetectorCalibration |  | This procedure measures the performance of the dectector | 0 … ∞ | Element
| 1 | Image | Plane |  | Each Image has to have at least one Plane and one Channel. The Plane element holds microscope stage and image timing data for a given channel/z-section/timepoint. The total number of Plane elements for a given Image must equal to SizeZ x SizeC x SizeT. | 1 … ∞ | Element
| 1 | Image | Channel |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. | 1 … ∞ | Element
| 1 | Image | TiffData (or BidData or MetadataOnly) |  | This described the location of the pixel data in a tiff file. | 1 … ∞ | Element | choose one of these three elements
|  |  | BinData |  | The contents of this element are base64-encoded. These are not CDATA sections, just a base64 stream. | or | Element
|  |  | MetadataOnly |  | This place holder means there is on pixel data in this file. | or | Element
| 1 | Image | Image/Description |  | A description for this image. | 0 .. 1 | Element
| 2 | Image (was Pixels) | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | TIRFSettings |  |  | Defines the TIRF conditions that were used for this experiment | 0 … ∞
| 3 | TIRFSettings | ID |  | A unique identifier for this component. | R
| 3 | TIRFSettings | ThroughLens |  | The TIRF modality used in this case. True=Through the lens TIRF modality. False=TIRF modality in which the laser light is not travelling through the objective lens. | R | boolean
| 3 | TIRFSettings | TIRFAngle |  | The incidence angle used in this case. | R | float with unit:TIRFAngleUnit
| 3 | TIRFSettings | Geometry |  | The TIRF geometry used in this case. | R | enum | Spot, Ring, Other
| 3 | TIRFSettings | TIRFSettings/Description |  | A description for the TIRF Settings | 0 .. 1 | plain-text multi-line string
| 3 | TIRFSettings | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 4 | ImagingEnvironment |  |  | This describes the environment in which the biological sample was kept, during the acquisition of this image. | 0 … ∞
| 4 | ImagingEnvironment | Temperature |  | The temperature the acquisition was performed at. | R | float with unit:TemperatureUnit
| 4 | ImagingEnvironment | AirPressure |  | The airpressure the acquisition was performed at. |  | float with unit:AirPressureUnit
| 4 | ImagingEnvironment | Humidity |  | The humidity the acqusition was performed at. |  | float with unit:HumidityUnit
| 4 | ImagingEnvironment | CO2Percent |  | The CO2 concentration, expressed in percentage, the acquisition was performed at. | R | float with unit:CO2PercentUnit
| 1 | MicroscopeSettings |  |  | This element refers to the microscope used to acquire this image and it holds the settings that were applied to the microscope during this acquisition. | 0 … ∞
| 1 | MicroscopeSettings | ID (MicroscopeID) |  | This is the unique identifier of the microscope used in this case. | R
| 3 | MicroscopeSettings | FieldOfView |  | The Field of View defines the area of the specimen that is viewable via the eyepiece. This field caputers the measures size of the Field of View. |  | float with unit:FieldOfViewUnit[default mm]
| 1 | MicroscopeSettings | TotalEffectiveMagnification |  | This is the total effective magnification used in this case. This number can be obtained by multiplying Optovar/Magnification times the Objective/NominalMagnification or even better the Objective/CalibratedMagnification if available. |  | float with unit:none
| 3 | MicroscopeSettings | OptovarRef |  | This empty element refers to the existing Optovar component used in this case. | 0 .. 1 | Extension of Reference
| 3 | MicroscopeSettings | ObjectiveRef |  | This empty element refers to the existing Objective component used in this case. | 1 | Extension of Reference
| 2 | ObjectiveSettings |  |  | This element refers to the objective used to acquire this image and it holds the settings that were applied to the objective during this acquisition. | 0 … ∞
| 2 | ObjectiveSettings | ID (ObjectiveID) |  | This is the unique identifier of the objective used in this case. | R
| 5 | ObjectiveSettings | CorrectionCollar |  | When using high numerical aperture dry objectives (numerical aperture of 0.8 or greater), cover glass thickness variations of only a few micrometers result in dramatic image degradation due to aberration, which grows worse with increasing cover glass thickness. To compensate for this error, the more highly corrected objectives are equipped with a correction collar to allow adjustment of the central lens group position to coincide with fluctuations in cover glass thickness. The CorrectionCollar is normally an adjustable ring on the objective, whose position can be modified. This field records the position of the collar during the acquisition of this image. Each correction collar has an arbitrary scale on it so the values is unit-less. |  | float with unit:none
| 2 | ObjectiveSettings | ImmersiumMedium | ObjectiveSettingsImmersionMedium | The actual immersion medium that was used for the acquisition of this image. | R | enum | Air, Glycerol, Multi, Oil, Water, WaterDripping, Other
| 2 | ObjectiveSettings | ImmersionRefractiveIndex |  | The nominal refractive index of the immersion medium utilized in this case, as specified by the manufacturer. | R | float with unit:none
| 4 | ObjectiveSettings | MeasuredRefractiveIndex |  | The temperature dependenent measured Refractive Index of ther immersion medium utilized in this case. |  | float with unit:none
| 4 | ObjectiveSettings | Temperature |  | The temperature of the objective if different from the ImagingEnvironment/Temperature. |  | float with unit:ObjectiveTemperatureUnit
| 1 | Plane |  |  | Each Image has to have at least one Plane and one Channel. The Plane element holds microscope stage and image timing data for a given channel/z-section/timepoint. The total number of Plane elements for a given Image must equal to SizeZ x SizeC x SizeT. | 1 … ∞
| 1 | Plane | ID |  | A unique identifier for this Plane. | R
| 1 | Plane | TheZ |  | The Z-section to which this plane corresponds. This is numbered from 0 | R | integer with unit:none
| 1 | Plane | TheT |  | The timepoint to which this plane corresponds. This is numbered from 0 | R | integer with unit:none
| 1 | Plane | TheC |  | The channel to which this plane corresponds. This is numbered from 0 | R | integer with unit:none
| 2 | Plane | Timestamp |  | This records the total time intercoursed since the beginning of image acquisitoin. | R | float with unit:TimestampUnit
| 2 | Plane | ExposureTime |  | The length of the exposure utilized to acquire this plane. | R | float with unit:ExposureTimeUnit.
| 4 | Plane | SamplePosition |  | This is used to specify the location of the sample during the acquisition of this plane, with respect to the microscope's reference frame. | 0 .. 1 | Element
| 3 | Plane | FocalPosition |  | This element is used to record the Z-axis position of this plane with respect to the microscope's reference frame | 0 .. 1 | Element
| 4 | Plane | MechanicalCalibration |  | This describes the procedures that were used for mechanical calibration as well as the calibration results. | 0 .. ∞ | Element
| 5 | Plane | HashSHA1 |  | This optional element is a hash of the plane's image data. It is a choice between all the supported hash types. Currently the only method supported is SHA1 | 0 ... 1 | Element
| 3 | Plane | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | FocalPosition |  |  | This element is used to record the Z-axis position of this plane with respect to the microscope's reference frame | 0 .. 1
| 3 | FocalPosition | ID |  | A unique identifier for this Z-axis focal position. | R
| 3 | FocalPosition | Name |  | A user assigned name for this Z-axis focal position. | R | Denomination
| 3 | FocalPosition | StagePositionZ |  | The Z-axis focal position of this plane. | R | float with unit:StagePositionZ
| 3 | FocalPosition | SamplePositioningRef |  | This empty element refers to the focusing device that was used for the acquisition of this image. | 0 .. 1 | Extension of Reference
| 3 | FocalPosition | AutoFocusRef |  | This empty element refers to the optional auto focus device that was used for the acquisition of this image. | 0 .. 1 | Extension of Reference
| 4 | SamplePosition |  |  | This is used to specify the location of the sample during the acquisition of this plane, with respect to the microscope's reference frame. | 0 .. ∞
| 4 | SamplePosition | ID |  | A unique identified for this specific sample holder position. | R
| 4 | SamplePosition | Name |  | A user assigned label for this specific sample holder position. | R | Denomination
| 4 | SamplePosition | SamplePositionX |  | The X position of the SampleHolder within the microscope's reference frame | R | float with unit:PositionXUint
| 4 | SamplePosition | SamplePositionY |  | The Y position of the SampleHolder within the microscope's reference frame | R | float with unit:PositionYUint
| 4 | SamplePosition | SamplePositionZ |  | The Z position of the SampleHolder within the microscope's reference frame | R | float with unit:PositionZUint
| 4 | SamplePosition | SamplePositioningRef |  | This empty element refers to the sample holder that was used for the acquisition of this image. | 0 … 1 | Extension of Reference
| 1 | Channel |  |  | Each Image has to have at least one Plane and one Channel. The Channel element stores information about how the each image channel was acquired. The total number of Channel elements for a given Image must equal to SizeC. | 1 … ∞
| 1 | Channel | ID |  | A unique identifier for this Channel. | R
| 1 | Channel | Name |  | A name for the channel that is suitable for presentation to the user. | R | Denomination
| 5 | Channel | SamplesPerPixel |  | The number of samples the detector takes to form each pixel value. [units:none] Note: This is not the same as "Frame Averaging" - see Integration in DetectorSettings. |  | PositiveInteger
| 2 | Channel | IlluminationType | ChannelIlluminationType | The method of illumination used to capture the channel. | R | enum | Transmitted, Epifluorescence, Confocal, LightSheet, TIRF, NonLinear, Oblique, Other
| 3 | Channel | PinholeSize |  | The Pinhole Size utilized for this channel. |  | float with unit:PinholeSizeUnit.
| 1 | Channel | AcquisitionMode | ChannelAcquistionMode | AcquisitionMode describes the type of microscopy performed for each channel. | R | enum | WideField, LaserScanningConfocalMicroscopy, SpinningDiskConfocal, SlitScanConfocal, MultiPhotonMicroscopy, StructuredIllumination, SingleMoleculeImaging, TotalInternalReflection, FluorescenceLifetime, SpectralImaging, FluorescenceCorrelationSpectroscopy, NearFieldScanningOpticalMicroscopy, SecondHarmonicGenerationImaging, PALM, STORM, STED, TIRF, FSM, LCM, Other, BrightField, SweptFieldConfocal, SPIM
| 1 | Channel | ContrastMethod | ChannelContrastMethod | ContrastMethod describes the technique used to achieve contrast for each channel. | R | enum | Brightfield, Phase, DIC, HoffmanModulation, ObliqueIllumination, PolarizedLight, Darkfield, Fluorescence, Other
| 3 | Channel | ExcitationWavelength |  | The excitation wavelength used for this channel. | R | float with unit:ExcitationWavelengthUnit
| 3 | Channel | EmissionWavelength |  | The emission wavelength used for this channel. | R | float with unit:EmissionWavelengthUnit
| 2 | Channel | Fluorophore |  | This is the name of the fluorophore used to produce this channel [plain text string]. | R | string
| 5 | Channel | NDFilter |  | The NDfilter attribute is used to specify the combined effect of any neutral density filters used. The amount of light the filter transmits at a maximum expressed as a fraction with value from 0.0 to 1.0. NOTE: This was formerly described as "units optical density expressed as a PercentFraction". This was how the field had been described in the schema from the beginning but all the use of it has been in the opposite direction, i.e. as a amount transmitted, not the amount blocked. This change has been made to make the model reflect this usage. |  | float with unit:none
| 5 | Channel | PockelCellSetting |  | The PockelCellSetting used for this channel. This is the amount the polarization of the beam is rotated by. |  | integer with unit:none
| 1 | Channel | Color |  | A color used to render this channel - endocded as RGBA. The default value is "-1" corresponding #FFFFFFFF, so solid white. | R | string
| 3 | Channel | LightSourceSettings |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 0 … ∞ | Element
| 2 | Channel | LightPath |  | A description of the light path | 1 | Element
| 3 | Channel | DetectorSettings-APE |  | This element holds the setting applied to a Detector e as well as a reference to the Detector (the ID attribute) in question. The ID referes to the Detector used in this case. | 0 … ∞ | Element
| 3 | Channel | ExcitationCalibration |  | This describes the procedure that was used for excitation calibration and the resulting calibration measurements. | 0 … ∞ | Element
| 3 | Channel | FieldCalibration |  | This describes the procedure used to measure the field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 0 … ∞ | Element
| 3 | Channel | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | LightSourceSettings |  |  | This holds the setting applied to a LightSource e as well as a reference to the LightSource (the ID attribute) in question. The ID referes to the LightSource used in this case. | 1
| 3 | LightSourceSettings | ID (LightSourceID) |  | A unique identifier for this component. | R
| 3 | LightSourceSettings | Wavelength |  | The nominal wavelength of the excitation light utilized in this case. | R | float with unit:WavelengthUnit
| 3 | LightSourceSettings | PowerSettings |  | The light power output settings utilized in this case. | R | float with unit:PowerSettingsUnit
| 3 | LightSourceSettings | Attenuation |  | The observed attenuation of the excitation light power expressed as a fractional value from 0.0 to 1.0. |  | float with unit:none
| 2 | LightPath |  |  | A description of the light path. | 1
| 2 | LightPath | ID |  | A unique identifier for this component. | R | LSID
| 2 | LightPath | Name |  | A user defined name for this component. | Denomination | string
| 4 | LightPath | LightSourceCouplingRef |  | This empty element refers to a LightSourceCoupling component used in this light path. | 0 … ∞ | Extension of Reference
| 2 | LightPath | FilterSetRef | Reference | This empty element refers to a FilterSet component used in this light path. | 0 … ∞ | Extension of Reference
| 2 | LightPath | ExcitationFilterRef | Reference | This empty element refers to an ExcitationFilter component used in this light path. | 0 … ∞ | Extension of Reference
| 2 | LightPath | StandardDichroicRef | Reference | This empty element refers to a StandardDichroic component used in this light path. | 0 … ∞ | Extension of Reference
| 2 | LightPath | EmissionFilterRef | Reference | This empty element refers to an EmissionFilter component used in this light path. | 0 … ∞ | Extension of Reference
| 2 | LightPath | AdditionalOptics |  | A description of additional optics placed in this light path. | 0 … 1 | Element
| 3 | LightPath | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 2 | AdditionalOptics |  |  | A description of additional optics placed in the light path | 0 … ∞
| 2 | AdditionalOptics | ID |  | A unique identifier for this component. | R
| 2 | AdditionalOptics | Name | Denomination | A user defined name for this component. | R | string
| 2 | AdditionalOptics | AdditionalLensRef | Reference | This empty element refers to an AdditionalLens component placed in this light path. | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics | PrimsRef | Reference | This empty element refers to a Prism component placed in this light path. | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics | PolarizationOpticsRef | Reference | This empty element refers to a PolarizationOptics component placed in this light path. | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics | AdditionalFilterRef | Reference | This empty element refers to an AdditionalFilter component placed in this light path. | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics | AdditionalMirrorRef | Reference | This empty element refers to an AdditionalMirror component placed in this light path. | 0 … ∞ | Extension of Reference
| 2 | AdditionalOptics | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | DetectorSettings-APE |  |  | This element refers to the detector used to acquire this image and it holds the settings that were applied to the detectir during this acquisition. | 0 .. 1
| 3 | DetectorSettings-APE | ID (DetetectorID) |  | A unique identifier for this component. | R | DetectorID
| 3 | DetectorSettings-APE | EffectiveOffset |  | This field reports the actual offset utilized for the acquisition of this image. The offset is an electronic baseline added to the output signal from the sensor to ensure that the displayed signal level is always a positive number of counts. | R | float with unit:none
| 3 | DetectorSettings-APE | Integration |  | This is the number of sequential frames that have been averaged to obtain this image, to improve the overall signal-to-noise ratio. |  | integer with unit:none
| 3 | DetectorSettings-APE | DetectorFieldOfView |  | The detector field of view is the diameter of the field visible by the camera as measured in the specimen plane. It is expressed in mm and it is defined as the detector field number divided by the total effective magnification of the system. In turn the Field Number is defined as the diameter of the view field in mm measured at the intermediate image plane. |  | float with unit:CameraFieldOfViewUnit[default mm]
| 3 | CameraSettings |  |  | This element refers to the Camera used to acquire this image and it holds the settings that were applied to the Camera during this acquisition.
| 3 | CameraSettings | Binning | Binning | This represents the number of pixels that are combined to improve readout rates and signal to noise ratios albeit at the expense of reduced spatial resolution. | R | enum | 1x1, 2x2, 4x4, 8x8, Other
| 3 | CameraSettings | AnalogToDigitalConverter | PixelType | The bit depth used to report the light signal in digital units. | R | enum | int8, int16, int32, uint8, unit16, unit32, float, double, complex, double-complex, bit
| 3 | CameraSettings | EMGain |  | This is the electron multiplication gain (as apposed to the inherent gain) that was applied to the camera for the acquisition of this image. | R | float with unit:none
| 5 | CameraSettings | ReadoutRate |  | The effective pixel read out rate used for the acquisition of this image. This value depends on the bytes per second that can be read from the detector (like a baud rate). |  | float with unit:ReadOutRateUnit
| 3 | CameraSettings | FrameRate |  | This is the effective frame rate used for the acquisition of this image. |  | float with unit:FrameRateUnit
| 3 | CameraSettings | OperatingTemperature |  | This is the effective cooling temperature used during the acquisition of this image. |  | float with unit:TemperatureUnit; default:Celsius[°C]
| 3 | PhotoMultiplierSettings |  |  | This element refers to the PhotoMultiplier used to acquire this image and it holds the settings that were applied to the PhotoMultiplier during this acquisition.
| 3 | PhotoMultiplierSettings | PMTGain |  | In general, the gain is a measure of the ability a two-port circuit to add energy to an input signal and therefore increase its power or amplitude. Specifically, PMTGain refers to the ratio of output to input current set for the acquisition of this image. | R | float with unit:none
| 3 | PhotoMultiplierSettings | PMTVoltage |  | This refers to the voltage applied to the photomultiplier for the acquisition of this image. | R | float with unit:PMTVoltageUnit
| 3 | PhotoMultiplierSettings | EffectiveZoom |  | The zoom factor, aka "Confocal Zoom" or "Scan Zoom", utilized for the acquisition of this image. | R | foat with unit:none
| 3 | OpticalCalibration |  |  | This describes the procedure that was used for optical calibration and the resulting calibration measurements. | 0 … ∞ |  |  | Alternative name: PSFCalibration
| 3 | OpticalCalibration | ID |  | A unique identifier for this component. | R
| 3 | OpticalCalibration | Date |  | The date in which this calibration was performed. | R
| 3 | OpticalCalibration | MedianSpotSize_FWHM |  | Observed median lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MedianSpotSize
| 3 | OpticalCalibration | MedianTheoretical_FWHM |  | Theoretical (expected) medians lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MedianTheoretical
| 3 | OpticalCalibration | MinSpotSize_FWHM |  | Observed minimum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MinSpotSize
| 3 | OpticalCalibration | MinTheoretical_FWHM |  | Theoretical (expected) minimum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MinTheoretical
| 3 | OpticalCalibration | MaxSpotSize_FWHM |  | Observed maximum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the xy-plane. |  | float with unit:FWHM_MaxSpotSize
| 3 | OpticalCalibration | MaxTheoretical_FWHM |  | Theoretical (expected) maximum lateral spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the xy-plane. |  | float with unit:FWHM_MaxTheoretical
| 3 | OpticalCalibration | ZSpotSize_FWHM |  | Observed z-axial spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile. This is a measure of the observed resolution of the system in the z-plane. |  | float with unit:FWHM_ZSpotSize
| 3 | OpticalCalibration | ZTheoretical_FWHM |  | Theoretical (expected) z-axial spot size expressed as the Full Width at Half Maximum (FWHM) of the fitted Point Spread Function (PSF) intensity profile, on the basis of the given optical conditions. This is a measure of the expected resolution of the system in the z-plane. |  | float with unit:FWHM_Ztheoretical
| 3 | OpticalCalibration | Planarity |  | This is a measure of the flatness of the focal plane across the Field of View (FOV). The planarity of the focal plane is calculated from the observed z-position of the best focus of each spot. It indicates focal aberrations or tilt. |  | float with unit:Planarity
| 3 | OpticalCalibration | Asymmetry |  | This field reports the mean +/- standard deviation over the Field of View (FOV) of the observed Point Spread Function (PSF) Asymmetry (i.e. the ratio of the minimum lateral FWMH over the maximum lateral FWHM). It indicates the presence of optical aberrations. |  | float with unit:none
| 3 | OpticalCalibration | LateralChromaticShift |  | The lateral offset of the observed spot position between two color channels. |  | float with unit:LateralChromaticShift
| 3 | OpticalCalibration | AxialChromaticShift |  | The axial offset of the observed spot position between two color channels. |  | float with unit:AxialChromaticShift
| 3 | OpticalCalibration | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | OpticalCalibrationImageRef |  |  | An image obtained by observing the optical standard used in this instance and acquired using the same instrument and settings employed to acquire the image this calibration refers to. | 1 … ∞ | Image
| 3 | OpticalCalibrationImageRef | ID (ImageID) |  | A unique identifier for this component. | R
| 3 | OpticalCalibrationReportRef |  |  | A report generated by the OpticalCalibrationSoftware referred to in OpticalCalibration. | 0 … ∞ | FileURI
| 3 | OpticalCalibrationReportRef | ID (OpticalCalibrationReportID) |  | A unique identifier for this component. | R
| 3 | OpticalCalibrationReportRef | OpticalCalibrationSoftware |  | The name of the software tool utilized to generate the OpticalCalibration report. As an example, such a tool might analyzes microscopy images of fluorescent beads in order to determine quantitatively the resolution and chromatic aberrations of a fluorescence microsocope over the full field of view. An example of such a tool is PSFj (http://www.knoplab.de/psfj/sample-page/). | R
| 3 | OpticalStandard-APE |  |  | The optical standard used in this case. This might be a pre-manufactured slide, such as those produced by Argolight (http://argolight.com/), Geller MicroAnalytical Laboratory (Magnification Reference Standard MRS-4 and others; http://www.gellermicro.com/mag_standards/mrs.html), Zeiss and others. Alternatively, the standard might be a slide containing multi-colored beads of known size and emission characteristics. Finally, slides coated with specific DNA origami patterns are increasingly being used for this purpose as well. | 1 … ∞
| 3 | OpticalStandard-APE | Manufacturer | Manufacturer Spect | The manufacturer of this component. | R | string
| 3 | OpticalStandard-APE | Model | Manufacturer Spect | The model of this component. | R | string
| 3 | OpticalStandard-APE | SerialNumber | Manufacturer Spect | The serial number of this component. | R | string
| 3 | OpticalStandard-APE | LotNumber | Manufacturer Spect | The lot number of this component. |  | string
| 3 | OpticalStandard-APE | SpecsFile |  | Provides the URI for a PDF file with the specifications for the component as provided by the manufacturer. | R | file
| 3 | OpticalStandard-APE | ID |  | A unique identifier for this component. | R
| 3 | OpticalStandard-APE | Name |  | A user defined name for this component. | R | Denomination
| 3 | OpticalStandard-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | ColoredBeadsSlide |  |  | This element describes a slide with colored beads that were used for this calibration procedure. |  | Element
| 3 | ColoredBeadsSlide | Name |  | A user assigned name for this beads preparation | R | Denomination
| 3 | ColoredBeadsSlide | BeadSize |  | The size of the colred beads utilized to make this slide. | R | float with unit:BeadSizeUnit
| 3 | ColoredBeadsSlide | BeadConcentration |  | The concentration of the colored beads preparation utilized to make this slide. | R | float with unit:BeadConcentrationUnit
| 3 | ColoredBeadsSlide | BeadVolume |  | The volume of the colored beads preparation utilized to prepare this slide |  | float with unit:BeadVolumeUnit
| 3 | ColoredBeadsSlide | BeadEmissionWavelenght |  | The peak wavelenght of the light emitted by these beads. In case of multi-color beads list different peak weavelength separated by commas. | R | string
| 3 | ColoredBeadsSlide | Type | BeadType | The type of colred beads used in this case | R | enum | Single color, Multi-color, Other
| 3 | CalibrationSlide |  |  | This element describes a slide used to measure the different optical properties such as magnification, distance, and different kinds of optical distorsions
| 3 | CalibrationSlide | Type | CalibrationSlideType | The type of calibration slide used in this case. Examples include slides produced by Argolight, Geller MicroAnalytical Laboratory, Zeiss and others. | R | enum | ARGO-HM, ARGO-LM, ARGO-SIM, Geller-MRS4, Zeiss universal D=0, Other
| 3 | CalibrationSlide | CalibratedMeasure | CalibratedMeasure | The optical characteristic this slide is used to calibrate, examples include magnification, distance, optical distorsions and others. | R | enum | Magnification, Distance, Multi, Distorsion, Other
| 3 | DNAOrigami |  |  | This element describes an optical standard constructed using the DNA origani technique.
| 3 | DNAOrigami | Type |  | This field describes the type of DNA origami-based optical standard used in this case. Currently this is fild is a simple string. However it might be substituted by an enumeration at a later stage. | R | string
| 3 | FieldCalibration |  |  | This describes the procedure used to measure the field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 0 … ∞
| 3 | FieldCalibration | ID |  | A unique identifier for this component. | R
| 3 | FieldCalibration | Date |  | The date in which this calibration procedure was performed. | R
| 3 | FieldCalibration | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | FlatFieldImageRef |  |  | An image describing field dependent variations in the pixel-to-pixel sensitivity of the detector and/or by distortions in the optical path. | 1 … ∞ | Image
| 3 | FlatFieldImageRef | ID (ImageID) |  |  | R
| 3 | IntensityCalibration-APE |  |  | This describes the procedure(s) that were used for excitation and detector calibration and the resulting calibration results. | 0 … ∞ |  |  | This is for Field, Excitation and Detector calibration
| 3 | IntenstiyCalibration-APE | ID |  | A unique identifier for this component. | R
| 3 | IntenstiyCalibration-APE | Date |  | The date in which this calibration procedure was performed. | R
| 3 | IntenstiyCalibration-APE | IntensityCalibrationToolRef |  | This element is a reference to the tool, if any, used to perform this calibratin procdure | 0 … ∞ | Extension of Reference
| 3 | IntenstiyCalibration-APE | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
| 3 | ExcitationCalibration |  |  | This procedure measures the performance and allows calibration of the wavelength and irradiance of the excitation light.
| 3 | ExcitationCalibration | ObservedExcitationWavelength |  | Observed peak excitation wavelength. | R | float with unit:ObservedExcitationWavelengthUnit
| 3 | ExcitationCalibration | ObservedPowerAtObjective |  | The observed excitation light power measured at the back of the objective. | R | float with unit:ObservedPowerAtObjectiveUnit
| 3 | ExcitationCalibration | ObservedExcitationIrradiance |  | Observed mean excitation light irradiance (radiant power per unit area) as measured at the sample plane over the Field of View (FOV). | R | float with unit:ObservedExcitationPowerUnit
| 3 | ExcitationCalibration | ObservedExcitationVariance |  | Observed excitation light irradiance (radiant power per unit area) variation over the Field of View (FOV). | R | float with unit:none
| 3 | DetectorCalibration |  |  | This procedure measures the performance of the dectector
| 3 | DetectorCalibration | Method |  | This is the method used for this calibration procedure.This is a description of method used for this calibration procedure. | R | string
| 3 | CalibrationMap |  |  | This is a 2D matrix containing values computed as a resulut of the calibration procedure. This matrix is the same size of the image this calibration refers to. | 1 … ∞
| 3 | CalibrationMap | MetricType |  | This field captures the type of metric that was computed as part of the calibratin procedure in this instance. | R | enum | Gain, LinearOffset, DetectionLimit, LinearRangeMax, Variance, Mean, Other
| 3 | CalibrationMap | IlluminationIntensity |  | This attribute only applies to matrices of type Variance and Mean. It represents the illumination intensity that was used to acquire the images used to contruct this matrix. The value can range from 0 (i.e. dark image) to any intenisty value experssed either as percentage or as absolute power. | R | float with unit:none
| 3 | CalibrationMap | NrOfImages |  | This field represents the number of images that were employed to generate this calibration matrix. | R | integer with unit:none
| 3 | Row |  |  |  | 1 … ∞
| 3 | Row | RowNr |  | The index of this row | R | integer with unit:none
| 3 | Cell |  |  |  | 1 … ∞
| 3 | Cell | CellNr |  | The index of this cell | R | integer with unit:none
| 3 | Cell | CellValue |  | The value of the metric contained in this cell | R | float with unit:CellValueUnit
| 4 | MechanicalCalibration |  |  | This describes the procedures that were used for mechanical calibration and the resulting calibration results. | 0 … ∞
| 4 | MechanicalCalibration | ID |  | A unique identifier for this component. | R
| 4 | MechanicalCalibration | Date |  | The date in which this calibration was performed. | R
| 4 | MechanicalCalibration | MeasuredLateralDrift |  | Measured shift in (x, y) position over time. | R | float with unit:MeasureLaterlDriftUnit
| 4 | MechanicalCalibration | MeasuredFocalDrift |  | Measured shift in z position over time. | R | float with unit:MeasureFocalDriftUnit
| 4 | MechanicalCalibration | MeasuredRepeatability |  | Measured precision with which this device is capable of returning to a previously visited position (useful for tiling). | R | float with unit:MeasuredRepeatablityUnit
| 4 | MechanicalCalibration | MeasuredSettlingTime |  | Measured time it takes for this device to stabilize at the target position. | R | float with unit:MeasuredSettlingTimeUnit
| 4 | MechanicalCalibration | SamplePositioningRef |  | This empty element refers to the focusing device that was used for the acquisition of this image. | 0 .. 1 | Extension of Reference
| 4 | MechanicalCalibration | AutoFocusRef |  | This empty element refers to the optional auto focus device that was used for the acquisition of this image. | 0 .. 1 | Extension of Reference
| 4 | MechanicalCalibration | AnnotationRef | CommentAnnotation | This is a simple multi-line comment or annotation describing this component. | 0 .. 1 | Extension of Reference
Note on A1: This is a verbose description of the all ELEMENTS and FIELDS of the OME data model 4DN extension. It is quite difficult to interpret if you are not familiar with this sort of things. We recommend that before plowing through this document you familiarize yourself with the attached graphical depiction of the model, which is provided alongside this document. The graphical depiction is presented in what is called the Entity-Relationship formalism in which each box represents a real-life entity one needs to collect information about.
Note on B2: This column provides the TIER LEVEL assignment for each element and attribute Please refer to the Tier system worksheet for a description of the tiers.
Note on C158: Need to enter a Ref to a optical diagram describing how the different components are arranged
Note on D172: HOW DOES THIS GET MEASURED? -Caterina Strambio De Castillia