## Tier system_v02-00
-- required metadata --
Category | Tier Nr. | Name | Description | Example | Example Labelling | Optical calibration | Intensity calibration | Mechanical calibration | Experiment/ Sample | Microscope hardware specifications | Image acquisition settings
Descriptive | 1 | Minimum Information About Light Microscopy/ Documentation/ Material and Methods | Reporting qualitative effects without quantification; identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required) | recommended annually (not required) |  | - experiment description and date - sample description - experimenter name - mounting medium | - microscope manufacturer, name and type - light source manufacturer, wavelenght and type - objective manufacturer, magnification, NA and correction - filter/dichroic transmittance range - detector manufacturer and type | - acquisition date - immersion medium name and refractive index - calibrated magnification - exposure time; pixel dwell time - channel name, color, contrast method and acquisition mode, illumination type, fluorophore - image dimension order and number, channel order - physical pixel size x, y, and z
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of refraction-limited particles, super-resolution microscopy; tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing; cell and single particle tracking, dynamic expression level quantification | All of the above + Fluoresent Protein (FP) labelling, Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling | required monthly | highly-recommended quarterly to monthly |  | - imaging temperature, air pressure, humidity and CO2 percentage - mounting medium measured refractive index - high precision coverglass | detailed microscope table, light source and light source coupling, magnification, focusing and autofocus, light path, filter, dichroic, additional optics and detector specification & settings | - acquisition time settings - sample positioning settings - objective temperature - measured immersion medium refractive index
| 3 | Manufacturing/ Full Documention/ Technical Development | Full reproducibility of microscopic set-up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition | required quarterly to monthly |  | all the metadata specified by the data model - including any novel technology-specific metrics

## Tier system_v02-00 SUMMARY
-- required metadata --
Category | Tier Nr. | Name | Description | Example | Example Labelling | Quality Control | Experiment/ Sample | Microscope hardware specifications | Image acquisition settings
Descriptive | 1 | Minimum Information About Light Microscopy/ Documentation/ Material and Methods | Reporting qualitative effects without quantification; identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required) | - experiment description and date - sample description - experimenter name - mounting medium | - microscope manufacturer, name and type - light source manufacturer, wavelenght and type - objective manufacturer, magnification, NA and correction - filter/dichroic transmittance range - detector manufacturer and type | - acquisition date - immersion medium name and refractive index - calibrated magnification - exposure time; pixel dwell time - channel name, color, contrast method and acquisition mode, illumination type, fluorophore - image dimension order and number, channel order - physical pixel size x, y, and z
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of refraction-limited particles, super-resolution microscopy; tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing; cell and single particle tracking, dynamic expression level quantification | All of the above + Fluoresent Protein (FP) labelling, Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling | required monthly/ highly recommended | - imaging temperature, air pressure, humidity and CO2 percentage - mounting medium measured refractive index - high precision coverglass | detailed microscope table, light source and light source coupling, magnification, focusing and autofocus, light path, filter, dichroic, additional optics and detector specification & settings | - acquisition time settings - sample positioning settings - objective temperature - measured immersion medium refractive index
| 3 | Manufacturing/ Full Documention/ Technical Development | Full reproducibility of microscopic set-up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition/monthly/quarterly | all the metadata specified by the data model - including any novel technology-specific metrics

## Tier system_v02-00 MINIMAL
Category | Tier Nr. | Name | Description | Example | Example Labelling | Quality Control
Descriptive | 1 | Minimum Information About Light Microscopy/ Documentation/ Material and Methods | Reporting qualitative effects without quantification; identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, study of cellular sub-compartments | FISH, Immuno Fluorescence | recommended quarterly (not required)
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of refraction-limited particles, super-resolution microscopy; tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing; cell and single particle tracking, dynamic expression level quantification | All of the above + Fluoresent Protein (FP) labelling, Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling | required monthly/ highly recommended
| 3 | Manufacturing/ Full Documention/ Technical Development | Full reproducibility of microscopic set-up and image acquisition | Development of novel unproved technology or of new gold-standard; full reproducibility ofmicroscopy set up and image acquisition settings | all of the above | required for every acquisition/monthly/quarterly